## Sheet1
Date | Time |  | Emax (MV/m) Ptran | Etran_per | error EPTrans (MV/m) |  | E (MV/m) emitted | Eemit_per | E (MV/m) Forward | EPF_per | Measured Qo | error Qo | + error Qo | -error Qo | Qo at 2K | Qo at 2.07K | Qo at 2.08K | RF Heat Load (W) | RF Heat Load at 2.07K (W) | RF Heat Load at 2.08K (W) | Static Heat Load (W) | Static dPdt 0 (torr/s) | Htr dPdt (torr/s) | RF dPdt (torr/s) | Initial He Pressure (torr) | Final He Pressure (torr) | Static dPdt 1 (torr/s) | Static dPdt 2 (torr/s) | Average Pressure | temp from pressure (K) | Measurement Interval (s) | Actual Heater Power (W) | Heater Voltage Input (V) | 29 Watt Line | Heat load by Curve Fit (W) | Qo by Heat Load Curve fit | Heat Load from Qo Curvefit | error from measured | error from measured
2016-10-07 | 10:21 | 18105251.74 | 5.0061 | 4.8806 | 0.24433 | 18036777.78 | 18.0368 |  |  |  | 16498865556.0193 | 55394434613.9104 | 23501134443.9807 | 16498865555.0193 | 16021040881.0789 | 12794146377.5584 | 12397231395.8516 | 0.22348 | 0.28819 | 0.29741 | 7.5536 | 0.025408 | 0.064962 | 0.02053 | 15.3717 | 23.5865 | 0.020291 | 0.019418 | 22.7249 | 1.991 | 20 | 13.922
2016-10-07 | 10:26 | 17087452.91 | 5.9878 | 4.8806 | 0.29224 | 43884191.35 | 43.8842 |  |  |  | 12208124926.7152 | 21223433614.3429 | 21223433614.3429 | 12208124925.7152 | 15485227187.9084 | 12450121072.8849 | 12073950822.0914 | 0.4321 | 0.4237 | 0.4369 | 6.6997 | 0.019129 | 0.059092 | 0.019856 | 24.7493 | 30.6379 | 0.0187 | 0.018572 | 29.8411 | 2.0764 | 20 | 14.2309
2016-10-07 | 10:41 | 14074935.56 | 8.1473 | 4.8806 | 0.39764 | 16508086.47 | 16.5081 |  |  |  | 10318383176.3621 | 8238120576.7721 | 8238120576.7721 | 8238120576.7721 | 11859660845.506 | 9993776844.9976 | 9749943434.6002 | 0.94649 | 0.97723 | 1.0017 | 9.7982 | 0.032148 | 0.06822 | 0.025407 | 22.1415 | 29.1814 | 0.02347 | 0.021972 | 28.2277 | 2.057 | 20 | 14.2379
2016-10-07 | 10:46 | 13092027.01 | 8.985 | 4.8806 | 0.43852 | 20031433.45 | 20.0314 |  |  |  | 10025971135.015 | 6422195065.3434 | 6422195065.3434 | 6422195065.3434 | 15390311541.2983 | 12388692297.2546 | 12016169395.1467 | 1.1847 | 0.95875 | 0.98848 | 7.5289 | 0.021267 | 0.059479 | 0.022296 | 30.2345 | 35.9662 | 0.019883 | 0.018234 | 35.1634 | 2.1403 | 20 | 14.2359
2016-10-07 | 11:00 | 10080942.9 | 11.038 | 4.8806 | 0.53872 | 17155105.06 | 17.1551 |  |  |  | 8228862967.2319 | 2944769977.9443 | 2944769977.9443 | 2944769977.9443 | 11434150858.7952 | 9689910241.2676 | 9460509019.0398 | 2.1784 | 1.8499 | 1.8948 | 6.4872 | 0.01955 | 0.061498 | 0.024905 | 30.2656 | 35.8078 | 0.018973 | 0.017901 | 34.9657 | 2.1379 | 20 | 14.2249
2016-10-07 | 11:09 | 8103547.94 | 11.9631 | 4.8806 | 0.58387 | 11335447.37 | 11.3354 |  |  |  | 8561259347.9403 | 2741165062.0283 | 2741165062.0283 | 2741165062.0283 | 9773881978.8963 | 8470533424.506 | 8294711307.4707 | 2.4595 | 2.4858 | 2.5385 | 8.7013 | 0.033971 | 0.075708 | 0.031612 | 23.1281 | 29.9142 | 0.024289 | 0.022468 | 28.8833 | 2.0649 | 20 | 13.9131
2016-10-07 | 11:14 | 7015687.421 | 13.2 | 4.8806 | 0.64424 | 16009652.83 | 16.0097 |  |  |  | 6148649429.3313 | 1258408609.6938 | 1258408609.6938 | 1258408609.6938 | 8235524736.1458 | 7290330207.2418 | 7159711851.3414 | 4.1693 | 3.5164 | 3.5805 | 6.6323 | 0.02154 | 0.065071 | 0.032169 | 31.0737 | 37.8337 | 0.019841 | 0.018483 | 36.9361 | 2.1615 | 20 | 14.2259
2016-10-07 | 12:49 | 22086291.46 | 13.8836 | 4.8806 | 0.6776 | 48994980.29 | 48.995 |  |  |  | 8303224946.6674 | 1995313873.675 | 1995313873.675 | 1995313873.675 | 9336237421.5127 | 8139851982.9855 | 7977358082.4528 | 3.4155 | 3.484 | 3.555 | 8.0989 | 0.027812 | 0.062449 | 0.029633 | 22.3028 | 29.3963 | 0.019982 | 0.019128 | 28.4819 | 2.0601 | 20 | 13.0657
2016-10-07 | 12:53 | 20089963.56 | 14.9684 | 4.8806 | 0.73055 | 56366485.48 | 56.3665 |  |  |  | 6754093628.8894 | 1229570966.8134 | 1229570966.8134 | 1229570966.8134 | 8893048438.7029 | 7800907891.0152 | 7651540599.4201 | 4.8807 | 4.2257 | 4.3082 | 7.5038 | 0.018702 | 0.054272 | 0.029907 | 30.1922 | 36.3452 | 0.017713 | 0.018415 | 35.4013 | 2.1431 | 20 | 14.8633
2016-10-07 | 12:59 |  | 16.1622 | 4.8806 | 0.78881 | [5 empty] | 6558465481.3152 | 1055666149.0267 | 1055666149.0267 | 1055666149.0267 | 7397238194.3365 | 6625657569.3649 | 6517594379.1125 | 5.8599 | 5.8005 | 5.8967 | 10.2298 | 0.032609 | 0.066503 | 0.037654 | 23.264 | 32.2037 | 0.023033 | 0.020402 | 29.8258 | 2.0762 | 20 | 14.2239
2016-10-07 | 13:09 |  | 7.035 | 4.8806 | 0.34335 | [5 empty] | 12709705140.7099 | 16684570518.2196 | 16684570518.2196 | 12709705139.7099 | 14676155208.8845 | 11921712911.2245 | 11576353026.3691 | 0.57291 | 0.61078 | 0.629 | 10.3685 | 0.036071 | 0.073006 | 0.026369 | 21.3174 | 28.5462 | 0.025485 | 0.023851 | 27.5022 | 2.0483 | 20 | 14.2259
2016-10-07 | 13:12 |  | 10.0841 | 4.8806 | 0.49216 | [5 empty] | 9127663976.998 | 4270453025.7114 | 4270453025.7114 | 4270453025.7114 | 12821951700.4544 | 10668479015.6869 | 10391067713.1912 | 1.6391 | 1.4024 | 1.4398 | 7.8116 | 0.023062 | 0.063216 | 0.025858 | 29.3915 | 35.1739 | 0.021766 | 0.020543 | 34.2484 | 2.1293 | 20 | 14.2198
2016-10-07 | 13:18 |  | 13.1003 | 4.8806 | 0.63937 | [5 empty] | 8284424227.1097 | 2193124120.1169 | 2193124120.1169 | 2193124120.1169 | 9196451168.6058 | 8033391665.7604 | 7875079118.6672 | 3.0478 | 3.1431 | 3.2063 | 8.843 | 0.035796 | 0.074971 | 0.031357 | 21.7332 | 29.0195 | 0.021687 | 0.021215 | 28.0089 | 2.0544 | 20 | 14.2239
[15 empty] | 20.9 | 8148981900 | 8024170800
[16 empty] | 30.2278 | 30.698

## Qo history data
|  | Qo Acc (7/11) |  | Emax (MV/m) Ptran |  | Qo Com (9/11) |  |  |  | Qo ReCom (12/13) |  | Eacc | std Eacc | Qo VTA (12/10) | std Qo
5.9712 | 0.23502 | 11504763008.331 | 4968586137.714 | 5.6191 | 0.32591 | 15330223084.9156 | 7242753419.7627 | 21.608 | 0.87747 | 7841251815.3159 | 622614427.1684 | 2.0791 | 0.14616 | 12826134900 | 1872501560
7.1382 | 0.28094 | 12800015353.5187 | 3945749349.8644 | 7.2937 | 0.42303 | 15126895064.0224 | 4758082964.7006 | 21.0544 | 0.85499 | 8052153314.7463 | 656931149.3528 | 2.4925 | 0.17515 | 13510183300 | 1960955490
8.0907 | 0.31843 | 12560494335.5226 | 2757344758.43 | 8.1546 | 0.47296 | 15557622166.6779 | 3697264887.2078 | 19.0631 | 0.77413 | 9517516715.0293 | 780818563.1355 | 2.9765 | 0.2091 | 13798455200 | 1993482180
8.1828 | 0.31831 | 12684282784.3833 | 3072771128.9065 | 9.204 | 0.53383 | 13697462257.1376 | 2741524074.1053 | 18.1053 | 0.73523 | 9973033088.6041 | 865296370.8713 | 3.5339 | 0.24821 | 13797522700 | 1986833730
10.0237 | 0.38991 | 11840987950.3893 | 1776548102.2372 | 10.0381 | 0.58221 | 13111326942.5434 | 2108130892.1602 | 17.0875 | 0.6939 | 10288077511.4943 | 907141448.5826 | 3.9468 | 0.27718 | 13746617100 | 1976558410
10.1678 | 0.40016 | 11233889411.0203 | 1999815232.9717 | 11.1075 | 0.64424 | 12449744007.42 | 1878491223.1556 | 16.0517 | 0.65184 | 10770318123.0931 | 989700834.1457 | 4.5485 | 0.31941 | 13669783200 | 1962324480
11.2431 | 0.44248 | 11296390053.2654 | 1579872805.3534 | 11.9852 | 0.69514 | 11876158398.5042 | 1540555264.1255 | 15.0332 | 0.61048 | 11203722269.822 | 1083124376.6972 | 4.9466 | 0.34735 | 13524843800 | 1939429890
12.1848 | 0.47954 | 11226182432.4243 | 1286579851.3297 | 12.8576 | 0.74574 | 11509701758.8905 | 1484333373.5134 | 14.0749 | 0.57156 | 10960857500.7398 | 1118496417.5947 | 5.4689 | 0.38401 | 13435195600 | 1924927320
14.0813 | 0.55418 | 9797623169.5898 | 1126128125.9697 | 13.6848 | 0.79372 | 10981122235.239 | 1274681633.5575 | 13.092 | 0.53165 | 11744880237.4793 | 1328442328.7097 | 5.4489 | 0.3826 | 13472621200 | 1930194260
14.1874 | 0.55186 | 9970949919.8022 | 1167561764.68 | 14.6951 | 0.85232 | 10654012749.3048 | 1262342038.7296 | 12.0432 | 0.48906 | 12278346701.1539 | 1570509158.7072 | 5.9429 | 0.41728 | 13415134800 | 1920771340
15.0343 | 0.59168 | 9502928084.9879 | 918641560.4463 | 15.5236 | 0.90037 | 10195240854.995 | 1110807158.9149 | 11.0603 | 0.44914 | 12707477913.8433 | 1870377276.7349 | 6.5661 | 0.46103 | 13304832900 | 1903651170
16.0515 | 0.63171 | 9102958086.1081 | 893226144.8875 | 16.4841 | 0.95608 | 9764985390.1284 | 1089099112.1293 | 10.0809 | 0.40937 | 13035825401.6728 | 2205136225.0804 | 6.9359 | 0.48698 | 13234804400 | 1892978370
16.1958 | 0.62998 | 9484559950.1546 | 914170224.3364 | 17.4211 | 1.0104 | 8142475193.4632 | 905873112.4012 | 9.0545 | 0.36769 | 13547661722.0531 | 2838024551.3311 | 7.4225 | 0.52114 | 13122400700 | 1876123590
17.155 | 0.67514 | 9148587608.1102 | 803465444.3819 | 18.228 | 1.0572 | 8281818407.3383 | 881832684.8055 | 8.1035 | 0.32907 | 14559766325.2092 | 4005055630.378 | 7.9332 | 0.55699 | 12944217400 | 1851385290
18.0871 | 0.70354 | 8304047175.7975 | 682842757.9734 | 19.2338 | 1.1156 | 7145080307.3847 | 785503022.8761 | 7.0157 | 0.2849 | 14570591762.0833 | 5262263091.3523 | 8.5402 | 0.5996 | 12842823000 | 1834208880
18.1634 | 0.71482 | 8938328092.1343 | 712212172.3563 | 18.2117 | 1.0563 | 8565933556.2857 | 947659773.8746 | 23.2384 | 0.94368 | 6880813476.7226 | 539851699.839 | 8.9929 | 0.63138 | 12714123900 | 1815235340
|  |  |  | 19.0969 | 1.1076 | 8441835637.8 | 890888615.6072 | 23.0175 | 0.93471 | 6782775362.8778 | 538451405.1742 | 9.4993 | 0.66692 | 12608350100 | 1799635240
|  |  |  | 20.0298 | 1.1617 | 7249522731.7408 | 794651376.6691 | 22.0863 | 0.89689 | 7452177856.8227 | 596684114.1736 | 10.013 | 0.70299 | 12491920800 | 1782624750
[8 empty] | 20.09 | 0.81582 | 8738984573.0944 | 721290919.9119 | 10.5628 | 0.74158 | 12362907600 | 1763824330
|  |  |  | 6.8808 | 0.39909 | 13556444504.745 | 4288846243.5296 |  |  |  |  | 10.9867 | 0.77134 | 12264677400 | 1749912120
|  |  |  | 7.9984 | 0.46391 | 13798418440.9248 | 3043508920.7146 |  |  |  |  | 11.529 | 0.80941 | 12171389800 | 1736267560
|  |  |  | 10.0028 | 0.58016 | 12092936884.0158 | 1990336828.9524 |  |  |  |  | 12.0066 | 0.84293 | 12052294900 | 1719245440
|  |  |  | 12.012 | 0.6967 | 11331633757.3071 | 1404214106.1142 |  |  |  |  | 12.5174 | 0.87879 | 11912859300 | 1698889320
|  |  |  | 13.8532 | 0.80348 | 10499385605.146 | 1262616430.6154 |  |  |  |  | 12.9864 | 0.91172 | 11789217700 | 1681184310
|  |  |  | 15.754 | 0.91373 | 9840458154.1428 | 1107291026.5056 |  |  |  |  | 13.5444 | 0.95089 | 11686278900 | 1666761220
|  |  |  | 17.7798 | 1.0312 | 9145265069.4617 | 1001607419.0216 |  |  |  |  | 14.0968 | 0.98967 | 11582307100 | 1652169270
|  |  |  | 19.8784 | 1.1529 | 8155434062.5106 | 878205072.6662 |  |  |  |  | 14.5018 | 1.0181 | 11500127700 | 1640603460
|  |  |  | 20.7418 | 1.203 | 7796735646.4187 | 837700862.9132 |  |  |  |  | 15.0894 | 1.0594 | 11391572600 | 1625429280
[12 empty] | 15.4959 | 1.0879 | 11335521600 | 1617612630
[12 empty] | 16.1022 | 1.1305 | 11270486000 | 1608370920
[12 empty] | 16.4579 | 1.1554 | 11198552600 | 1598425930
[12 empty] | 16.92 | 1.1879 | 11156362800 | 1592713250
[12 empty] | 17.5022 | 1.2287 | 11075115500 | 1581522470
[12 empty] | 18.0006 | 1.2637 | 11028397100 | 1575246750
[12 empty] | 18.4831 | 1.2976 | 10978618800 | 1568589530
[12 empty] | 18.9858 | 1.3329 | 10915793200 | 1560089990
[12 empty] | 19.4622 | 1.3663 | 10896318800 | 1557770080
[12 empty] | 20.0039 | 1.4044 | 10818076000 | 1547137160
[12 empty] | 20.5625 | 1.4436 | 10778630100 | 1542026900
[12 empty] | 20.9079 | 1.4678 | 10756700400 | 1539472670
[12 empty] | 21.6463 | 1.5197 | 10665616900 | 1527380650
[12 empty] | 21.9848 | 1.5434 | 10662256400 | 1527330750
[12 empty] | 22.4863 | 1.5786 | 10602391600 | 1520488790
[12 empty] | 23.1764 | 1.6271 | 10556321700 | 1513351650
[12 empty] | 23.4251 | 1.6445 | 10533303300 | 1510432320
[12 empty] | 24.0053 | 1.6853 | 10488359900 | 1504779540
[12 empty] | 24.4841 | 1.7189 | 10389832200 | 1491459780
[12 empty] | 25.027 | 1.757 | 10325303500 | 1482950500
[12 empty] | 25.3695 | 1.781 | 10286222100 | 1477945500
[12 empty] | 25.6542 | 1.801 | 10246279700 | 1472679750
[12 empty] | 25.9198 | 1.8197 | 10211179300 | 1468089450
[12 empty] | 26.4763 | 1.8587 | 10140376100 | 1458992480
[12 empty] | 27.0328 | 1.8978 | 10061685200 | 1448849430