## Slide 3: History
Linac Single Pass 100% in Hall pre2009
PREXI – 2009
RF modules changed to transport type in HallA post PREX – ‘Hybrid’ type – timing,IF LSP & RF transport
HallC qweak experiment – ‘Hybrid’ type BPMs, Qweak ADCs
2016 – appearance in 1H04A,E jumps (with varying severity, more apparent at higher currents 60-80uA sometimes) – Qweawk ADCs
2018-2019 – lower current running (20-30uA), jumps appeared more subtly until 50uA running, where more apparent – Qweak ADCs
Thus far reproducing 30mV jumps requires beam on and only appears in bpm1H04A,E not other bpms.

## Slide 4
170,500us
~4.167ms,33.330ms
240Hz linesync, 30Hz(29.556Hz) freeclock
3.968,31.328ms
0.99,8.282ms
4 subblocks
30Hz multiplets can be formed in data processing to cancel 60Hz data when analyzing helicity correlated asymmetries

## Slide 6
30Hz Data

## Slide 7
30Hz Data

## Slide 8
30Hz Data

## Slide 9
Aside: in 2009, in the injector, when RF modules used with LSP IF/timing cards, saw tails

## Slide 10
240Hz Data

## Slide 11
240Hz Data

## Slide 14
240Hz /(960Hz) beam on data 2019
BPM trigger plugged into DAQ recently
separate run

## Slide 21
Trigger on data type:
-Synched to beam sync
-switch clock pattern alternating
-switch clock pattern interrupts
No trigger on data type:
-Not synched to beam sync - freeclock
-switch clock pattern constant
-No switch clock pattern interrupts

## Slide 22
Trigger on data type(0.8sec):
-Synched to beam sync
-switch clock pattern alternating
-switch clock pattern interrupts
No trigger on data type (0.2sec):
-Not synched to beam sync - freeclock
-switch clock pattern constant
-No switch clock pattern interrupts

## Slide 23
Aside: Trigger pickup
5V 280us long trigger signal pickup in wires channels of 20-30mV. Averages over 33ms to 0.2mV. Same on all wires. Simply a pedestal shift.

## Slide 24
Run1691_6kHzminuswiresAM

## Slide 25
Run1690_sawtoothAMmodoncurrent6kHz_240Hzdata.png

## Slide 40
https://logbooks.jlab.org/entry/3677268
https://logbooks.jlab.org/entry/3679224

## Slide 41
 | Volts | Volts | mV
--bpm4a-- | Mean(0.8sec data type) | Mean(0.2sec data type) | Shift (0.2stype-0.8stype)
xp | 2.638 | 2.638 | 0.14
xm | 1.966 | 1.966 | 0.12
yp | 2.607 | 2.607 | -0.79
ym | 2.505 | 2.508 | 2.24
xp-xm | 0.673 | 0.673 | -0.08
yp-ym | 0.102 | 0.099 | -3.19
 | mm | mm | um
x_unrotated | 2.75 | 2.75 | -0.6
y_unrotated | 0.38 | 0.36 | -11.8
x | 1.68 | 1.69 | 8.3
y | 2.21 | 2.20 | -8.8

## Slide 42
--bpm4e-- | Volts | Volts | mV
xp | 2.574 | 2.576 | 1.47
xm | 1.946 | 1.946 | -0.22
yp | 2.556 | 2.557 | 1.58
ym | 2.437 | 2.436 | -1.54
xp-xm | 0.628 | 0.630 | 1.53
yp-ym | 0.118 | 0.121 | 3.07
 | mm | mm | um
x_unrotated | 2.61 | 2.62 | 5.5
y_unrotated | 0.45 | 0.46 | 11.6
x | 1.53 | 1.53 | -4.2
y | 2.16 | 2.18 | 11.9

## Slide 43
--bpm14-- | Volts | Volts | mV
xp | 2.745 | 2.746 | 0.26
xm | 2.644 | 2.645 | 0.33
yp | 2.692 | 2.692 | 0.11
ym | 2.729 | 2.729 | 0.29
xp-xm | 0.101 | 0.101 | -0.01
yp-ym | -0.037 | -0.037 | -0.13
 | mm | mm | um
x_unrotated | 0.35 | 0.35 | 0.0
y_unrotated | -0.13 | -0.13 | -0.5
x | 0.34 | 0.34 | 0.4
y | 0.16 | 0.16 | -0.2

## Slide 44
--bpm18-- | Volts | Volts | mV
xp | 2.470 | 2.470 | 0.42
xm | 2.005 | 2.005 | 0.04
yp | 2.454 | 2.454 | 0.01
ym | 2.323 | 2.324 | 0.45
xp-xm | 0.464 | 0.464 | 0.26
yp-ym | 0.131 | 0.131 | -0.35
 | mm | mm | um
x_unrotated | 1.95 | 1.95 | 0.7
y_unrotated | 0.52 | 0.51 | -1.5
x | 1.02 | 1.02 | 1.7
y | 1.74 | 1.74 | -0.5