JLab 5 MeV Mott Polarimeter – Data Analysis
Joe Grames, Riad Suleiman, Daniel Moser, Marcy Stutzman
May 2017
Very rough draft

Abstract
This note describes the general technique used to analyze individual Mott data runs, followed by how background “dilution” and other sources of error are accounted for. Following this, final asymmetries and rates from Run I and II for a given foil thickness are presented.

General Technique
The Mott DAQ produces a raw data file for each data run that is then decoded into a ROOT tree such that each scalar has a unique branch. The analysis consists of ROOT-interpretted C++ code that an individual run's ROOT tree is passed to.
There are three main sub-routines in the Mott analysis code that are executed sequentially – the first loop in which energy spectra are cut and fit in order to determine “good” elastic scatterings from the target foil, the second loop in which the determined “good” scatterings are broken down by helicity and asymmetries are calculated, and the scaler loop in which charge asymmetry is calculated. Rates are also calculated in the second loop.

General Technique – First Loop
From a run's ROOT tree, in the first loop sub-routine, “raw” data histograms are filled. Eight histograms corresponding to each of the 8 PMTs – Left, Right, Up, Down for Energy and dEnergy – are filled after pedestal subtraction. These histograms are helicity-independent, that is, events with either helicity state are present.
Pedestal subtraction is ….... (Riad)
PMT histograms are binned 10Ch/Bin from 0 to 13000. Time-of-flight (ToF) histograms for each detector are also filled at this time. These histograms are binned 10 bins/ns, or 100 ps/bin (TDC resolution is 34ps/channel), from 40 to 80 ns. Additionally, for display, 2-dimensional Energy vs ToF, dE vs ToF, and Energy vs dE histograms are created. The energy axis follows the same binning as the 1D histograms, while ToF axis is expanded to the full 90 ns window and 2 bins/ns.

Figure 1: "Raw" Energy Spectra
Figure 2: dE Spectra

Figure 3: Time-of-Flight Spectra, Vertical Log Scale
Figure 4: Energy vs Time-of-Flight

Figure 5: dE vs Time-of-Flight
Figure 6: Energy vs dE

Figures 1 through 6 are sample spectra from the Left detector of Mott Run 8545 – Run II, 31MHz beam repitition rate, transversely vertically linearly polarized electrons, scattering off of a 350 nm gold foil. Corresponding spectra for the other three detectors are similar. Figure 1 shows a typical E-detector “raw”, or uncut, spectra. Figure 2 shows a typical dE-detector spectra. Figures 3 shows a typical Time-of-Flight spectra. Run 8545 was performed with a hardware timing veto wired into the FADC, and so from roughly 61 to 74 ns there appears to be no data. Without this veto, a second peak corresponding to scatterings from the dump would be present. The peak around 54 ns is scatterings from the target foil. Figure 4 shows “raw” energy spectra vs ToF, Figure 5 dE vs ToF, and Figure 6 E vs dE.
After filling histograms, when running at a suitable beam repitition rate, each detector's Time-of-Flight spectra's target peak is fit with a Gaussian. To do this, the maximum bin between 49 and 55 ns is found and used as the seed value for the mean of the Gaussian fit. The amplitude seed value is 1000 counts, and the sigma seed value is 1 ns. The fit is restricted to the 49 to 55 ns range. The default ROOT TH1 class fitter is used – a Chi^2 function minimized using Minuit and the MIGRAD minimizer. Figure 7 shows run 8545, Left detector Time-of-Flight spectra, vertical log-scale, fit with a Gaussian in light red.

Figure 7: Time-of-Flight Spectra, Target Peak Fit with Gaussian
From this fit, the time-window that “good” Mott scattering events from the target foil occur within is determined as from (mean – 2 sigma) to (mean + 2 sigma). The choice of this +/- 2 sigma window about the mean is explained in detail in section Time-of-Flight Cuts.
Next, the uncut, “raw” energy spectra are Time-of-Flight cut – for each Left, Right, Up, Down detector, a new energy histogram is filled, but only if the event occurs within our specified time-window. Figure 8 shows run 8545's Left E-detector's uncut spectra in gray, and Time-of-Flight cut spectra in blue. Note the reduction in number of events, from 255709 (as can be found on figures 1-7) to 221413 – we've cut away ~13% of all events.

Figure 8: Energy Spectra, Time-of-Flight Cut in Blue
When making Mott measurements with beam repitition rates of 249.5 MHz or 499 MHz, typical CEBAF repitition rates, the beam bunches are temporally spaced too close together to resolve a target scattering peak in the Time-of-Flight spectra. In this case, a flag can be passed to the analysis code to forgo the fit and subsequent cut.
Next, ToF-cut energy spectra (or simply “raw” energy spectra if no ToF-cuts are possible or wanted) are horizontally normalized such that their peaks each line up at a specified energy channel “center,” chosen to be 8000. This is implemented by calculating a 'squeeze fraction' equal to the center of the bin that the maximum count value occurs at divided by the channel to center on. Then, any bin edge or bin center can be calculated simply as:
NewBin = (MaxCountBinCenter / center) * OldBin
thereby squeezing/centering our four different detector energy spectra about a chosen channel.
Figure 9 shows unnormalized horizontally ToF-cut energy spectra from run 8545 in gray, and horizontally normalized ToF-cut energy spectra in blue. Figure 10 then shows the horizontally normalized, ToF-cut energy spectra of each of the four detectors atop one another.
Figure 9: Horizontally Normalizing, from gray to blue, ToF-Cut Energy Spectra

Figure 10: Horizontally Normalized ToF-Cut Energy Spectra
Next, each of the four detectors' horizontally normalized, potentially ToF-cut, energy spectra is fit with a Gaussian. The mean is given the “center” bin used to horizontally normalize about as its seed value, meanwhile the amplitude and sigma are each given a seed value heuristically determined to be 300. The fit is restricted to +/-500 channels about the “center” bin. Again, the default ROOT TH1 fitter is used. Figure 11 shows run 8545's Left detector energy spectra, ToF-cut, horizontally normalized about channel 8000, fit with a Gaussian in magenta. From the returned fit parameters, a “good” elastic scatterings off the target foil energy window is determined as from (mean – (½)*sigma) to (mean + 2*sigma). The choice of this -0.5 to +2 sigma about the mean energy window is explained in detail in section Energy Cuts.

Figure 11: Energy Spectra Horizontally Normalized, ToF-Cut, Fit with Gaussian
Note: Entries no longer = number of events in the statistics box, Entries = number of channels
Just prior to the end of the first loop subroutine, fit parameters and associated uncertainties from both fits along with 'squeeze fractions' are written to a formatted output file. The Time-of-Flight and Energy windows determined by the fits, windows that define our “good” Mott scatterings off the target foil, are passed to the second loop along with the calculated 'squeeze fractions' to horizontally normalize energy spectra. Figures 12 and 13 show run 8545's Left detector Energy vs Time-of-Flight plot with the determined Time-of-Flight window/cut shown by the vertical light red lines and the Energy window/cut shown by the horizontal magenta lines.

General Technique – Second Loop
In the second loop subroutine, from a given run's ROOT tree, eight new energy spectra histograms are filled – four detectors, and now breaking down scatterings by positive or negative helicity state. After pedestal subtraction, only scatterings that make it within our energy window determined in the first loop, and within our time-of-flight window if one is employed, are added. Using the 'squeeze fractions' passed from the first loop, these “good” elastic scatterings from the target foil are added directly to their horizontally normalized bin. These new, helicity-dependent energy spectra histograms are binned the same as they were in the first loop subroutine.