## Slide 1: 300 kV DC High Voltage Photogun with Inverted Insulator Geometry
Carlos Hernandez-Garcia
Jefferson Lab Accelerator Advisory Committee Mtg., September 13-15, 2017

## Slide 2: Inverted Insulator Photoguns
CEBAF 130 kV, used since 2010
350 kV gun for GTS and UITF
Longer “R30” insulator
Spherical electrode

## Slide 3: Building the 350 kV Gun
C. Hernandez-Garcia, M. Poelker, and J. Hansknecht, “High Voltage Studies of Inverted-geometry Ceramic Insulators for a 350kV dc Polarized Electron Gun,” IEEE Transactions on Dielectrics and Electrical Insulation, Vol. 23, No. 1; February 2016
Start with “dummy” electrode (no photocathode) and test different insulators and cathode screening electrodes

## Slide 4: 300 kV Inverted Gun and CsK2Sb Photocathode
Slide 4
4
8 MV/m
11 MV/m
First insulator conditioned to 325 kV, then rebuilt gun with optimized triple-junction shield
1mA
Learning how to use the new photocathode
Need mask and ion precipitator
Summer 2016

## Slide 5: Careful design of triple-junction-shield
Gradients shown for 350 kV
Simulationsby Y. Wang
Ceramic
HV plug
Design overall shape for low gradient here !!

## Slide 6: Rebuilt 300kV gun: new insulator and electrode, masked cathode and beamline ion precipitators
Active area
Non-magnetized beam current
Magnetized beam current
BEAM!
Cathode damage
Active area
4.5mA
3 mA
Notes: Interior Page Design 2

## Slide 7: Hand-Polishing Nightmare!
New CEBAF 200 kV Gun

## Slide 8: Centrifugal Barrel Polishing to Rescue
Months to hours !!!
  From machine shop
30 min. plastic cones
30 min. crushed corncob
C. Hernandez-Garcia, D. Bullard, F. Hannon, Y. Wang, and M. Poelker, “High voltage performance of a dc photoemission electron gun with centrifugal barrel-polished electrodes,” accepted Rev. Sci. Instrum, 22 August 2017
  Barrel polishing machine at JLab SRF

## Slide 9: Excellent Surface Finish
Diamond paste polishing Ra 30 nm*
* M. BastaniNejad, A. A. Elmustafa, E. Forman, S. Covert, J. Hansknecht, C. Hernandez-Garcia, M. Poelker, L. Das, M. Kelley, P. Williams, “Evaluation of electropolished stainless steel electrodes for use in DC high voltage photoelectron guns,” J. Vac. Sci. Technol. A, Vol. 33, No. 4, Jul/Aug 2015
similar to traditional hand polishing with sand paper and diamond grit

## Slide 10: The ultimate metric of success: High Voltage
Reached
360 kV in 70 hours
Vacuum and radiation levels indistinguishable from bkgd at 350kV
10 MV/m at 350 kV

## Slide 11: 3 photoguns with barrel polished electrodes
CEBAF 200 kV
Conditioned to 216 kV,
No FE at 200kV, ready for installation
UITF 350 kV
Under assembly
GTS 350 kV
In operation since Nov. 2016 with CsK2Sb photocathode:
3 mA magnetized beam
4.5 mA non-magnetized

## Slide 12: Summary
K2CsSb Photocathode Preparation Chamber, Gun, Solenoid and Beamline are all operational
Photogun operates reliably at 300 kV
Cathode solenoid can trigger field emission but we have learned how to prevent this
Have successfully magnetized electron beam and measured rotation angle
Preparing to install a modelocked drive laser, to generate mA magnetized beam with RF structure
Then switch to 32 mA 225 kV HV power supply….
Thanks to: P. Adderley, J. Benesch, B. Bullard, J. Grames, J. Guo, F. Hannon, J. Hansknecht, R. Kazimi, G. Krafft, M. A. Mamun, Y. Wang, S. Wijiethunga, J. Yoskovitz, S. Zhang
ODU Graduate Students, Postdoc

## Slide 13: Backup slides

## Slide 14: Face the Realities of HV Breakdown
The first test with R30 pure alumina ceramic suffered breakdown at ~315kV with subsequent puncture at 330kV.
A second test with a new R30 pure alumina suffered breakdown at ~315 kV.
* M. BastaniNejad, A. A. Elmustafa, E. Forman, J. Clark, S. Covert, J. Grames, J. Hansknecht, C. Hernandez-Garcia, M. Poelker and R. Suleiman, “Improving the performance of stainless-steel DC high voltage photoelectron gun cathode electrodes via gas conditioning with helium or krypton,” Nucl. Instr. and Meth. in Phys. Res. A, Vol. 762, pp. 135–141, 2014
Field emission managed via krypton-processing*, i.e., voltage applied with ~10-5 Torr krypton added to gun chamber

## Slide 15: Optimizing the Electrode Design
Pure alumina insulator replaced with doped alumina for higher voltage stand-off and robust operations
Electrode ‘shed’ (triple junction shield) re-designed by Yan Wang to lower gradient from 12 MV/m to 10 MV/m at 350kV
Pure alumina insulator
Doped alumina insulator

## Slide 16: Inverted-Insulator Photoguns
CEBAF 200 kV Gun
350 kV gun for GTS and UITF
with optimized triple point shields and mildly conductive insulators
Both designs, maximum field strength < 10 MV/m
Notes: Interior Page Design 2

## Slide 17: Barrel polished electrodes…
Gun electrodes traditionally polished using sand paper and diamond paste
Tedious process, takes about 30 days to complete
Now performed in 1 day using SRF barrel polishing machine
  From machine shop
30 minutes in plastic cones
30 minutes with corncob
C. Hernandez-Garcia, D. Bullard, F. Hannon, Y. Wang, and M. Poelker, “High voltage performance of a dc photoemission electron gun with centrifugal barrel-polished electrodes,” accepted Rev. Sci. Instrum, 22 August 2017