## Slide 1: CEBAF Vertical Wien Magnetic Field Map
Results of magnet mapping on November 16, 2020
Riad Suleiman

## Slide 2: Field Map, Background
Degaussed, Power Supply OFF

## Slide 3: Field Map, I=0.00 A
Hysteresis ±19.9 A

## Slide 4: Field Map, I=0.00 A
With Background Subtraction

## Slide 5: Field Map, No Background Subtraction

## Slide 6: Field Map, with Background Subtraction
With Background Subtraction

## Slide 7: Field Map, No Background Subtraction

## Slide 8: Field Map, with Background Subtraction
With Background Subtraction

## Slide 9: Field Map, 19.9 A, Repeat Measurements
With Background Subtraction

## Slide 10: Normalized Field Map, 19.9 A, Repeat Measurements
With Background Subtraction

## Slide 11: Field Map, 19.9 A, x Transverse

## Slide 12: Core (Center) Magnetic Field vs Current
With Background Subtraction

## Slide 13: Core Magnetic Field Residuals (B-p1*I)

## Slide 14: Hysteresis Loop

## Slide 15: Modified Hysteresis Loop
A second Modification:
subtract -0.017 A offset from Magnet Current

## Slide 16: Magnetic Field Effective Length
No Background Subtraction

## Slide 17: Magnetic Field Effective Length
With Background Subtraction
Vertical Wien Effective Length = 31.07±0.01 cm

## Slide 18: Vertical Wien Magnetic Field Map
Use B core (center) vs I measurement with these two modifications:
Subtract -0.017 A offset from Magnet Current:
Subtract background core field offset of -0.214 G
Use measured effective length to calculate Bdl:

## Slide 19: Summary
Vertical Wien field map background of -3.12 G cm is found by mapping degaussed magnet with power supply off
With background Bdl subtracted, Magnetic Effective Length of 31.07±0.01 cm is now independent of magnet current (Electric Effective Length is 31.20 cm)
These three measurements are important in understanding field maps:
  Magnet at I=0 and on loop
  Magnet degaussed with power supply off
  Magnet off stand but leaving everything else as is
Hysteresis loop study of Vertical Wien magnet showed that magnet current during mapping was off by -0.017 A – fix by subtracting this offset from map current
New Field map was generated

## Slide 21: How to calculate power supply current offset
How to calculate power supply current offset when examining field maps:
  Start with B (or Bdl) vs I data
  First, correct first background fields: B = B - bckg
  Fit with B=a1*I, or B = a0 + a1*I or higher order polynomial (you want the linear coefficient)
  Calculate hysteresis curve: h= B - a1*I
  Plot h vs I
  Calculate opposite hysteresis curve with: neg_h = -h and neg_I = -I
  Plot neg_h vs neg_I
  If end-points do not overlap then there is a current offset
  Guess an offset, re-calculate I as: I = I - CurrentOffset
  Re-calculate h as: h = B - a1*(I - CurrentOffset)
  Re-calculate neg_I as: neg_I = -I
  Re-calculate neg_h as: neg_h =-h
  Repeat for few values of current offset, till there is overlap at end-point currents