## Slide 1: Small group Mott meeting
4 February 2016

## Slide 2: Recent work
First, write up thickness measurement section, refine error bars on data very slightly to avoid double counting
Second, the small χ2 imply overestimated error bars. Do we really need the 2% foil uniformity and 5% batch uniformity specs?
Third, churn through the Run 2 data vs. thickness using original error bars.
Fourth, do we need to talk about the linear fit lopping off points at the end?

## Slide 3
 |  | Au_5385_B | Au_3057_C | Au_5134_B | Au_7028_B | Au_5275_C | Au_5613_D | Au_7029_B | Au_6809_B
 | nominal thickness (nm) | 1000 | 870 | 750 | 625 | 500 | 355 | 225 | 50
 | mean thickness (all data, nm) | 943.7 | 836.8 | 774.6 | 561.2 | 482.0 | 389.4 | 215.2 | 52.0
Stat. | Stdev, nom. identical data (nm) | 29.0 | 7.1 | 9.1 | 8.0 | 9.7 | 4.5 | 1.9 | 2.3
 | stdev image reanalysis (nm) | 22.5 | 7.7 | 9.4 | 7.5 | 4.0 | 2.7 | 1.8 | 2.1
Syst. | Image analysis: ± 4 Pixel | 20.0 | 8.0 | 10.0 | 8.0 | 8.0 | 8.0 | 2.6 | 2.6
 | Resolution (1.2 nm inherent) | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
 | Tilt (0.4%) | 4.6 | 4.2 | 3.9 | 2.8 | 2.5 | 1.9 | 1.1 | 0.3
 | Focus (1%) | 9.4 | 8.4 | 7.7 | 5.6 | 4.8 | 3.9 | 2.2 | 0.5
 | Different spots (Lebow: 2%) | 18.9 | 16.7 | 15.5 | 11.2 | 9.6 | 7.8 | 4.3 | 1.0
 | Sibling difference (Lebow:5%) | 47.2 | 41.8 | 38.7 | 28.1 | 24.1 | 19.5 | 10.8 | 2.6
Totals |  |  |  |  |  |  |  |  | 
 | stat uncertainty (nm) | 36.7 | 10.5 | 13.1 | 11.0 | 10.5 | 5.2 | 2.6 | 3.1
 | syst uncertainty (nm) | 55.6 | 46.7 | 43.8 | 31.9 | 27.7 | 22.9 | 12.2 | 4.1
 | total uncertainty (nm) | 66.6 | 47.9 | 45.7 | 33.7 | 29.6 | 23.5 | 12.5 | 5.1

## Slide 4: Pade(n,m) orders: Asy vs. Thick With both Lebow Uncertainties
Pade(n,m) | intercept | dA | R2 | red. χ2 | d.o.f. | Conf | Ftest
(1,0) | 43.8025 | 0.1169 | 0.991 | 1.28 | 9 | <50? | --
(2,0) | 44.0176 | 0.1018 | 0.997 | 0.594 | 8 | >70 | 11.45
(3,0) | 44.1777 | 0.128 | 0.997 | 0.546 | 7 | ~80 | 3.15 (rej F test)
(0,1) | 44.0382 | 0.0786 | 0.997 | .554 | 8 | >80 | 11.23
(0,2) | 44.0484 | 0.1057 | 0.997 | 0.737 | 7 | >60 | 0.022 (rej ftest)
(1,1) | 44.049 | 0.1061 | 0.997 | 0.737 | 7 | >60 | 9.67
(1,2) | 44.0295 | 0.0986 | 0.997 | 0.870 | 6 | >50 | 0.083 (rej. Ftest)
(2,1) | 45.0432 | 4.014 | 0.9977 | 0.6104 | 6 | >70 | 2.25 (rej. Ftest)

## Slide 5: Both 5% batch and 2% uniformity Lebow uncertainties not included

## Slide 6: Pade(n,m) orders: Asy vs. Thick, both systematic Lebow removed
Pade(n,m) | intercept | dA | R2 | red. χ2 | d.o.f. | Conf. | Ftest
(1,0) | 43.645 | .155 | 0.993 | 3.55 | 9 |  | --
(2,0) | 43.95 | .172 | 0.997 | 2.23 | 8 |  | 6.35
(3,0) | 44.1523 | .215 | 0.997 | 1.97 | 7 |  | 2.05
(0,1) | 44.022 | 0.127 | 0.997 | 2.16 | 7 |  | 6.79
(0,2) | 43.987 | 0.181 | .997 | 2.44 | 7 |  | .07
(1,1) | 43.9854 | 0.180 | .997 | 2.85 | 7 |  | 5.22
(1,2) | 43.9778 | 0.173 | .997 | 3.29 | 6 |  | .198
(2,1) | 49.1791 | 104 | .998 | 2.49 | 6 |  | 1.85

## Slide 7: Only 2% uniformity spec from Lebow included

## Slide 8: Pade(n,m) orders: Asy vs. Thick With both only 2% Lebow Uncertainty
Pade(n,m) | intercept | dA | R2 | red. χ2 | d.o.f. | Conf | Ftest
(1,0) | 43.68 | .147 | .993 | 2.87 | 9 |  | --
(2,0) | 43.97 | .147 | .997 | 1.56 | 8 |  | 8.60
(3,0) | 44.156 | .179 | .997 | 1.32 | 7 |  | 2.43
(0,1) | 44.027 | 0.110 | .997 | 1.48 | 8 |  | 9.47
(0,2) | 44.009 | .153 | .997 | 1.68 | 7 |  | 0.027
(1,1) | 44.008 | .153 | .997 | 1.96 | 7 |  | 7.15
(1,2) | 3e10 | 5e10 | .997 | 2.26 | 6 |  | .205
(2,1) | 46.21 | 17.2 | .998 | 1.64 | 6 |  | 2.18

## Slide 9: Run two data: Pade terms

## Slide 10: Run 1 Pade(n,m) orders: Asy vs. Thick
Pade(n,m) | intercept | dA | R2 | red. χ2 | d.o.f. | Ftest
(1,0) | 43.8025 | 0.1169 | 0.991 | 1.28 | 9 | --
(2,0) | 44.0176 | 0.1018 | 0.997 | 0.594 | 8 | 11.45
(0,1) | 44.0382 | 0.0786 | 0.997 | .554 | 8 | 11.23
(1,1) | 44.049 | 0.1061 | 0.997 | 0.737 | 7 | 9.67
Pade(n,m) | intercept | dA | R2 | red. χ2 | d.o.f. | Ftest
(1,0) | 43.8437 | 0.1475 | .984 | 2.28 | 9 | --
(2,0) | 44.1131 | 0.125 | .993 | 1.01 | 8 | 12.21
(0,1) | 44.0759 | 0.102 | .994 | 1.04 | 8 | 11.61
(1,1) | 44.1735 | 0.132 | .993 | 1.19 | 7 | 9.67
Run 2 Pade(n,m) orders: Asy vs. Thick

## Slide 11: Linear fit: how many points?
Points kept (of 11) | Asym | dA
5 | 43.9754 | 0.115
6 | 43.916 | 0.147
7 | 43.915 | 0.128
8 | 43.853 | 0.137
9 | 43.833 | 0.128
all | 43.803 | 0.117