## Slide 1: DAQ of UITF 200 keV Mott Polarimeter
New 200 keV Mott DAQ
Riad Suleiman

## Slide 2: Vacuum Chamber

## Slide 3: Measuring Mott Asymmetry
How to measure the Mott Asymmetry A?
  For one helicity state, measure the left and right detector counting rate, and
  Flip the electron polarization, measure the counting rate again, and
  Calculate the cross-ratio (r),
  Then, the Mott Asymmetry (A),
This cancels false asymmetries from detector efficiency, beam current, target thickness, and solid angle

## Slide 4: Measuring Instr. Asymmetry
How to measure the Instrumental Asymmetry A1?
  For one helicity state, measure the left and right detector counting rate, and
  Flip the electron polarization, measure the counting rate again, and
  Calculate the cross-ratio (x1),
  Then, the Instrumental Asymmetry (A1),
This measures asymmetries from detector efficiency and solid angle but cancels beam current and target thickness

## Slide 5: Measuring Instr. Asymmetry
How to measure the Instrumental Asymmetry A2?
  For one helicity state, measure the left and right detector counting rate, and
  Flip the electron polarization, measure the counting rate again, and
  Calculate the cross-ratio (x2),
  Then, the Instrumental Asymmetry (A2),
This measures asymmetries from beam current and target thickness but cancels detector efficiency and solid angle

## Slide 6: Sherman Function and Differential Cross Section, Au

## Slide 7: Sherman Function and Differential Cross Section, Ag
200 keV | S(130˚) | DCS (b/sr)
Au | -0.446 | 659
Ag | -0.206 | 128

## Slide 8: ORTEC Electronics
Have these modules:
  (1) ORTEC 710 Quad High Voltage Bias Supply (1 - 1000 V)
  (2) ORTEC 142A Preamplifier for detector input capacitance 0 to 100 pF (conversion gain 45 mV/MeV)
  (2) ORTEC 142B Preamplifier for detector input capacitance 100 to 400 pF (conversion gain 20 mV/MeV)
  (2) ORTEC Model 590A Amplifier and Timing Single-Channel Analyzer (SCA)
  (2) ORTEC Model 570 Amplifier

## Slide 9: ORTEC Detectors
(2) ORTEC ULTRA Detectors (BU-013-050-1000-S):
  Ion-Implanted Silicon Charged Particles Detectors
  Ultra-thin entrance window (500 Å) for optimum energy
  resolution (FWHM, 13 keV, 7 keV)
  B Mount
  Detector size of 50 mm2
  Depletion Depth (Range) of 1000 um for energies 500 keV
  Bias Voltage: +115 V
(2) ORTEC ULTRA Detectors (BU-012-050-100):
  Ion-Implanted Silicon Charged Particles Detectors
  Ultra-thin entrance window (500 Å) for optimum energy
  resolution (FWHM, 12 keV , 6 keV)
  B Mount
  Detector size of 50 mm2
  Depletion Depth (Range) of 100 um for energies 200 keV
  Bias Voltage: +50 V
Zero Transmission Rage vs. Energy for Electrons in Silicon

## Slide 10: Single-Channel Analyzer (SCA)
SCA Out:
  Standard NIM of +5 V, 500 ns wide if the amplified signal passes the threshold of the Lower Level discriminator
  Connect to Scaler
Amp Out:
  Energy output (Amplified: x5 to x1250)
  Connect to fADC

## Slide 11
Summer 2008 Test at CEBAF Injector
100 / 500 keV

## Slide 13
Left Detector
Right Detector
Left Preamp T Output
Right Preamp T Output
Note: the timing output is 50 ns wide signals with 5 ns rise time (picture pending)

## Slide 14: The old 100 keV DAQ

## Slide 15
100 keV, ITS Cave
100 keV, CEBAF Tunnel
500 keV, CEBAF Tunnel

## Slide 16: Asymmetry Data (100 keV, 300n A)
  Detectors HV +115 V
  Coarse Gain = 500, Fine Gain = 1.5
  Lower Level threshold = 1.25 V
  Window Width = 0.1 V for spectra data, = 0.5 for asymmetry data
  HWien = 90 degrees, MFB0D01A = -964 G-cm, MFB0D01B = 1013 G-cm
P = 31.3 % (Bulk GaAs)

## Slide 17: Asymmetry Data (500 keV, 300n A)
  Detectors HV +115 V
  Coarse Gain = 100, Fine Gain = 1.5
  Lower Level threshold = 1.4 V
  Window Width = 0.1 V for spectra data, = 0.3 for asymmetry data
  HWien = 90 degrees, MFB0D01A = -2369 G-cm, MFB0D01B = 2369 G-cm
P = 67.7 % (Super Lattice GaAs)

## Slide 18: Target Thickness Extrapolation
Single-Atom Sherman Function must be corrected for plural scattering (a few large angle scattering) in the target:
alpha = 0.0005/A for 100 keV electrons and 0.0001/A for 500 keV electrons
If possible, run with the thinnest target

## Slide 19: Target Thickness Scan at 100 keV

## Slide 20
Design of
UITF 200 keV Mott Polarimeter

## Slide 21: DAQ Schematic Diagram
BU-012-050-100
TI
Right (120˚)
Detector
Left (120˚)
Detector
142A
Preamp
VXS crate
SBC
E Out
570 Amp
Amp
Out
Fanout
T Out
OR
Integ.
fADC
SCA
Out
Mott Trigger
Counting
fADC
4 Helicity Signals
Helicity
Trigger
BU-013-050-1000-S
Right (50˚)
Detector
Left (50˚)
Detector
E Out
590A
Amp/Disc
142B
Preamp
4 Helicity Signals,
1 FC, 2 BPMs (X±, Y±)

## Slide 22: Helicity Signals
Pair Sync
Helicity Pattern
Reported or Delayed Helicity
T_Settle

## Slide 23: New DAQ for Mott Polarimeter
Measure Mott asymmetry (event counting)
Measure Charge asymmetry (per helicity)
Measure position differences (per helicity)
DAQ Triggers:
  Mott Detector
  Helicity
DAQ Hardware:
  VXS crate (Ordered)
  XVR-16 from Abaco Single Board Computer (SBC) (Ordered)
  Jefferson Lab Flash Analog-to-Digital Convertor (fADC250) (quantity = 2) (On-site)
  Trigger Interface (TI) (On-site)
  Front Panel Signal Distribution module (On-site)
  Desktop (Ready to Order)

## Slide 24: Installation Timeline
DAQ design and procurement: October – December
CODA and Firmware: January
Data decoding and analysis: January - February
DAQ tests: January – February
Ready for beam: March 1, 2021

## Slide 25: DAQ Readouts
Mott Trigger:
  Counting fADC
  Integration fADC
Helicity Trigger:
  Integration fADC

## Slide 26: Signals to DAQ
Signals to Counting fADC:
  (2) Amp Out
  (2) Timing output from preamplifier
  (4) Helicity Signals
Signals to Integration fADC (helicity-gated):
  (2) SCA Out
  (4) Helicity Signals
  (1) FC
  (8) 2 BPMs (X±, Y±)

## Slide 27: Target and Rates
Wien Vertical at 90˚
Gold Targets (Z=79): 100, 300, 500 Å
Silver Targets (Z=47): 300, 500, 2000 Å
Solid angle of 0.11 msr (hole with diameter of 0.02” at 1.6875”)
Backward scattering angle of 130 degrees:
Viewer and a thru hole
 | 200 keV
Au Target (A) | 100
Current (nA) | 50
Rate per detector(Hz) | 1200
Time (s) | 300

## Slide 28: Collimators
Electron Shielding: The front collimator is 1 mm thick Al. Do not change the Al Collimator.
x-ray Shielding: Add 1/8 inch lead jacket around the Al Collimator.
Electron Energy | Al Range (g/cm2) | Al Range (mm)
100 keV | 0.0075 | 0.03
500 keV | 0.095 | 0.35

## Slide 29: Cables, …
(4) Bias cable, SHV connector, 0 - 1000 V, male connectors
(2) E 142A preamp 93-ohm cable, RG62A/U BNC, male connectors
(2) T 142A preamp 50-ohm cable, RG58A/U BNC, male connectors
(2) Test 142A preamp 50-ohm cable, RG58A/U BNC, male connectors
(2) 142A preamp power cable, 9-pin D connectors (amphenol 17-10090), female in cave, male in service rack
(2) E 142B preamp 93-ohm cable, RG62A/U BNC, male connectors
(2) T 142B preamp 50-ohm cable, RG58A/U BNC, male connectors
(2) Test 142B preamp 50-ohm cable, RG58A/U BNC, male connectors
(2) 142B preamp power cable, 9-pin D connectors (amphenol 17-10090), female cave, male in service rack