## Slide 1: High Polarization Photocathodes: Operational Experience at Jefferson Lab Marcy L. Stutzman, P. Adderley, J. Clark, S. Covert, C.Hernandez-Garcia, J. Grames, J. Hansknecht, M. Poelker R. Suleiman
Notes: Title Page Design 3

## Slide 2: Thomas Jefferson National Accelerator Facility
12 GeV electron accelerator for Nuclear Physics (DOE)
Mission: To understand the transition from hadronic (protons and neutrons) to partonic (quarks and gluons) physics.
Hall A
Hall B
Hall C
Hall D
Source
Polarized source
Electron beam, with polarization nearly 90% and current to 200 μA
Investigate spin dependent nuclear processes
7/8 mile racetrack design accelerator
4 experimental halls

## Slide 3: Beam required at Jefferson Lab
24 Hours
High polarization for 3 (4) experimental halls simultaneously
24/7 beam operation
Beam Current: 200 µA
Charge Delivered: 16 Coulombs polarized electrons
Outline
Polarized source
Strained superlattice photocathode
Ion bombardment
Polarization issues
Surface charge limit
Lifetime enhancement
  Laser position, size
  Activation area
  High voltage
Future directions, ongoing research

## Slide 4: Photoemission Source
-130kV
light in
electrons out
photocathode
-130 kV DC (vs. RF) electrode bias
x-ray standard “inverted” insulator
Pumps with NEG modules and SEM style ion pump
Base pressure approaching XHV ≡ P < 1x10-10 Pa
  400°C heat treatment to reduce outgassing rate
  Commercial XHV gauge characterization to validate pressure measurement
Field emission reduced through polishing, electrode design
“synchronous photoinjection” = illuminate only when beam is needed
High polarization strained superlattice GaAs/GaAsP photocathode

## Slide 5: Photoemission from GaAs
Bare GaAs surface
Large Work Function
Positive Electron Affinity (PEA)
PEA=4.07 eV
Egap=1.4 eV
Conduction
Band
Valence
Band
GaAs
Vacuum
Dipole layers of Cs + NF3 or Oxygen
Reduces Work Function
Negative Electron Affinity (NEA)
+
-
-
χNEA ~ 0.2 eV
GaAs
Vacuum
Light
NEA Surface Activation
GaAs Spin Polarized Electron Source, D.T. Pierce, R.J. Celotta, G.-C. Wang, W.N. Unertl, A. Galejs, C.E. Kuyatt, and S.R. Mielczarek, Review of Scientific Instruments 51(4), 478 (1980).

## Slide 6: Polarized electrons using circularly polarized light
Crystalline (bulk) GaAs
Circularly polarized light
780 nm ~ band gap
Helicity reversal: 1kHz
Maximum theoretical polarization: 50%
GaAs Spin Polarized Electron Source, D.T. Pierce, R.J. Celotta, G.-C. Wang, W.N. Unertl, A. Galejs, C.E. Kuyatt, and S.R. Mielczarek, Review of Scientific Instruments 51(4), 478 (1980).

## Slide 7: Fiber-Based Drive Laser: telecom
Gain-switched seed
ErYb-doped fiber amplifier
Frequency-doubler
1560nm
780nm
Flip helicity of circularly polarized laser light
and electron polarization at 1 kHz

## Slide 8: High P: Breaking GaAs Degeneracy
Split degeneracy of P3/2 : Introduce strain on GaAs crystal by growing it on substrate (GaAsP) with different lattice constant
High polarization by circularly polarized laser excitation from
P3/2 to S1/2 : Egap < E < Egap+δ
Higher QE: Alternating layers of GaAs and GaAsP
Strained Superlattice GaAs
+3/2
Egap
δ=20–50
meV
mj = -1/2
+1/2
S1/2
P3/2
mj = -3/2
-1/2
+1/2
σ+
σ-
Bulk GaAs
Max 100% theoretical polarization

## Slide 9: Strained Superlattice QE and Polarization
+3/2
Egap
δ=20–50
meV
mj = -1/2
+1/2
S1/2
P3/2
mj = -3/2
-1/2
+1/2
σ+
σ-
Highest polarization at wavelength just over the band gap
Higher QE with shorter wavelengths
J.L. McCarter, … M.L. Stutzman et al., NIMA 618 (2010) 30.

## Slide 10: GaAs Photocathode Evolution
QE ~ 20%, 120 μA/mW
Unpolarized @ 532 nm
QE ~ 5%, 30 μA/mW
Pol ~ 35% @ 780 nm
Bulk GaAs
QE ~ 0.2%, 1 μA/mW
Pol ~ 75% @ 850 nm
Strained GaAs:
GaAs on GaAsP
100 nm
Superlattice GaAs:
Layers of GaAs on GaAsP
No strain relaxation
QE ~ 1%, 6 μA/mW
Pol ~ 85% @ 780 nm (1.59 eV)
100 nm
14 Layers
2 μm
625 μm
Timeline of Polarized Source at CEBAF
0.36
GaAs
GaAsP
GaAs
0.36
GaAs
GaAsP
GaAs
350 μm
350 μm

## Slide 11: High Polarization Photocathode
DOE Small Business Innovation Research (SBIR) partner (SLAC, JLab)
Grows high polarization photocathodes
GaAs/GaAsP, boron doped, gas phase MBE
Ongoing research for higher polarization or more robust photocathode materials

## Slide 12: Operational Lifetime Limitations
  Any gas in chamber can be ionized by electron beam, accelerated back toward the photocathode and
  limit photocathode operational lifetime
Static vacuum approaching 1x10-10 Pa
  Baked, leak free, all metal vacuum systems
    Residual water vapor destroys surface chemistry
  Hydrogen outgassing reduction
    Extended 400°C heat treatment of chamber
    Elimination of thick flanges
  Pump with NEG (non-evaporable getter) and ion pumps
Dynamic vacuum issues
  Eliminate field emission from electrodes
  Eliminate beam halo, scraping

## Slide 13
Ion Bombardment

## Slide 14: Photocathode Lifetime
Ion bombardment – with characteristic QE “trench” from laser spot to electrostatic center of photocathode – damages NEA of GaAs
High energy ions are focused to electrostatic center: create QE “hole”
  Don’t run beam from electrostatic center.
QE can be restored, but takes about 8 hours to heat and reactivate
Photocathode “QE scan”
  Active area = 5 mm
  Laser size = 0.35 mm
What causes photocathode damage?

## Slide 15: Off-axis implantation: SIMS analysis
Chart: Carbon
 | 1 | 10 | 100 | 1000 | 10000 | 100000
Chart: Oxygen
 | 1 | 100 | 10000 | 1000000 | 100000000
Chart: Fluorine
 | 1 | 10 | 100 | 1000 | 10000 | 100000
Chart: Hydrogen
 | 1 | 10 | 100 | 1000 | 10000 | 100000
Unused sample
Used sample
,

## Slide 16: Electrostatic center implantation: SIMS
Chart: Carbon
 | 1 | 10 | 100 | 1000 | 10000 | 100000
Chart: Oxygen
 | 1 | 100 | 10000 | 1000000 | 100000000
Chart: Fluorine
 | 1 | 10 | 100 | 1000 | 10000 | 100000
Chart: Hydrogen
 | 1 | 10 | 100 | 1000 | 10000 | 100000
V. Shutthanandan, …, M.L. Stutzman, et al., Surface science analysis of GaAs photocathodes following sustained electron beam delivery, PRSTAB 15 063501 (2012)

## Slide 17: Liftout FIB / TEM analysis
Superlattice layers
Focused Ion Beam used to mill cross sectional depth profile for TEM
Au/Pd evaporated, Pt sputtered to protect surface (misalignment = shadows)
Uniform layers, no darkening
Fred Stevie NCSU, former AVS M-A chair
F.A.Steve et al., Surf. Interface Anal. 31, 345, (2001).
Unused sample

## Slide 18: Liftout FIB / TEM analysis
Superlattice
layers
Used sample
Implantation damage evident
Dislocations – strain relaxation during heat cycles?
Why is implantation non-uniform?

## Slide 19: Liftout FIB / TEM analysis
Used sample
The residual gasses in the chamber are implanted
The implantation is non-uniform
Open questions
  What species are most harmful?
  Damage evident, but polarization and QE are fully restored with heat cycle and re-activation – why?

## Slide 20: Beam Polarization at CEBAF

## Slide 21: Variation in polarization: open questions
Why does heating/reacitivating the wafer cause polarization to decrease so significantly?
Why did polarization drop so significantly before the spot move?

## Slide 22: Polarization variation?
Low QE:
Run constant current. Low QE requires more laser power (~ 500mW).
Is the wafer heating, shifting the bandgap?
High QE:
Do we have depolarization in the Band Bending Region while electrons move through surface chemicals?

## Slide 23: Surface Charge Limit
Surface Charge Limit (or Surface Photovoltage Effect) affects cathode yield
Doubling of laser power doesn’t double electron beam current
Bulk GaAs,
532 nm,
100 kV
SCL high bunch charge at SLAC
Non-uniform emission during electron bunch
More pronounced at higher currents
JLab data
T.Maruyama et al., AIP Conf.Proc. 675 (2003).
Notes: SCL data:
https://logbooks/entry/1519971
https://logbooks/entry/1519969
https://logbooks/entry/1564645
Bulk GaAs Doping: Zn at 8 e+18 atoms/cm3

## Slide 24: Surface Charge Limit
Vacuum level increases with high current
Surface becomes less “Negative Electron Affinity”
Yield (quantum efficiency) drops
Egap
Mitigation to reduce SCL?
Heavily dope surface layer
Increase gun voltage
χNEA ~ 0.2 eV typical
Maruyama et al., Nucl. Instrum. Meth. A 492 (2002) 199.
Notes: SCL data:
https://logbooks/entry/1519971
https://logbooks/entry/1519969
https://logbooks/entry/1564645
Bulk GaAs Doping: Zn at 8 e+18 atoms/cm3

## Slide 25: Operational parameters to Prolong Photocathode Lifetime?
Laser spot position variation
Laser spot size variation
Activated area of photocathode
Gun voltage
“Charge and fluence lifetime measurements of a DC high voltage GaAs photogun at high average current.,” J. Grames, … M.L. Stutzman, et al., Phys. Rev. ST Accel. Beams 14, 043501 (2011)

## Slide 26: Laser spot position vs. electrostatic center
Use physical mask to limit area where Cs is deposited
Vary size of activated area
Vary laser position
5 mm
7 mm
11 mm
Conclusions
Smaller active area better
  Less stray beam
Lifetime bad at EC
Notes: What is the source of the anode current? Photo-emission from active area by diffusion of light.

## Slide 27: Lifetime vs. Laser spot size
Enhanced Charge Lifetime for QWeak: Increase laser size from 0.35 mm to 0.7 mm (diameter)
Notes: What is the new beam diameter at CEBAF? Joe is recalling 1.0 mm.
This POLOG entry:
https://logbooks/entry/1643854
says at least a factor of two.
I will keep using 0.7 mm.

## Slide 28
At 200 kV, only 60% of ions are created compared to 100 kV
Higher Gun Voltage to Improve Lifetime?
Issues at higher operational voltages
Better emittance, beam handling
Better lifetime???
Field emission
surface polishing of electrodes
noble gas processing
Insulator breakdown,
punch-through

## Slide 29: How to Prolong Photocathode Lifetime?
Larger Laser Size
  Distribute damage over larger area
  Less effect from surface charge limit
Laser Position on Photocathode and Active Area
  Run the laser farther from the electrostatic center
Higher Gun Voltage
  Ionization cross section for residual gas peaks around 150eV
  Higher gun voltage reduces volume in which max ionization occurs
  Better beam profile and less potential beam scraping
  Field emission and HV insulator problems
“Charge and fluence lifetime measurements of a DC high voltage GaAs photogun at high average current.,” J. Grames, … M.L. Stutzman, et al., Phys. Rev. ST Accel. Beams 14, 043501 (2011)

## Slide 30
Open issues
Better GaAs based photocathodes?
  Longer lifetime
  Higher polarization
Ongoing research
Jefferson Lab SRF institute well suited for SRF and GaAs characterization, similations
Other Photocathode materials?
  100% emission from half metals?
  Novel emission processes (surface plasmon polaritons)
  Can robust materials such as CsK2Sb emit polarized e-?
SBIR partnership with SVT Associates
Polarimeter improvments
Vacuum research toward XHV
Collaborations with university partners (William and Mary)
Multialkali research program at JLab (Md. A. Mamun)
Higher gun voltage
  Less Surface Charge Limit
  Better emittance
Insulator research and characterization
Gas processing and surface finish studies to reduce field emission (M.BastaniNejad, Md. Mamun, E. Forman)
Superconducting Radio Frequency (SRF) photoguns with GaAs photocathode?
  High brightness for light source
  High bunch charge
  High polarization

## Slide 31: Thank you for your attention
Acknowledgements
Jefferson Science Associates U.S. DOE Contract
No. DE-AC05-06OR23177
(L-R)Jim Clark, Matt Poelker, Steve Covert,
Phil Adderley, Riad Suleiman, John Hansknecht,
Marcy Stutzman, Joe Grames,
Not shown: Carlos Hernandez Garcia, Shukui Zhang, Don (Bubba) Bullard
Recent Students: M.BastaniNejad, Md. A. Mamun, E. Forman, J.L.McCarter, K. Surles-Law,

## Slide 32: backup

## Slide 33: Electric Field Enhancement of QE
NEA of GaAs depends on Gun HV. QE increases with external Electric Field at GaAs surface, Es,
Where χ (~200 meV) is the zero-field NEA value (Physics Letters A 282, 309) and potential barrier lowering due to Electric Field (Schottky Effect) is
Where εs (= 13.1) is GaAs relative permittivity.
Gun HV (kV) | Es (MV/m) | δU(Es) (meV)
100 | 2.0 | 50
200 | 4.0 | 70
χ
δU(Es)
Egap
Notes: Schottky Effect: Applied Physics Letters 23, 123, (1973).

## Slide 34: Space Charge Limit
Child’s Law (1D)
Child’s Law (2D) (PRL 87, 278301)
Short Bunch (PRL 98, 164802)
Maximum current density that can be transported across cathode-anode
gap is (for an infinite charge plane):
For electron emission from a finite circular spot on the cathode:
For CEBAF electron beam (499 MHz):
[A/cm2]

## Slide 35: NEA Activation of GaAs
“Activate” GaAs photocathode by applying about one mono-layer of Cesium and NF3 to very clean surface
QE
0
NF3
Cs
R | GaAs Light Reflection Coefficient (~ 0.3)
d | GaAs layer thickness (~ 0.1 μm)
α(Һω) | Photo-absorption Coefficient (~ 5x103 cm-1)
B(χ) | Surface Tunneling Probability(~ 0.2)
Time
GaAs
Cs
F
F
F
F
Cs
Cs
Cs