## Slide 1: UITF: T-cathode electrode, BPM data @200 [kV]
05/15/2019
Gabriel Palacios, Carlos Hernandez
Matt Poelker, Joe Grames

## Slide 2: Details
UITF gun
T-cathode electrode @200 [kV]
Bulk GaAs activated with with Cs and NF3
Data taken with BPM 201 (on top of corrector 101H&V) and BPM 203 (after corrector 101H&V)

## Slide 3: BPM linear response region
In the following plots the BPMs responses where mapped as a function of different 101 corrector H and V positions (increments of 1 [step], 3 [s] waiting time).
From -62 [Gauss cm] to 62 [Gauss cm] along the x-axis (With 101V corrector in 55 [Gauss cm]).
From 0[Gauss cm] to 142 [Gauss cm] along the y-axis (With 101H corrector OFF).
R
y
x

## Slide 4: BPM203 linear response region
After readjusting the BPMs and checking hardware to be working OK, we analyze the linear response region of BPM203 by scanning correctors MBHK201H along the x-axis, and the MBHK201V along the y-axis.
  We find that the BPM203 saturates outside the intervals:
  -6 [mm]<x< 6[mm]
  -6 [mm]<y< 6[mm]
  Therefore inside the intervals is the linear response region.

## Slide 5
Original

## Slide 6: BPM201 vs laser position
Assuming the BPM201 linear response region is the same as BPM203, we scan the laser position on the photocathode along the x and y-axis and plot the BPM201 response.

## Slide 7
Bona fide results inside the saturation interval:
-6 [mm]<BPM201 x< 6[mm]
-6 [mm]< BPM201 y< 6[mm]
Original

## Slide 8: Corrected positions according to S&A

## Slide 9: Corrected positions according to S&A
z
x
BPM201
BPM203
12.6 mm
3.6 mm
Photocathode
7.5 mm
BFS
Solenoid?
8.9 mm
7.6 mm
1.3 mm
US face?
DS face?

## Slide 10: Corrected positions according to S&A
z
y
Photocathode
BPM201
BPM203
0.6 mm
0.3mm
1.2mm
BFS
Solenoid?
0.3mm
US face?
DS face?

## Slide 11
Relative to photocathode
+5.1 mm
+0.9 mm

## Slide 12
+5.1 mm
+0.9 mm
As we move left and right from the photocathode center, the BPM201 “sees” the beam moving slightly in the same direction, up until the 3000 [step], where the x positions displace further away for each step.
As we move above and below the center, the BPM201 “sees” the beam moving in a similar way, but this time the vertical displacement increases as we go below the center. Moreover, most positions on the photocathode result in BPM201 negative vertical readings. The beam experiences a negative vertical kick (as predicted by our simulations!)
Relative to photocathode

## Slide 13: Centered in lense
MBHK101V, corrector centered in the first viewer (201)
  vs BPM201 x readback
  vs BPM201 y readback

## Slide 14
Check position of MBHK101

## Slide 15
Check position of MBHK101

## Slide 16: Active area
In the following color maps the laser was scanned on a square box of 3600X3600 [steps], centered in the photocathode (at position 9250,8750) as a function of different BPMs response (in increments of 300 [steps], 3 [s] waiting time).
This corresponds to a diameter of 5 [mm] (radius R=2.5 [mm]).
Then: L=2.5/√2≈1.77
So: 1.77 [mm]≈1770 [steps] (assuming 1000 [steps]≈1 [mm])
Then I used 1800 [steps] to each side from the origin.
R
L
L
y
x

## Slide 17: Active area
Relative to photocathode
Own coordinate systems

## Slide 18: Fin

## Slide 19: Extras

## Slide 20
Relative to photocathode
+5.1 mm
+0.9 mm
-3.9 mm
+1.8 mm

## Slide 21: BPM203 linear response region
  Then, relative to the photocathode, the linear response regions are
  -10 [mm]<x< 2[mm]
  -4 [mm]<y< 8[mm]