## Setup and Temperatures
Magnet ID: | DL 001
Location: | MDL0L02
Pole Gap: | 1.05"
Pole Length: | 4"
Meas. Position: | +Z is downstream of Z center
+X is towards the power terminal leads
Coil Temperature (Ϲ)
Time (min) | Top | Bottom
0 | 26 | 26
1 | 30 | 30
2 | 34 | 33
3 | 37 | 37
4 | 40 | 39
5 | 42 | 42
10 | 53 | 52
20 | 68 | 67
30 | 73 | 74
40 | 81 | 80
50 | 85 | 82
60 | 86 | 85
+Z
+X

## Center BdL Summary
X cl | 16.78 | cm
Z cl | 103.78 | cm
NMR Corr | Pos. Curr | Neg. Curr
| 100% | 99.99%
|  | 0 | 103 | 206 | 309 | 412 | 515 | 618 | 721 | 824 | 927 | 1030 | 1133 | 1236 | 1339 | 1442 | 1545 | 1648 | 1751 | 1854 | 1957 | 2060
Current (A)
| Z-Pos (cm) | 10.0137 | 9.0147 | 8.0132 | 7.0114 | 6.0102 | 5.0105 | 4.0102 | 3.0094 | 2.0091 | 1.0092 | 0.003348 | -0.99049 | -1.9895 | -2.9892 | -3.9938 | -4.9927 | -5.9903 | -6.9895 | -7.991 | -8.9956 | -9.9925
0.4995 | -30.01 | 0.80026 | 0.80026 | 0.80026 | 0.80026 | 0.80026 | 0.65021 | 0.70023 | 0.60019 | 0.60019 | 0.60019 | 0.50016 | 0.49996 | 0.44996 | 0.39997 | 0.34997 | 0.34997 | 0.29997 | 0.24998 | 0.24998 | 0.19998 | 0.19998
0.999 | -29.011 | 0.85028 | 0.85028 | 0.85028 | 0.80026 | 0.80026 | 0.70023 | 0.70023 | 0.65021 | 0.60019 | 0.55018 | 0.50016 | 0.49996 | 0.44996 | 0.39997 | 0.34997 | 0.29997 | 0.24998 | 0.19998 | 0.24998 | 0.14999 | 0.14999
0.999 | -28.012 | 0.90029 | 0.90029 | 0.90029 | 0.85028 | 0.85028 | 0.75024 | 0.75024 | 0.65021 | 0.60019 | 0.60019 | 0.50016 | 0.44996 | 0.44996 | 0.39997 | 0.34997 | 0.29997 | 0.19998 | 0.19998 | 0.14999 | 0.099992 | 0.049996
0.9995 | -27.013 | 1.0003 | 1.0003 | 0.95031 | 0.90029 | 0.90029 | 0.80026 | 0.80026 | 0.75024 | 0.65021 | 0.65021 | 0.50016 | 0.49996 | 0.39997 | 0.34997 | 0.29997 | 0.24998 | 0.19998 | 0.14999 | 0.099992 | 0.049996 | 0
1.0005 | -26.013 | 1.1004 | 1.0503 | 1.0503 | 1.0003 | 0.90029 | 0.85028 | 0.80026 | 0.75024 | 0.65021 | 0.60019 | 0.50016 | 0.44996 | 0.34997 | 0.29997 | 0.24998 | 0.19998 | 0.14999 | 0.049996 | 0.049996 | -0.049996 | -0.049996
1.0015 | -25.012 | 1.2004 | 1.2004 | 1.1504 | 1.1004 | 1.0003 | 0.95031 | 0.85028 | 0.80026 | 0.70023 | 0.65021 | 0.50016 | 0.44996 | 0.39997 | 0.24998 | 0.19998 | 0.14999 | 0.049996 | 0 | -0.049996 | -0.14999 | -0.19998
0.9985 | -24.01 | 1.3504 | 1.3004 | 1.2504 | 1.2004 | 1.1504 | 1.0003 | 0.95031 | 0.80026 | 0.70023 | 0.65021 | 0.55018 | 0.44996 | 0.34997 | 0.24998 | 0.19998 | 0.099992 | 0 | -0.099992 | -0.14999 | -0.29997 | -0.34997
0.9985 | -23.015 | 1.5005 | 1.4505 | 1.4505 | 1.3504 | 1.2504 | 1.1504 | 1.0003 | 0.90029 | 0.80026 | 0.65021 | 0.55018 | 0.44996 | 0.34997 | 0.19998 | 0.099992 | -0.049996 | -0.099992 | -0.24998 | -0.34997 | -0.44996 | -0.54995
1.002 | -22.013 | 1.7506 | 1.7006 | 1.6005 | 1.5005 | 1.4005 | 1.3004 | 1.1504 | 1.0003 | 0.85028 | 0.70023 | 0.55018 | 0.39997 | 0.24998 | 0.14999 | 0 | -0.099992 | -0.29997 | -0.44996 | -0.54995 | -0.69994 | -0.79993
1.0005 | -21.011 | 2.0006 | 2.0006 | 1.9006 | 1.7506 | 1.6005 | 1.4005 | 1.2504 | 1.1004 | 0.90029 | 0.70023 | 0.55018 | 0.39997 | 0.19998 | 0.049996 | -0.14999 | -0.29997 | -0.49996 | -0.64994 | -0.79993 | -0.94992 | -1.0999
0.999 | -20.012 | 2.5008 | 2.3508 | 2.2007 | 2.0507 | 1.8506 | 1.6505 | 1.4505 | 1.2504 | 1.0003 | 0.80026 | 0.55018 | 0.34997 | 0.14999 | -0.049996 | -0.29997 | -0.54995 | -0.74994 | -0.94992 | -1.0999 | -1.2999 | -1.4999
0.999 | -19.013 | 3.001 | 2.9009 | 2.6509 | 2.4508 | 2.2507 | 1.9506 | 1.7006 | 1.4505 | 1.1504 | 0.90029 | 0.60019 | 0.29997 | 0 | -0.24998 | -0.54995 | -0.84993 | -1.0499 | -1.3499 | -1.5999 | -1.8498 | -2.0498
1.0005 | -18.014 | 3.7512 | 3.6012 | 3.3511 | 3.051 | 2.7009 | 2.3508 | 2.0006 | 1.7006 | 1.3504 | 0.95031 | 0.60019 | 0.24998 | -0.14999 | -0.49996 | -0.84993 | -1.1999 | -1.5499 | -1.8998 | -2.1998 | -2.4998 | -2.8498
1.0015 | -17.012 | 4.8516 | 4.6015 | 4.2514 | 3.8512 | 3.4011 | 2.951 | 2.5008 | 2.0507 | 1.6005 | 1.1004 | 0.60019 | 0.099992 | -0.34997 | -0.84993 | -1.2499 | -1.7499 | -2.2498 | -2.6998 | -3.0997 | -3.5497 | -3.9497
0.999 | -16.011 | 6.4021 | 6.0019 | 5.5518 | 5.0016 | 4.4014 | 3.8012 | 3.201 | 2.5508 | 1.9006 | 1.3004 | 0.65021 | 0 | -0.64994 | -1.2499 | -1.9498 | -2.5498 | -3.1997 | -3.8497 | -4.3496 | -4.9496 | -5.5495
0.9985 | -15.014 | 8.7528 | 8.1526 | 7.4524 | 6.6522 | 5.8519 | 5.0016 | 4.1513 | 3.3011 | 2.4008 | 1.5505 | 0.70023 | -0.19998 | -1.0499 | -1.9998 | -2.8998 | -3.6997 | -4.5996 | -5.4495 | -6.2495 | -7.0494 | -7.8493
1.0005 | -14.014 | 12.254 | 11.3537 | 10.3534 | 9.203 | 8.1026 | 6.9022 | 5.6518 | 4.4514 | 3.201 | 2.0006 | 0.70023 | -0.54995 | -1.7998 | -3.0497 | -4.2996 | -5.4995 | -6.7494 | -7.9993 | -9.1492 | -10.2991 | -11.449
1.0015 | -13.013 | 17.9558 | 16.5554 | 14.9548 | 13.2543 | 11.6038 | 9.8532 | 8.0526 | 6.202 | 4.4514 | 2.6008 | 0.80026 | -1.0499 | -2.8998 | -4.6996 | -6.5494 | -8.3493 | -10.1991 | -11.999 | -13.6988 | -15.4487 | -17.1485
0.7505 | -12.011 | 27.4089 | 25.1582 | 22.6073 | 20.0065 | 17.4056 | 14.7048 | 11.9539 | 9.203 | 6.4021 | 3.6512 | 0.90029 | -1.9498 | -4.6996 | -7.4994 | -10.2991 | -13.0489 | -15.8487 | -18.6484 | -21.3482 | -23.998 | -26.6477
0.4985 | -11.512 | 34.4112 | 31.5102 | 28.3092 | 25.0581 | 21.7571 | 18.3059 | 14.8048 | 11.3537 | 7.9026 | 4.4014 | 0.90029 | -2.5998 | -6.0995 | -9.5992 | -13.0989 | -16.5986 | -20.0483 | -23.598 | -26.9977 | -30.3474 | -33.6472
0.4995 | -11.014 | 43.8642 | 40.013 | 35.9617 | 31.7603 | 27.4589 | 23.1575 | 18.7561 | 14.3046 | 9.9032 | 5.4518 | 1.0003 | -3.4497 | -7.9493 | -12.399 | -16.8486 | -21.2982 | -25.6978 | -30.1974 | -34.4971 | -38.8467 | -43.0964
0.5 | -10.513 | 56.7684 | 51.7668 | 46.415 | 40.9133 | 35.4115 | 29.7096 | 24.0078 | 18.3059 | 12.6041 | 6.8022 | 1.0503 | -4.6996 | -10.4991 | -16.2486 | -22.0481 | -27.7477 | -33.4472 | -39.2467 | -44.8962 | -50.4957 | -56.0453
0.5 | -10.014 | 74.5242 | 67.822 | 60.7697 | 53.5173 | 46.215 | 38.8126 | 31.3101 | 23.8077 | 16.2553 | 8.7028 | 1.2004 | -6.3995 | -13.9488 | -21.5482 | -29.0975 | -36.5969 | -44.0963 | -51.6456 | -59.095 | -66.5444 | -73.8437
0.5005 | -9.513 | 99.6323 | 90.5794 | 81.0763 | 71.3731 | 61.5199 | 51.6167 | 41.5635 | 31.5102 | 21.457 | 11.3537 | 1.3004 | -8.8493 | -18.8984 | -28.9975 | -39.0967 | -49.1458 | -59.145 | -69.2941 | -79.2433 | -89.1925 | -99.0416
0.499 | -9.013 | 135.4439 | 123.0399 | 109.9356 | 96.6813 | 83.277 | 69.7726 | 56.1682 | 42.5138 | 28.8093 | 15.1549 | 1.4505 | -12.299 | -25.9478 | -39.6466 | -53.2955 | -66.9443 | -80.5432 | -94.292 | -107.8409 | -121.4397 | -134.8886
0.5005 | -8.515 | 186.8606 | 169.3549 | 151.3491 | 132.9931 | 114.4871 | 95.8311 | 76.9749 | 58.1689 | 39.3127 | 20.5066 | 1.6505 | -17.1985 | -36.0469 | -54.8954 | -73.6938 | -92.5422 | -111.2906 | -130.289 | -148.7874 | -167.6858 | -186.1842
0.5005 | -8.012 | 261.5348 | 236.7267 | 211.4185 | 185.5601 | 159.6017 | 133.5433 | 107.1847 | 80.9262 | 54.6177 | 28.2592 | 1.9006 | -24.3979 | -50.7457 | -77.0435 | -103.4412 | -129.689 | -155.9868 | -182.2846 | -208.4824 | -234.8301 | -260.9279
0.4985 | -7.514 | 367.169 | 332.2577 | 296.3961 | 260.0843 | 223.5725 | 186.9606 | 150.0986 | 113.1867 | 76.2247 | 39.2627 | 2.2507 | -34.6471 | -71.5939 | -108.5408 | -145.5377 | -182.3346 | -219.1814 | -256.0783 | -292.8752 | -329.9721 | -366.669
0.501 | -7.015 | 514.4167 | 465.5009 | 414.8845 | 364.018 | 313.0515 | 261.3347 | 209.818 | 158.1012 | 106.2344 | 54.6177 | 2.7009 | -48.9459 | -100.6415 | -152.4371 | -204.2327 | -255.8283 | -307.274 | -359.2696 | -410.6152 | -462.6108 | -514.0065
0.5005 | -6.512 | 715.732 | 647.1597 | 576.887 | 505.8139 | 434.6909 | 363.1677 | 291.3944 | 219.4211 | 147.4978 | 75.5745 | 3.4011 | -68.3942 | -140.3381 | -212.332 | -284.4759 | -356.0199 | -427.7138 | -499.6077 | -571.3516 | -643.6955 | -714.9895
0.499 | -6.014 | 977.2667 | 882.9361 | 787.1051 | 689.9236 | 592.7421 | 495.4606 | 397.3788 | 299.197 | 201.0151 | 102.8333 | 4.3514 | -93.6421 | -191.9338 | -290.1754 | -388.6171 | -486.6588 | -584.3005 | -682.1923 | -780.234 | -878.7756 | -976.5673
0.5005 | -5.514 | 1289.4179 | 1164.5274 | 1038.4365 | 910.4951 | 782.0034 | 653.4118 | 524.1199 | 394.6779 | 265.0359 | 135.4939 | 5.5518 | -123.8395 | -253.4285 | -383.0676 | -513.1066 | -642.3456 | -771.3847 | -900.6238 | -1029.9128 | -1159.7518 | -1288.6909
0.4995 | -5.013 | 1577.8614 | 1425.412 | 1270.1116 | 1113.6109 | 956.41 | 798.9589 | 640.9077 | 482.6064 | 324.105 | 165.6037 | 6.6021 | -151.5872 | -310.1238 | -468.8603 | -628.0468 | -786.1835 | -944.1201 | -1102.0567 | -1259.9934 | -1418.5299 | -1575.2667
0.5005 | -4.515 | 1753.8684 | 1584.0134 | 1411.3574 | 1237.5011 | 1062.7444 | 887.7377 | 712.1808 | 536.2238 | 360.0667 | 184.0096 | 7.2524 | -168.5857 | -344.8208 | -521.1559 | -698.3409 | -874.226 | -1049.7611 | -1225.3963 | -1400.8314 | -1576.5166 | -1750.3018
0.5005 | -4.012 | 1827.9424 | 1650.835 | 1470.8767 | 1289.5179 | 1107.5089 | 925.0998 | 742.0905 | 558.6811 | 375.1716 | 191.6621 | 7.5024 | -175.7351 | -359.4196 | -543.254 | -727.9384 | -911.2729 | -1094.3074 | -1277.2919 | -1460.1264 | -1643.0109 | -1823.9456
0.4985 | -3.514 | 1852.9505 | 1673.2923 | 1490.8332 | 1307.0236 | 1122.5138 | 937.6539 | 752.1438 | 566.1835 | 380.1732 | 194.2129 | 7.5524 | -178.1849 | -364.3692 | -550.7034 | -737.8875 | -923.8718 | -1109.3561 | -1294.8904 | -1480.1247 | -1665.409 | -1848.7935
0.5005 | -3.015 | 1860.703 | 1680.3446 | 1497.0852 | 1312.5754 | 1127.1653 | 941.5551 | 755.2448 | 568.5343 | 381.7737 | 195.0132 | 7.5024 | -178.9849 | -365.969 | -553.1032 | -741.0873 | -927.8715 | -1114.0557 | -1300.3899 | -1486.4242 | -1672.4584 | -1856.5929
0.5 | -2.513 | 1863.1038 | 1682.4953 | 1498.9858 | 1314.1759 | 1128.5658 | 942.7055 | 756.1451 | 569.1845 | 382.2239 | 195.2633 | 7.4524 | -179.2348 | -366.419 | -553.8531 | -742.0372 | -929.0214 | -1115.5556 | -1302.0898 | -1488.324 | -1674.6083 | -1858.9427
0.5 | -2.015 | 1863.804 | 1683.0955 | 1499.536 | 1314.6761 | 1128.9659 | 943.0056 | 756.4452 | 569.3845 | 382.3239 | 195.3133 | 7.4524 | -179.3348 | -366.569 | -554.0531 | -742.3372 | -929.3713 | -1115.9555 | -1302.5897 | -1488.874 | -1675.2082 | -1859.6426
0.501 | -1.513 | 1864.0041 | 1683.1955 | 1499.686 | 1314.7761 | 1129.0659 | 943.1056 | 756.4952 | 569.4345 | 382.3739 | 195.3133 | 7.4524 | -179.3848 | -366.619 | -554.1031 | -742.3872 | -929.4713 | -1116.0555 | -1302.7397 | -1489.024 | -1675.3582 | -1859.8426
0.4985 | -1.013 | 1864.0041 | 1683.2455 | 1499.636 | 1314.7761 | 1129.0159 | 943.1056 | 756.4952 | 569.4345 | 382.3239 | 195.2633 | 7.4024 | -179.3848 | -366.619 | -554.0531 | -742.3872 | -929.5213 | -1116.1555 | -1302.7897 | -1489.074 | -1675.4082 | -1859.8926
0.5 | -0.516 | 1864.0041 | 1683.2455 | 1499.686 | 1314.7761 | 1129.0659 | 943.1056 | 756.4452 | 569.4345 | 382.3739 | 195.2133 | 7.4024 | -179.3848 | -366.619 | -554.0531 | -742.4372 | -929.5213 | -1116.1055 | -1302.7897 | -1489.124 | -1675.4082 | -1859.8926
0.501 | -0.013 | 1864.0041 | 1683.2455 | 1499.636 | 1314.7761 | 1129.0659 | 943.0556 | 756.3951 | 569.4345 | 382.3739 | 195.2133 | 7.4024 | -179.4348 | -366.619 | -554.1031 | -742.4372 | -929.5713 | -1116.1555 | -1302.7897 | -1489.124 | -1675.4582 | -1859.8926
0.4985 | 0.486 | 1864.0541 | 1683.2455 | 1499.636 | 1314.7761 | 1129.1159 | 943.1056 | 756.3951 | 569.4846 | 382.3739 | 195.1633 | 7.4024 | -179.3848 | -366.669 | -554.1031 | -742.4372 | -929.5713 | -1116.1555 | -1302.8397 | -1489.2239 | -1675.5082 | -1859.9426
0.5005 | 0.984 | 1864.0041 | 1683.2455 | 1499.636 | 1314.7261 | 1129.1159 | 943.1056 | 756.4452 | 569.4846 | 382.3739 | 195.2133 | 7.4024 | -179.3848 | -366.669 | -554.1531 | -742.4372 | -929.5713 | -1116.1555 | -1302.8897 | -1489.2239 | -1675.5582 | -1859.9926
0.5005 | 1.487 | 1863.9541 | 1683.1955 | 1499.636 | 1314.6761 | 1129.0659 | 943.1056 | 756.4452 | 569.4846 | 382.3739 | 195.2633 | 7.4024 | -179.3848 | -366.619 | -554.1531 | -742.4372 | -929.5213 | -1116.1555 | -1302.8397 | -1489.174 | -1675.5582 | -1859.9426
0.4995 | 1.985 | 1863.754 | 1682.9954 | 1499.4359 | 1314.476 | 1128.9659 | 943.0056 | 756.3951 | 569.4345 | 382.3239 | 195.2133 | 7.4524 | -179.3348 | -366.619 | -554.1031 | -742.3872 | -929.4213 | -1116.0055 | -1302.7397 | -1489.074 | -1675.3582 | -1859.7926
0.501 | 2.486 | 1863.0038 | 1682.3952 | 1498.9358 | 1314.0259 | 1128.5658 | 942.7055 | 756.1451 | 569.2345 | 382.2239 | 195.2133 | 7.4524 | -179.2848 | -366.469 | -553.9031 | -742.0872 | -929.0714 | -1115.6056 | -1302.2398 | -1488.524 | -1674.7082 | -1859.0926
0.499 | 2.987 | 1860.553 | 1680.0945 | 1496.8851 | 1312.3753 | 1127.1153 | 941.5051 | 755.1447 | 568.5343 | 381.7737 | 194.9132 | 7.5024 | -179.0348 | -365.969 | -553.1532 | -741.0873 | -927.8215 | -1114.1557 | -1300.4899 | -1486.5242 | -1672.5084 | -1856.5929
0.5 | 3.484 | 1852.4504 | 1672.8421 | 1490.383 | 1306.7235 | 1122.2637 | 937.4538 | 751.9937 | 566.1335 | 380.1232 | 194.1629 | 7.4524 | -178.2349 | -364.3192 | -550.6534 | -737.7876 | -923.7718 | -1109.2561 | -1294.8404 | -1480.0747 | -1665.1591 | -1848.3935
0.5005 | 3.987 | 1826.3419 | 1649.2845 | 1469.4762 | 1288.2675 | 1106.4586 | 924.2996 | 741.3903 | 558.1809 | 374.8215 | 191.4621 | 7.4024 | -175.6351 | -359.1696 | -542.904 | -727.4384 | -910.7229 | -1093.6074 | -1276.492 | -1459.0765 | -1641.4611 | -1821.9958
0.4985 | 4.485 | 1750.6174 | 1580.7623 | 1408.4064 | 1234.7502 | 1060.5437 | 885.9371 | 710.6803 | 535.0734 | 359.3165 | 183.5095 | 7.2023 | -168.2858 | -344.1209 | -520.306 | -697.091 | -872.7261 | -1048.0113 | -1223.3465 | -1398.2316 | -1572.7169 | -1745.5023
0.501 | 4.984 | 1571.7094 | 1419.36 | 1264.3598 | 1108.5593 | 952.1586 | 795.4578 | 638.1068 | 480.4557 | 322.6046 | 164.8534 | 6.5021 | -151.0372 | -308.9239 | -467.0105 | -625.697 | -783.4337 | -940.8204 | -1098.307 | -1255.0438 | -1411.0306 | -1565.8675
0.5005 | 5.487 | 1279.6147 | 1155.8246 | 1029.5837 | 902.7926 | 775.4513 | 647.8099 | 519.7685 | 391.2768 | 262.8352 | 134.2935 | 5.3517 | -122.8896 | -251.4287 | -380.2178 | -509.3569 | -637.746 | -765.8852 | -894.1743 | -1021.5635 | -1148.5528 | -1273.9922
0.499 | 5.985 | 970.2144 | 876.4841 | 780.9031 | 684.5719 | 588.1906 | 491.3592 | 394.2278 | 296.7462 | 199.4146 | 101.933 | 4.2014 | -93.0421 | -190.4339 | -288.0256 | -385.9673 | -483.3091 | -580.5509 | -677.7426 | -774.3345 | -870.2763 | -965.5683
0.5005 | 6.485 | 709.68 | 641.1578 | 571.4852 | 501.0624 | 430.4395 | 359.5665 | 288.4935 | 217.2204 | 145.9973 | 74.7742 | 3.2511 | -67.8443 | -139.0382 | -210.4322 | -281.9761 | -353.2201 | -424.3641 | -495.658 | -566.2521 | -636.4461 | -706.0402
0.499 | 6.986 | 510.9156 | 462.0998 | 411.7334 | 361.017 | 310.2505 | 259.0339 | 208.0174 | 156.6508 | 105.3341 | 54.0175 | 2.5508 | -48.5959 | -99.8415 | -151.2372 | -202.7828 | -254.0785 | -305.2242 | -356.7198 | -407.5655 | -458.0612 | -508.207
0.5005 | 7.483 | 365.3184 | 330.5571 | 294.6955 | 258.4838 | 222.172 | 185.5601 | 149.0483 | 112.3364 | 75.5745 | 38.9126 | 2.1507 | -34.5471 | -71.194 | -107.9909 | -144.7877 | -181.4846 | -218.0815 | -254.9284 | -291.3753 | -327.5223 | -363.4192
0.501 | 7.987 | 259.284 | 234.8261 | 209.4179 | 183.8096 | 158.1012 | 132.0428 | 106.1344 | 80.0259 | 53.9675 | 27.909 | 1.7506 | -24.2979 | -50.3457 | -76.4935 | -102.5913 | -128.7391 | -154.7369 | -181.0347 | -206.8825 | -232.5303 | -257.8782
0.498 | 8.485 | 185.6102 | 168.2045 | 150.1987 | 131.8427 | 113.4868 | 94.8307 | 76.2747 | 57.5687 | 38.9126 | 20.2065 | 1.5005 | -17.1985 | -35.897 | -54.5954 | -73.2938 | -91.9922 | -110.6906 | -129.539 | -148.0375 | -166.3859 | -184.4844
0.5 | 8.983 | 134.3936 | 121.9895 | 108.9853 | 95.781 | 82.4767 | 69.0224 | 55.568 | 42.0136 | 28.4092 | 14.9048 | 1.3504 | -12.249 | -25.7978 | -39.3967 | -52.9455 | -66.5944 | -80.0932 | -93.7421 | -107.1409 | -120.4398 | -133.5387
0.5 | 9.485 | 98.732 | 89.629 | 80.226 | 70.5729 | 60.8197 | 50.9165 | 41.0633 | 31.1101 | 21.1569 | 11.1536 | 1.2004 | -8.7993 | -18.7984 | -28.8476 | -38.7967 | -48.7959 | -58.795 | -68.8442 | -78.6433 | -88.3925 | -97.9917
0.4995 | 9.983 | 73.924 | 67.2218 | 60.2195 | 53.0672 | 45.8148 | 38.3624 | 30.96 | 23.5076 | 16.0052 | 8.5528 | 1.1004 | -6.4495 | -13.9488 | -21.4482 | -28.9475 | -36.4469 | -43.9463 | -51.4956 | -58.845 | -66.1444 | -73.2438
0.501 | 10.484 | 56.1682 | 51.1666 | 45.9149 | 40.4631 | 35.0113 | 29.3595 | 23.7077 | 18.0559 | 12.304 | 6.7022 | 1.0003 | -4.7496 | -10.4491 | -16.1986 | -21.8981 | -27.6477 | -33.2972 | -39.0467 | -44.5962 | -50.1458 | -55.5453
0.4995 | 10.985 | 43.314 | 39.6128 | 35.5615 | 31.4602 | 27.2088 | 22.8574 | 18.456 | 14.1546 | 9.7031 | 5.3517 | 0.90029 | -3.4997 | -7.9493 | -12.399 | -16.7986 | -21.1982 | -25.6478 | -30.0475 | -34.2971 | -38.5967 | -42.6964
0.5 | 11.483 | 33.961 | 31.1601 | 27.9591 | 24.758 | 21.507 | 18.1059 | 14.6047 | 11.2536 | 7.7525 | 4.3014 | 0.85028 | -2.6498 | -6.1495 | -9.6492 | -13.0989 | -16.5486 | -20.0483 | -23.498 | -26.8477 | -30.1474 | -33.3972
0.7495 | 11.985 | 27.0588 | 24.808 | 22.3572 | 19.8064 | 17.2056 | 14.5047 | 11.7538 | 9.0529 | 6.302 | 3.5512 | 0.80026 | -1.9998 | -4.7496 | -7.4994 | -10.2991 | -13.0489 | -15.7487 | -18.5484 | -21.1482 | -23.798 | -26.3478
0.999 | 12.982 | 17.7558 | 16.3553 | 14.7548 | 13.1543 | 11.5037 | 9.7031 | 7.9026 | 6.152 | 4.3514 | 2.5508 | 0.80026 | -1.0999 | -2.8998 | -4.7496 | -6.5494 | -8.3993 | -10.1491 | -11.949 | -13.6488 | -15.3487 | -16.9486
1.001 | 13.983 | 12.1539 | 11.2536 | 10.1533 | 9.153 | 7.9526 | 6.7522 | 5.5518 | 4.4014 | 3.151 | 1.9006 | 0.70023 | -0.59995 | -1.8498 | -3.0497 | -4.2996 | -5.5495 | -6.7494 | -7.9993 | -9.0992 | -10.2491 | -11.299
1.001 | 14.984 | 8.6028 | 8.0026 | 7.3024 | 6.6021 | 5.7519 | 4.9516 | 4.1013 | 3.2511 | 2.3508 | 1.5505 | 0.65021 | -0.29997 | -1.0999 | -1.9998 | -2.8498 | -3.7497 | -4.5996 | -5.4995 | -6.2495 | -7.0494 | -7.7993
0.999 | 15.985 | 6.302 | 5.9019 | 5.4518 | 4.9016 | 4.3514 | 3.7512 | 3.151 | 2.5008 | 1.8506 | 1.2504 | 0.60019 | -0.049996 | -0.69994 | -1.2999 | -1.9498 | -2.5998 | -3.1997 | -3.8497 | -4.3496 | -4.9496 | -5.4995
0.999 | 16.982 | 4.8016 | 4.5015 | 4.1513 | 3.8012 | 3.4011 | 2.9009 | 2.4508 | 2.0006 | 1.5005 | 1.0503 | 0.55018 | 0.099992 | -0.34997 | -0.84993 | -1.2999 | -1.7998 | -2.2498 | -2.6998 | -3.0997 | -3.5497 | -3.8997
1.0015 | 17.983 | 3.7012 | 3.5512 | 3.2511 | 3.051 | 2.6509 | 2.3508 | 2.0006 | 1.6505 | 1.3004 | 0.95031 | 0.55018 | 0.19998 | -0.14999 | -0.49996 | -0.89992 | -1.1999 | -1.5499 | -1.9498 | -2.1998 | -2.5498 | -2.7998
1.001 | 18.985 | 2.951 | 2.8009 | 2.6509 | 2.4508 | 2.2007 | 1.9506 | 1.7006 | 1.4005 | 1.1504 | 0.85028 | 0.55018 | 0.24998 | 0 | -0.29997 | -0.59995 | -0.84993 | -1.0499 | -1.3999 | -1.5999 | -1.8498 | -2.0498
0.999 | 19.985 | 2.4508 | 2.3007 | 2.1507 | 2.0507 | 1.8506 | 1.6505 | 1.5005 | 1.2504 | 1.0003 | 0.80026 | 0.50016 | 0.34997 | 0.099992 | -0.099992 | -0.34997 | -0.54995 | -0.74994 | -0.94992 | -1.0999 | -1.3499 | -1.4999
0.999 | 20.983 | 2.0006 | 1.9506 | 1.8506 | 1.7506 | 1.6005 | 1.4005 | 1.2504 | 1.1004 | 0.90029 | 0.75024 | 0.50016 | 0.34997 | 0.19998 | 0.049996 | -0.099992 | -0.34997 | -0.49996 | -0.64994 | -0.79993 | -0.94992 | -1.0499
1.0015 | 21.983 | 1.7506 | 1.6505 | 1.6005 | 1.5005 | 1.4005 | 1.2504 | 1.1004 | 1.0003 | 0.80026 | 0.70023 | 0.50016 | 0.39997 | 0.24998 | 0.099992 | -0.049996 | -0.19998 | -0.29997 | -0.44996 | -0.54995 | -0.69994 | -0.69994
1.001 | 22.986 | 1.5005 | 1.4505 | 1.4005 | 1.3504 | 1.2504 | 1.1004 | 1.0003 | 0.90029 | 0.75024 | 0.65021 | 0.50016 | 0.44996 | 0.29997 | 0.14999 | 0.049996 | -0.049996 | -0.19998 | -0.29997 | -0.39997 | -0.49996 | -0.49996
0.9985 | 23.985 | 1.3004 | 1.3004 | 1.2504 | 1.2004 | 1.1004 | 1.0003 | 0.90029 | 0.85028 | 0.70023 | 0.65021 | 0.50016 | 0.44996 | 0.34997 | 0.19998 | 0.14999 | 0 | -0.099992 | -0.14999 | -0.19998 | -0.29997 | -0.29997
0.9995 | 24.983 | 1.1504 | 1.2004 | 1.1504 | 1.1504 | 1.0003 | 0.95031 | 0.85028 | 0.75024 | 0.65021 | 0.60019 | 0.50016 | 0.44996 | 0.34997 | 0.29997 | 0.19998 | 0.099992 | 0 | -0.049996 | -0.099992 | -0.14999 | -0.19998
1.001 | 25.984 | 1.0503 | 1.0503 | 1.0503 | 1.0003 | 0.90029 | 0.85028 | 0.80026 | 0.75024 | 0.60019 | 0.55018 | 0.50016 | 0.39997 | 0.39997 | 0.29997 | 0.24998 | 0.14999 | 0.099992 | 0.049996 | 0 | -0.049996 | -0.049996
1.001 | 26.985 | 0.95031 | 0.95031 | 0.95031 | 0.90029 | 0.85028 | 0.80026 | 0.75024 | 0.65021 | 0.60019 | 0.55018 | 0.45015 | 0.44996 | 0.39997 | 0.29997 | 0.24998 | 0.19998 | 0.14999 | 0.099992 | 0 | 0 | 0
0.999 | 27.986 | 0.90029 | 0.85028 | 0.85028 | 0.85028 | 0.80026 | 0.75024 | 0.70023 | 0.70023 | 0.55018 | 0.55018 | 0.50016 | 0.44996 | 0.39997 | 0.34997 | 0.24998 | 0.24998 | 0.19998 | 0.14999 | 0.099992 | 0.099992 | 0.099992
0.9995 | 28.983 | 0.80026 | 0.80026 | 0.80026 | 0.80026 | 0.75024 | 0.70023 | 0.65021 | 0.60019 | 0.55018 | 0.55018 | 0.45015 | 0.44996 | 0.39997 | 0.34997 | 0.29997 | 0.24998 | 0.24998 | 0.19998 | 0.14999 | 0.14999 | 0.14999
0.501 | 29.985 | 0.80026 | 0.75024 | 0.80026 | 0.75024 | 0.70023 | 0.70023 | 0.60019 | 0.60019 | 0.55018 | 0.55018 | 0.50016 | 0.39997 | 0.39997 | 0.34997 | 0.29997 | 0.24998 | 0.24998 | 0.19998 | 0.19998 | 0.14999 | 0.19998
| Current (A) | 10.0137 | 9.0147 | 8.0132 | 7.0114 | 6.0102 | 5.0105 | 4.0102 | 3.0094 | 2.0091 | 1.0092 | 0.003348 | -0.99049 | -1.9895 | -2.9892 | -3.9938 | -4.9927 | -5.9903 | -6.9895 | -7.991 | -8.9956 | -9.9925
| BdL (G-cm) | 24038.0631 | 21720.5306 | 19360.3629 | 16980.3787 | 14589.7536 | 12191.9526 | 9785.8494 | 7374.5829 | 4960.8155 | 2548.3863 | 125.9789 | -2283.5543 | -4698.197 | -7116.0522 | -9542.7929 | -11954.6359 | -14360.6554 | -16769.5002 | -19168.9954 | -21569.5661 | -23944.1907
| Eff. L (cm) | 12.8959 | 12.904 | 12.91 | 12.915 | 12.922 | 12.9281 | 12.9375 | 12.9507 | 12.9737 | 13.0544 | 17.0187 | 12.7264 | 12.8149 | 12.8425 | 12.8533 | 12.8604 | 12.8662 | 12.872 | 12.8727 | 12.8738 | 12.874
Hall BdL
0 | -9.9925 | -23944.1907
-1 | -8.9956 | -21569.5661
-2 | -7.991 | -19168.9954
-3 | -6.9895 | -16769.5002
-4 | -5.9903 | -14360.6554
-5 | -4.9927 | -11954.6359
-6 | -3.9938 | -9542.7929
-7 | -2.9892 | -7116.0522
-8 | -1.9895 | -4698.197
-9 | -0.99049 | -2283.5543
-10 | 0.003348 | 125.9789
-11 | 1.0092 | 2548.3863
-12 | 2.0091 | 4960.8155
-13 | 3.0094 | 7374.5829
-14 | 4.0102 | 9785.8494
-15 | 5.0105 | 12191.9526
-16 | 6.0102 | 14589.7536
-17 | 7.0114 | 16980.3787
-18 | 8.0132 | 19360.3629
-19 | 9.0147 | 21720.5306
-20 | 10.0137 | 24038.0631

## Full Grid 5.5A Summary
X cl | X cl | 16.93 | cm
Z cl | Z cl | 102.18 | cm
NMR Corr | NMR Corr | 100.1%
|  |  | 0 | 85 | 170 | 255 | 340 | 425 | 510 | 595 | 680 | 765 | 850 | 935 | 1020 | 1105 | 1190 | 1275 | 1360 | 1445 | 1530 | 1615 | 1700
X-Position (cm)
|  | Z-Pos (cm) | -5 | -4.502 | -4.002 | -3.502 | -3.001 | -2.501 | -2.001 | -1.501 | -1.001 | -0.501 | -0.001 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.501 | 4.001 | 4.501 | 5.001
| 0.497 | -30.008 | 0.45028 | 0.40025 | 0.40025 | 0.40025 | 0.40025 | 0.40025 | 0.35021 | 0.40025 | 0.30018 | 0.30018 | 0.35021 | 0.30018 | 0.25015 | 0.30018 | 0.30018 | 0.30018 | 0.25015 | 0.30018 | 0.30018 | 0.30018 | 0.30018
| 0.999 | -29.014 | 0.50031 | 0.45028 | 0.40025 | 0.40025 | 0.45028 | 0.45028 | 0.40025 | 0.45028 | 0.40025 | 0.30018 | 0.35021 | 0.30018 | 0.30018 | 0.30018 | 0.30018 | 0.30018 | 0.30018 | 0.35021 | 0.35021 | 0.35021 | 0.35021
| 1.002 | -28.01 | 0.50031 | 0.50031 | 0.45028 | 0.45028 | 0.45028 | 0.50031 | 0.45028 | 0.45028 | 0.45028 | 0.40025 | 0.40025 | 0.35021 | 0.35021 | 0.35021 | 0.30018 | 0.35021 | 0.35021 | 0.35021 | 0.35021 | 0.40025 | 0.35021
| 0.998 | -27.01 | 0.55034 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.45028 | 0.50031 | 0.45028 | 0.40025 | 0.40025 | 0.40025 | 0.40025 | 0.30018 | 0.40025 | 0.40025 | 0.45028 | 0.35021 | 0.45028 | 0.40025
| 0.999 | -26.014 | 0.55034 | 0.55034 | 0.60037 | 0.55034 | 0.60037 | 0.55034 | 0.55034 | 0.55034 | 0.50031 | 0.50031 | 0.45028 | 0.50031 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.50031 | 0.45028 | 0.45028
| 1.0005 | -25.012 | 0.6504 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.55034 | 0.55034 | 0.55034 | 0.55034 | 0.55034 | 0.50031 | 0.45028 | 0.50031 | 0.50031 | 0.55034 | 0.55034 | 0.55034 | 0.60037
| 1 | -24.013 | 0.75046 | 0.70043 | 0.70043 | 0.70043 | 0.75046 | 0.70043 | 0.70043 | 0.70043 | 0.6504 | 0.6504 | 0.6504 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037
| 0.9995 | -23.012 | 0.80049 | 0.80049 | 0.75046 | 0.80049 | 0.80049 | 0.80049 | 0.80049 | 0.80049 | 0.80049 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.70043 | 0.75046 | 0.70043 | 0.75046 | 0.75046 | 0.75046
| 0.999 | -22.014 | 0.95058 | 0.90055 | 0.90055 | 0.90055 | 0.95058 | 0.95058 | 0.95058 | 0.95058 | 0.90055 | 0.90055 | 0.85052 | 0.85052 | 0.85052 | 0.80049 | 0.80049 | 0.80049 | 0.85052 | 0.80049 | 0.80049 | 0.80049 | 0.80049
| 1.0015 | -21.014 | 1.1007 | 1.1507 | 1.2007 | 1.2007 | 1.2007 | 1.1507 | 1.2007 | 1.2007 | 1.2007 | 1.2007 | 1.1507 | 1.2007 | 1.1507 | 1.1507 | 1.1507 | 1.0506 | 1.1007 | 1.1007 | 1.0506 | 1.0006 | 1.0006
| 1.001 | -20.011 | 1.3508 | 1.3508 | 1.4009 | 1.4009 | 1.4009 | 1.4509 | 1.4009 | 1.4509 | 1.4009 | 1.4009 | 1.3508 | 1.3508 | 1.3508 | 1.3008 | 1.2508 | 1.3008 | 1.3008 | 1.3008 | 1.3008 | 1.3008 | 1.2508
| 0.9985 | -19.012 | 1.601 | 1.651 | 1.701 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.701 | 1.701 | 1.651 | 1.601 | 1.601 | 1.601 | 1.601 | 1.5509 | 1.5509 | 1.5009
| 0.9995 | -18.014 | 1.9512 | 2.0012 | 2.0513 | 2.0513 | 2.1013 | 2.1513 | 2.1513 | 2.2013 | 2.1513 | 2.1513 | 2.1513 | 2.0513 | 2.0513 | 2.0012 | 2.0012 | 1.9512 | 1.9512 | 2.0012 | 1.9512 | 1.9012 | 1.9012
| 1.0015 | -17.013 | 2.5015 | 2.5516 | 2.6016 | 2.6516 | 2.6516 | 2.7017 | 2.7517 | 2.7517 | 2.7517 | 2.7017 | 2.7517 | 2.7017 | 2.7017 | 2.6016 | 2.5516 | 2.5516 | 2.5516 | 2.5516 | 2.4515 | 2.4515 | 2.3514
| 1.001 | -16.011 | 3.252 | 3.302 | 3.4021 | 3.5021 | 3.5021 | 3.5522 | 3.6022 | 3.6522 | 3.5522 | 3.6022 | 3.6022 | 3.5522 | 3.5522 | 3.5021 | 3.4521 | 3.4521 | 3.4021 | 3.3521 | 3.252 | 3.202 | 3.1519
| 0.998 | -15.011 | 4.4027 | 4.5028 | 4.5528 | 4.7029 | 4.8029 | 4.7529 | 4.903 | 4.903 | 4.903 | 4.903 | 4.903 | 4.8029 | 4.8029 | 4.8029 | 4.8029 | 4.7029 | 4.6028 | 4.5028 | 4.4527 | 4.3527 | 4.2026
| 0.999 | -14.015 | 6.0537 | 6.2038 | 6.4039 | 6.504 | 6.604 | 6.7041 | 6.7541 | 6.8042 | 6.8042 | 6.8042 | 6.8042 | 6.8042 | 6.7541 | 6.7541 | 6.604 | 6.554 | 6.454 | 6.3539 | 6.2038 | 6.0037 | 5.8036
| 1.0015 | -13.013 | 8.8054 | 9.0055 | 9.2056 | 9.5058 | 9.6059 | 9.806 | 9.806 | 9.856 | 9.9061 | 9.9061 | 9.9061 | 9.806 | 9.806 | 9.756 | 9.6059 | 9.5058 | 9.3557 | 9.2056 | 8.9555 | 8.7554 | 8.4552
| 0.749 | -12.012 | 13.3082 | 13.7584 | 14.0086 | 14.3588 | 14.5589 | 14.759 | 14.8091 | 14.9091 | 14.9592 | 14.9592 | 14.9592 | 14.9091 | 14.8591 | 14.659 | 14.5589 | 14.3588 | 14.1587 | 13.8585 | 13.5583 | 13.2581 | 12.8078
| 0.4995 | -11.515 | 16.7102 | 17.2606 | 17.6608 | 17.961 | 18.2112 | 18.4113 | 18.5614 | 18.6114 | 18.6614 | 18.6614 | 18.6614 | 18.6114 | 18.4613 | 18.3612 | 18.1611 | 17.961 | 17.7609 | 17.4607 | 17.1105 | 16.6102 | 16.0598
| 0.5015 | -11.013 | 21.3631 | 21.9635 | 22.4638 | 22.814 | 23.1642 | 23.4143 | 23.5144 | 23.6145 | 23.7145 | 23.7145 | 23.6145 | 23.5644 | 23.4143 | 23.2142 | 23.0641 | 22.7639 | 22.5638 | 22.1636 | 21.7633 | 21.213 | 20.4625
| 0.499 | -10.512 | 27.767 | 28.4174 | 28.9677 | 29.5181 | 29.8683 | 30.1685 | 30.3186 | 30.4186 | 30.4687 | 30.4687 | 30.4186 | 30.2685 | 30.1685 | 29.9684 | 29.7682 | 29.418 | 29.1178 | 28.6175 | 28.0172 | 27.4168 | 26.5162
| 0.5 | -10.015 | 36.5224 | 37.523 | 38.1734 | 38.7737 | 39.174 | 39.4742 | 39.6743 | 39.7744 | 39.8244 | 39.8244 | 39.7744 | 39.5742 | 39.3741 | 39.224 | 38.9238 | 38.5236 | 38.1734 | 37.573 | 36.8226 | 36.0221 | 34.8714
| 0.5 | -9.512 | 49.1801 | 50.3809 | 51.2814 | 51.8818 | 52.332 | 52.6322 | 52.8324 | 52.8824 | 52.9324 | 52.8824 | 52.7823 | 52.6322 | 52.4321 | 52.232 | 51.8317 | 51.4315 | 50.8311 | 50.1807 | 49.3302 | 48.2295 | 46.9287
| 0.4985 | -9.015 | 67.1911 | 68.592 | 69.5926 | 70.3931 | 70.8434 | 71.1936 | 71.3937 | 71.4938 | 71.4938 | 71.4438 | 71.2937 | 70.9935 | 70.6933 | 70.3431 | 69.8928 | 69.3925 | 68.7921 | 68.0417 | 66.891 | 65.6402 | 63.8891
| 0.501 | -8.515 | 93.4572 | 95.2583 | 96.309 | 97.2095 | 97.7098 | 98.01 | 98.2101 | 98.2101 | 98.01 | 97.91 | 97.8099 | 97.3096 | 97.0094 | 96.7092 | 96.2089 | 95.6086 | 94.8081 | 93.8575 | 92.5567 | 90.9057 | 88.7043
| 0.5 | -8.013 | 132.0309 | 133.9821 | 135.4329 | 136.0833 | 136.5336 | 136.8338 | 136.8838 | 136.9839 | 136.7838 | 136.4836 | 136.2835 | 135.7331 | 135.3329 | 134.8326 | 134.2322 | 133.3817 | 132.4811 | 131.3304 | 129.6294 | 127.8783 | 125.0766
| 0.5 | -7.515 | 187.2647 | 189.416 | 191.067 | 191.6174 | 191.8675 | 192.1177 | 191.9676 | 191.9175 | 191.7174 | 191.4673 | 190.8669 | 190.1164 | 189.6662 | 188.6655 | 187.9651 | 187.0646 | 185.9639 | 184.563 | 182.912 | 180.7107 | 177.5587
| 0.5005 | -7.013 | 265.9129 | 268.2143 | 269.5651 | 270.4156 | 270.2155 | 269.9653 | 270.0154 | 269.7152 | 268.9647 | 268.4144 | 267.714 | 266.7634 | 265.9129 | 265.1624 | 264.1618 | 262.961 | 262.0105 | 260.4095 | 258.3082 | 255.7566 | 251.9543
| 0.4985 | -6.514 | 373.5788 | 375.48 | 376.4806 | 377.0809 | 376.7307 | 376.2804 | 375.4299 | 375.1297 | 374.3793 | 373.0285 | 372.0779 | 371.4275 | 369.9766 | 368.976 | 367.7752 | 366.5745 | 364.7234 | 363.0223 | 360.6709 | 357.9692 | 353.8667
| 0.5005 | -6.016 | 516.2161 | 517.5169 | 518.0673 | 517.567 | 517.1167 | 515.916 | 514.8653 | 514.2649 | 513.2643 | 512.2137 | 510.7628 | 509.1118 | 507.9611 | 506.5102 | 505.1093 | 503.5084 | 501.4571 | 499.5559 | 497.2045 | 494.4528 | 490.5004
| 0.5005 | -5.513 | 690.673 | 690.9732 | 690.623 | 689.7724 | 688.9719 | 687.8212 | 686.7206 | 685.5699 | 684.0689 | 682.8682 | 680.817 | 679.3661 | 677.565 | 676.064 | 674.7132 | 673.0122 | 671.3111 | 669.3099 | 667.1586 | 664.4069 | 660.3044
| 0.4995 | -5.015 | 856.3745 | 856.2744 | 855.7241 | 855.1737 | 854.2732 | 853.1725 | 852.0718 | 851.0212 | 849.8205 | 848.5197 | 847.0688 | 846.0682 | 844.6673 | 843.0163 | 842.1157 | 840.7149 | 839.214 | 837.8631 | 835.6618 | 833.9607 | 830.8588
| 0.501 | -4.514 | 964.4406 | 964.3906 | 964.2405 | 963.8403 | 963.1399 | 962.5895 | 961.9391 | 961.3888 | 960.7384 | 959.9379 | 959.3876 | 958.6871 | 958.0367 | 957.2362 | 956.5858 | 955.8854 | 955.0349 | 954.3345 | 953.3338 | 952.1831 | 950.7823
| 0.4995 | -4.013 | 1012.0198 | 1012.1198 | 1011.9197 | 1011.8197 | 1011.4695 | 1011.2193 | 1011.0192 | 1010.719 | 1010.4188 | 1010.2187 | 1009.8184 | 1009.4682 | 1009.2181 | 1008.9179 | 1008.6677 | 1008.3175 | 1008.1674 | 1007.8672 | 1007.417 | 1006.9167 | 1006.2162
| 0.501 | -3.515 | 1028.3798 | 1028.4799 | 1028.3798 | 1028.3798 | 1028.2797 | 1028.1797 | 1028.1297 | 1027.9796 | 1027.8295 | 1027.7794 | 1027.5293 | 1027.4292 | 1027.3292 | 1027.1791 | 1027.129 | 1027.029 | 1026.9789 | 1026.8289 | 1026.7288 | 1026.4286 | 1026.0784
| 0.501 | -3.011 | 1033.583 | 1033.7831 | 1033.633 | 1033.6831 | 1033.583 | 1033.583 | 1033.4829 | 1033.3829 | 1033.3829 | 1033.2828 | 1033.2328 | 1033.1327 | 1033.1327 | 1033.0327 | 1032.9826 | 1033.0327 | 1033.0327 | 1032.9826 | 1032.9326 | 1032.7825 | 1032.5824
| 0.4985 | -2.513 | 1035.234 | 1035.3841 | 1035.234 | 1035.284 | 1035.234 | 1035.184 | 1035.1339 | 1035.0839 | 1035.0339 | 1034.9839 | 1034.9338 | 1034.8838 | 1034.8838 | 1034.8338 | 1034.8338 | 1034.8838 | 1034.8838 | 1034.8338 | 1034.8338 | 1034.6837 | 1034.4835
| 0.5005 | -2.014 | 1035.7843 | 1035.8344 | 1035.7843 | 1035.7843 | 1035.6843 | 1035.6342 | 1035.5342 | 1035.5842 | 1035.4842 | 1035.4341 | 1035.4341 | 1035.3841 | 1035.3841 | 1035.3341 | 1035.3341 | 1035.3341 | 1035.3841 | 1035.3841 | 1035.3841 | 1035.234 | 1035.1339
| 0.5 | -1.512 | 1035.8844 | 1035.9344 | 1035.9344 | 1035.8844 | 1035.8344 | 1035.8344 | 1035.6843 | 1035.6342 | 1035.6342 | 1035.5842 | 1035.5342 | 1035.4842 | 1035.4842 | 1035.4842 | 1035.4341 | 1035.4842 | 1035.5342 | 1035.5342 | 1035.5342 | 1035.4341 | 1035.3341
| 0.4995 | -1.014 | 1035.8844 | 1035.9845 | 1035.9845 | 1035.9344 | 1035.8844 | 1035.8344 | 1035.7843 | 1035.7843 | 1035.6342 | 1035.6342 | 1035.5342 | 1035.4842 | 1035.4842 | 1035.4341 | 1035.4842 | 1035.4842 | 1035.5842 | 1035.5342 | 1035.5842 | 1035.5342 | 1035.4341
| 0.501 | -0.513 | 1035.8844 | 1036.0345 | 1036.0345 | 1035.8844 | 1035.8844 | 1035.8344 | 1035.8344 | 1035.7343 | 1035.6342 | 1035.6342 | 1035.5342 | 1035.5342 | 1035.5342 | 1035.4842 | 1035.4842 | 1035.5342 | 1035.6342 | 1035.5342 | 1035.6342 | 1035.5842 | 1035.4341
| 0.499 | -0.012 | 1035.9344 | 1036.0345 | 1035.9845 | 1035.8844 | 1035.9344 | 1035.7843 | 1035.7843 | 1035.8344 | 1035.6342 | 1035.5842 | 1035.5842 | 1035.5842 | 1035.5342 | 1035.4842 | 1035.4842 | 1035.5842 | 1035.6342 | 1035.5842 | 1035.6843 | 1035.6342 | 1035.4842
| 0.4995 | 0.485 | 1035.9344 | 1035.9845 | 1036.0345 | 1035.9344 | 1035.9344 | 1035.8344 | 1035.8344 | 1035.7843 | 1035.6843 | 1035.6342 | 1035.5842 | 1035.5842 | 1035.5342 | 1035.5342 | 1035.5342 | 1035.6342 | 1035.6342 | 1035.5842 | 1035.7843 | 1035.6342 | 1035.4341
| 0.5005 | 0.987 | 1035.8844 | 1036.0345 | 1035.9845 | 1035.9344 | 1035.9344 | 1035.8344 | 1035.8844 | 1035.8344 | 1035.6843 | 1035.6342 | 1035.5842 | 1035.5842 | 1035.5842 | 1035.5842 | 1035.5342 | 1035.5842 | 1035.6342 | 1035.6342 | 1035.7843 | 1035.6342 | 1035.4842
| 0.4995 | 1.486 | 1035.8844 | 1035.9344 | 1035.9344 | 1035.9344 | 1035.8844 | 1035.8844 | 1035.8344 | 1035.7843 | 1035.7843 | 1035.6342 | 1035.6342 | 1035.6342 | 1035.5342 | 1035.5842 | 1035.5342 | 1035.5842 | 1035.6342 | 1035.6342 | 1035.7843 | 1035.6342 | 1035.4842
| 0.501 | 1.986 | 1035.7843 | 1035.8344 | 1035.8344 | 1035.8344 | 1035.8344 | 1035.7843 | 1035.7843 | 1035.7343 | 1035.5842 | 1035.5842 | 1035.5842 | 1035.5342 | 1035.4842 | 1035.4842 | 1035.4842 | 1035.5342 | 1035.5342 | 1035.5842 | 1035.6342 | 1035.5842 | 1035.4341
| 0.4995 | 2.488 | 1035.4341 | 1035.4842 | 1035.4842 | 1035.4842 | 1035.4842 | 1035.3841 | 1035.4341 | 1035.3841 | 1035.3841 | 1035.3841 | 1035.3341 | 1035.284 | 1035.234 | 1035.234 | 1035.234 | 1035.284 | 1035.3341 | 1035.3841 | 1035.3841 | 1035.284 | 1035.1339
| 0.499 | 2.985 | 1034.1834 | 1034.2334 | 1034.3835 | 1034.3835 | 1034.3835 | 1034.3835 | 1034.4335 | 1034.2834 | 1034.2834 | 1034.2334 | 1034.2334 | 1034.2334 | 1034.1834 | 1034.2334 | 1034.1834 | 1034.2834 | 1034.2834 | 1034.3835 | 1034.3835 | 1034.2834 | 1034.1333
| 0.5005 | 3.486 | 1030.1309 | 1030.281 | 1030.381 | 1030.4811 | 1030.4811 | 1030.5311 | 1030.5812 | 1030.5812 | 1030.6312 | 1030.5812 | 1030.6312 | 1030.6812 | 1030.7813 | 1030.8313 | 1030.7813 | 1030.8813 | 1030.9314 | 1030.9814 | 1031.0815 | 1030.9814 | 1030.7813
| 0.499 | 3.986 | 1017.1729 | 1017.5732 | 1017.9234 | 1018.1235 | 1018.2236 | 1018.4237 | 1018.6238 | 1018.6238 | 1018.824 | 1018.874 | 1019.0241 | 1019.1742 | 1019.3243 | 1019.4744 | 1019.5744 | 1019.8246 | 1020.0247 | 1020.1748 | 1020.3249 | 1020.3249 | 1020.0747
| 0.5005 | 4.484 | 978.9996 | 979.9001 | 980.5005 | 981.0508 | 981.6012 | 982.2015 | 982.6518 | 982.952 | 983.3522 | 983.8025 | 984.3528 | 984.7031 | 985.2034 | 985.7537 | 986.0039 | 986.4541 | 986.9544 | 987.3547 | 987.855 | 987.905 | 987.855
| 0.501 | 4.987 | 885.6424 | 887.4435 | 888.8444 | 890.0451 | 891.3459 | 892.4466 | 893.3971 | 894.2977 | 895.2983 | 896.1989 | 897.2495 | 898.4002 | 899.3508 | 900.2513 | 901.252 | 902.3526 | 903.3032 | 904.0537 | 905.1043 | 905.4045 | 905.4045
| 0.4985 | 5.486 | 729.4968 | 732.2985 | 734.6499 | 736.2509 | 738.052 | 739.653 | 741.1539 | 742.6048 | 743.7555 | 744.9562 | 746.6072 | 748.2583 | 749.6091 | 750.8598 | 752.1606 | 753.7116 | 755.0124 | 756.1631 | 757.3638 | 757.9142 | 757.3638
| 0.4995 | 5.984 | 555.09 | 558.492 | 560.9435 | 562.9448 | 564.8459 | 566.3969 | 568.0479 | 569.0485 | 570.4994 | 571.7502 | 573.201 | 574.752 | 576.0528 | 577.5037 | 578.6544 | 579.8051 | 580.9058 | 581.9064 | 582.6568 | 582.2066 | 581.256
| 0.5005 | 6.485 | 404.5978 | 408.2 | 410.8516 | 412.6527 | 414.5039 | 415.7046 | 417.3056 | 418.2562 | 419.1567 | 420.2073 | 421.4081 | 422.6088 | 423.4593 | 424.51 | 425.4105 | 426.0109 | 426.8614 | 427.3617 | 427.4118 | 426.5612 | 424.51
| 0.4995 | 6.985 | 289.9276 | 293.3797 | 295.9813 | 297.5822 | 299.3834 | 300.2839 | 301.5347 | 302.4352 | 303.1356 | 304.0362 | 304.7867 | 305.6872 | 306.2876 | 306.988 | 307.4383 | 307.9386 | 308.1887 | 308.1387 | 307.7885 | 306.4877 | 303.9862
| 0.5005 | 7.484 | 205.3758 | 208.3276 | 210.679 | 212.33 | 213.7309 | 214.6815 | 215.6321 | 216.3325 | 216.8828 | 217.5332 | 218.1836 | 218.734 | 219.2343 | 219.6345 | 219.8847 | 220.0348 | 220.0848 | 219.7846 | 219.0341 | 217.6333 | 215.0317
| 0.5 | 7.986 | 144.6386 | 147.2402 | 149.0913 | 150.4922 | 151.793 | 152.4934 | 153.3439 | 153.8942 | 154.4446 | 154.8949 | 155.2951 | 155.6954 | 155.9455 | 156.2457 | 156.2957 | 156.3458 | 156.1957 | 155.6954 | 154.8949 | 153.444 | 151.1426
| 0.4995 | 8.484 | 102.4127 | 104.514 | 106.215 | 107.3658 | 108.2663 | 109.0168 | 109.7172 | 110.2175 | 110.5677 | 110.9179 | 111.3182 | 111.5683 | 111.6684 | 111.8185 | 111.8685 | 111.7684 | 111.4683 | 110.8179 | 110.1174 | 108.6165 | 106.5653
| 0.5005 | 8.985 | 73.2949 | 74.9959 | 76.3468 | 77.3474 | 78.0478 | 78.7983 | 79.2485 | 79.5987 | 79.8489 | 80.1491 | 80.2992 | 80.5493 | 80.5994 | 80.5994 | 80.5493 | 80.3992 | 80.1491 | 79.5987 | 78.8483 | 77.6476 | 75.9465
| 0.499 | 9.485 | 53.2326 | 54.5834 | 55.6841 | 56.4846 | 57.2351 | 57.7354 | 58.1856 | 58.4358 | 58.6359 | 58.836 | 58.9861 | 59.1362 | 59.1862 | 59.0862 | 58.9361 | 58.786 | 58.5358 | 57.9855 | 57.3351 | 56.3345 | 54.9837
| 0.5005 | 9.983 | 39.4241 | 40.5248 | 41.4754 | 42.0758 | 42.6261 | 43.0764 | 43.4266 | 43.6767 | 43.8268 | 43.9769 | 44.1771 | 44.1771 | 44.2271 | 44.1771 | 43.9769 | 43.7768 | 43.5267 | 43.0264 | 42.476 | 41.6255 | 40.5248
| 0.501 | 10.486 | 29.5681 | 30.4687 | 31.2191 | 31.7695 | 32.2698 | 32.5699 | 32.8701 | 33.1203 | 33.2203 | 33.3704 | 33.4205 | 33.4205 | 33.5205 | 33.3704 | 33.2203 | 33.0703 | 32.7701 | 32.4199 | 31.9195 | 31.2692 | 30.3186
| 0.4985 | 10.985 | 22.7139 | 23.3643 | 23.9647 | 24.465 | 24.8152 | 25.1654 | 25.4156 | 25.5156 | 25.6157 | 25.7658 | 25.8158 | 25.7658 | 25.8158 | 25.7658 | 25.5657 | 25.4156 | 25.2655 | 24.8152 | 24.465 | 23.8146 | 23.1642
| 0.5 | 11.483 | 17.8109 | 18.2612 | 18.7615 | 19.1617 | 19.4619 | 19.7121 | 19.9122 | 20.0123 | 20.1623 | 20.2124 | 20.2124 | 20.1623 | 20.2124 | 20.1623 | 20.0123 | 19.8622 | 19.7121 | 19.4119 | 19.0617 | 18.6114 | 18.011
| 0.751 | 11.985 | 14.0086 | 14.4589 | 14.8091 | 15.2093 | 15.4595 | 15.6096 | 15.7597 | 15.8597 | 15.9598 | 16.0098 | 16.1099 | 16.0098 | 16.0098 | 15.9598 | 15.8097 | 15.7597 | 15.5595 | 15.3094 | 15.0092 | 14.659 | 14.2087
| 0.9985 | 12.985 | 9.2056 | 9.4558 | 9.756 | 9.9061 | 10.1062 | 10.1562 | 10.3063 | 10.3563 | 10.4564 | 10.4564 | 10.5064 | 10.4564 | 10.4564 | 10.4064 | 10.2563 | 10.2063 | 10.1062 | 9.9561 | 9.756 | 9.4558 | 9.2557
| 0.999 | 13.982 | 6.2538 | 6.454 | 6.6541 | 6.8042 | 6.9042 | 7.0543 | 7.1544 | 7.1544 | 7.2544 | 7.2544 | 7.3045 | 7.2044 | 7.2044 | 7.1544 | 7.1544 | 6.9543 | 6.8542 | 6.7541 | 6.6541 | 6.454 | 6.2538
| 1.002 | 14.983 | 4.4527 | 4.6028 | 4.7529 | 4.8029 | 4.903 | 4.953 | 5.1031 | 5.0531 | 5.1031 | 5.1532 | 5.1532 | 5.1532 | 5.1031 | 5.0031 | 5.0031 | 4.903 | 4.853 | 4.7529 | 4.7529 | 4.6028 | 4.5028
| 1.001 | 15.986 | 3.302 | 3.3521 | 3.5021 | 3.5522 | 3.6022 | 3.6522 | 3.7523 | 3.7523 | 3.7523 | 3.7523 | 3.7023 | 3.7023 | 3.7023 | 3.6022 | 3.5522 | 3.5522 | 3.5522 | 3.5021 | 3.4521 | 3.302 | 3.252
| 0.9985 | 16.985 | 2.4515 | 2.5516 | 2.6016 | 2.6516 | 2.7517 | 2.8017 | 2.8517 | 2.8017 | 2.7517 | 2.8017 | 2.8017 | 2.7517 | 2.7517 | 2.7517 | 2.7517 | 2.7017 | 2.6516 | 2.6516 | 2.5516 | 2.5015 | 2.5015
| 0.999 | 17.983 | 1.9512 | 1.9512 | 2.0012 | 2.1013 | 2.1513 | 2.1513 | 2.1513 | 2.1513 | 2.1513 | 2.2013 | 2.1513 | 2.1513 | 2.1513 | 2.1513 | 2.1013 | 2.1013 | 2.0513 | 2.0012 | 2.0012 | 2.0012 | 1.9512
| 1.0015 | 18.983 | 1.5509 | 1.601 | 1.651 | 1.651 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.7511 | 1.701 | 1.701 | 1.701 | 1.701 | 1.651 | 1.701 | 1.601 | 1.651 | 1.601 | 1.601 | 1.5509
| 1.001 | 19.986 | 1.3008 | 1.3008 | 1.3508 | 1.3508 | 1.4009 | 1.4009 | 1.4509 | 1.4009 | 1.4009 | 1.4009 | 1.3508 | 1.4009 | 1.3508 | 1.3008 | 1.3008 | 1.3508 | 1.3508 | 1.3008 | 1.3508 | 1.3008 | 1.3008
| 0.9985 | 20.985 | 1.1507 | 1.1507 | 1.1507 | 1.1507 | 1.1507 | 1.2007 | 1.1507 | 1.2007 | 1.1507 | 1.2007 | 1.1507 | 1.1507 | 1.1507 | 1.1507 | 1.1007 | 1.1507 | 1.1507 | 1.1007 | 1.1007 | 1.1507 | 1.1007
| 0.9995 | 21.983 | 0.85052 | 0.85052 | 0.90055 | 0.85052 | 0.90055 | 0.90055 | 0.90055 | 0.90055 | 0.90055 | 0.90055 | 0.85052 | 0.85052 | 0.80049 | 0.80049 | 0.80049 | 0.80049 | 0.80049 | 0.85052 | 0.80049 | 0.80049 | 0.80049
| 1.0015 | 22.984 | 0.75046 | 0.70043 | 0.80049 | 0.80049 | 0.75046 | 0.80049 | 0.80049 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.70043 | 0.75046 | 0.70043 | 0.70043 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.75046 | 0.75046
| 1.001 | 23.986 | 0.60037 | 0.60037 | 0.6504 | 0.6504 | 0.6504 | 0.70043 | 0.70043 | 0.60037 | 0.6504 | 0.6504 | 0.60037 | 0.60037 | 0.55034 | 0.60037 | 0.55034 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037 | 0.60037
| 0.9985 | 24.986 | 0.55034 | 0.55034 | 0.55034 | 0.55034 | 0.55034 | 0.60037 | 0.60037 | 0.55034 | 0.55034 | 0.55034 | 0.50031 | 0.45028 | 0.50031 | 0.45028 | 0.50031 | 0.50031 | 0.55034 | 0.50031 | 0.50031 | 0.55034 | 0.55034
| 0.9995 | 25.983 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.50031 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.40025 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.50031 | 0.45028
| 1.0015 | 26.985 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.45028 | 0.50031 | 0.40025 | 0.45028 | 0.40025 | 0.40025 | 0.35021 | 0.40025 | 0.35021 | 0.35021 | 0.35021 | 0.40025 | 0.40025 | 0.40025 | 0.45028 | 0.45028
| 1.0005 | 27.986 | 0.40025 | 0.40025 | 0.40025 | 0.40025 | 0.40025 | 0.45028 | 0.50031 | 0.35021 | 0.35021 | 0.35021 | 0.30018 | 0.30018 | 0.30018 | 0.25015 | 0.30018 | 0.30018 | 0.35021 | 0.35021 | 0.35021 | 0.35021 | 0.35021
| 0.999 | 28.986 | 0.35021 | 0.35021 | 0.35021 | 0.35021 | 0.35021 | 0.40025 | 0.35021 | 0.30018 | 0.35021 | 0.30018 | 0.30018 | 0.30018 | 0.25015 | 0.25015 | 0.25015 | 0.30018 | 0.30018 | 0.30018 | 0.30018 | 0.35021 | 0.35021
| 0.499 | 29.984 | 0.35021 | 0.35021 | 0.30018 | 0.30018 | 0.30018 | 0.35021 | 0.30018 | 0.30018 | 0.30018 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.25015 | 0.30018 | 0.30018 | 0.30018
|  | X-Pos (cm) | -5 | -4.502 | -4.002 | -3.502 | -3.001 | -2.501 | -2.001 | -1.501 | -1.001 | -0.501 | -0.001 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.501 | 4.001 | 4.501 | 5.001
|  | BdL (G-cm) | 13282.4544 | 13308.6408 | 13327.5234 | 13339.8646 | 13349.1546 | 13354.9434 | 13359.8198 | 13361.7434 | 13362.0795 | 13362.0666 | 13361.2888 | 13359.6228 | 13357.6836 | 13354.3671 | 13350.2511 | 13346.1595 | 13340.1806 | 13330.6117 | 13317.5541 | 13297.2796 | 13265.9631

## NMR Transverse (Pre-Split)
Magnet= | DL001
Location= | DL0L02
Current= | 5.5 | Amps
Probe= | 2
Position= | +Z position is downstream of center, +X position is towards power terminals
X-Position (cm)
Trial | Z-Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
1 | 2.2 | 0.10351 | 0.10351 | 0.10351 | 0.10351 | 0.10351 | 0.10352 | 0.10352 | Z = 2.2
2 | 2.2 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.10351 | Z = 0
3 | 2.2 | 0.10349 | 0.10349 | 0.10349 | 0.10349 | 0.1035 | 0.1035 | 0.10351 | Z = -2
1 | 0 | 0.10352 | 0.10351 | 0.10351 | 0.10351 | 0.10352 | 0.10352 | 0.10353
2 | 0 | 0.10351 | 0.1035 | 0.1035 | 0.1035 | 0.10351 | 0.10351 | 0.10352
3 | 0 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.10351 | 0.10352
1 | -2 | 0.10348 | 0.10348 | 0.10349 | 0.10349 | 0.1035 | 0.1035 | 0.10351
2 | -2 | 0.10348 | 0.10348 | 0.10348 | 0.10349 | 0.10349 | 0.1035 | 0.1035
3 | -2 | 0.10348 | 0.10348 | 0.10348 | 0.10348 | 0.10349 | 0.1035 | 0.1035
X-Position (cm)
| Z Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
AVERAGE | 2.2 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.1035 | 0.10351 | 0.10351
| 0 | 0.10351 | 0.10351 | 0.1035 | 0.1035 | 0.10351 | 0.10352 | 0.10353
| -2 | 0.10348 | 0.10348 | 0.10348 | 0.10349 | 0.10349 | 0.1035 | 0.1035
X-Position (cm)
| Z Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
% DEV. FROM CENTER | -2 | 0.002995% | 0.003404% | 0.00438% | 0.00351% | 0.0008695% | -0.002673% | -0.0076%
| 0 | -0.003671% | -0.001288% | 0.0003865% | 0% | -0.003768% | -0.01208% | -0.0218%
| 2.2 | 0.02367% | 0.02209% | 0.02045% | 0.01681% | 0.01198% | 0.005346% | -0.0007085%

## NMR Transverse (Post-Split)
Magnet= | DL001
Location= | DL0L02
Current= | 5.5 | Amps
Probe= | 2
Position= | +Z position is downstream of center, +X position is towards power terminals
X-Position (cm)
Trial | Z-Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
1 | 2.2 | 0.10355 | 0.10355 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | Z = 2.2
2 | 2.2 | 0.10355 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | Z = 0
3 | 2.2 | 0.10355 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | Z = -2
1 | 0 | 0.10357 | 0.10356 | 0.10356 | 0.10356 | 0.10356 | 0.10356 | 0.10357
2 | 0 | 0.10356 | 0.10356 | 0.10355 | 0.10355 | 0.10355 | 0.10355 | 0.10356
3 | 0 | 0.10355 | 0.10355 | 0.10354 | 0.10354 | 0.10354 | 0.10355 | 0.10355
1 | -2 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10355 | 0.10355 | 0.10355
2 | -2 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10355
3 | -2 | 0.10353 | 0.10353 | 0.10353 | 0.10353 | 0.10353 | 0.10354 | 0.10354
X-Position (cm)
| Z Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
AVERAGE | 2.2 | 0.10355 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354
| 0 | 0.10356 | 0.10355 | 0.10355 | 0.10355 | 0.10355 | 0.10355 | 0.10356
| -2 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10354 | 0.10355
X-Position (cm)
| Z Pos (cm) | 3 | 2 | 1 | 0 | -1 | -2 | -3
% DEV. FROM CENTER | -2 | 0.001159% | 0.004571% | 0.008595% | 0.009432% | 0.01056% | 0.01037% | 0.008853%
| 0 | -0.01014% | -0.005794% | -0.000676% | 0% | 0.0001288% | -0.003766% | -0.009786%
| 2.2 | 0.008788% | 0.009174% | 0.01037% | 0.009947% | 0.008466% | 0.005601% | 0.001803%

## Field Map
| Magnet Loc: | MDL0L02
| Magnet ID: | DL001
| Meas. #: | DL001009.stp
| Meas. Date: | 2014-08-29
| Coil used: | Hall Probe Stepper
| Current (Amps) | Strength (Gauss)
| -------------- | -------------------
| -9.9925 | -23944.1907
| -8.9956 | -21569.5661
| -7.991 | -19168.9954
| -6.9895 | -16769.5002
| -5.9903 | -14360.6554
| -4.9927 | -11954.6359
| -3.9938 | -9542.7929
| -2.9892 | -7116.0522
| -1.9895 | -4698.197
| -0.99049 | -2283.5543
| 0.003348 | 125.9789
| 1.0092 | 2548.3863
| 2.0091 | 4960.8155
| 3.0094 | 7374.5829
| 4.0102 | 9785.8494
| 5.0105 | 12191.9526
| 6.0102 | 14589.7536
| 7.0114 | 16980.3787
| 8.0132 | 19360.3629
| 9.0147 | 21720.5306
| 10.0137 | 24038.0631
0 | 9.0147 | 21720.5306
-1 | 8.0132 | 19360.3629
-2 | 7.0114 | 16980.3787
-3 | 6.0102 | 14589.7536
-4 | 5.0105 | 12191.9526
-5 | 4.0102 | 9785.8494
-6 | 3.0094 | 7374.5829
-7 | 2.0091 | 4960.8155
-8 | 1.0092 | 2548.3863
-9 | 0.003348 | 125.9789
-10 | -0.99049 | -2283.5543
-11 | -1.9895 | -4698.197
-12 | -2.9892 | -7116.0522
-13 | -3.9938 | -9542.7929
-14 | -4.9927 | -11954.6359
-15 | -5.9903 | -14360.6554
-16 | -6.9895 | -16769.5002
-17 | -7.991 | -19168.9954
-18 | -8.9956 | -21569.5661
-19 | -9.9925 | -23944.1907

## NMR Calibration
NMR Calibration |  | 2014-08-29
X, Z = 0cm
CENTER CORE FIELD
| Current | NMR | Hall Probe | Correction
| 10 | 1864.189 | 1863.4 | 100%
| 9 | 1683.017 | 1682.7 | 100%
| 8 | 1499.463 | 1499.15 | 100%
| 7 | 1314.746 | 1314.35 | 100%
| 6 | 1129.209 | 1128.7 | 100%
| 5 | 943.101 | 942.75 | 100% | 100%
| -6 | 1116.248 | -1116.25 | 100%
| -7 | 1302.8 | -1302.9 | 99.99%
| -8 | 1489.068 | -1489.25 | 99.99%
| -9 | 1675.439 | -1675.6 | 99.99%
| -10 | 1859.815 | -1860.05 | 99.99% | 99.99%
NMR Calibration |  | 2014-08-21
Z=0cm
I = 5.5A
| X-Pos | NMR | Hall Probe | Correction
| -5 | 1035.555 | 1035.3 | 100%
| -4.5 | 1035.754 | 1035.5 | 100%
| -4 | 1035.818 | 1035.45 | 100%
| -3.5 | 1035.82 | 1035.45 | 100%
| -3 | 1035.814 | 1035.4 | 100%
| -2.5 | 1035.809 | 1035.35 | 100%
| -2 | 1035.802 | 1035.3 | 100%
| -1.5 | 1035.798 | 1035.2 | 100.1%
| -1 | 1035.794 | 1035.2 | 100.1%
| -0.5 | 1035.795 | 1035.2 | 100.1%
| 0 | 1035.854 | 1035.2 | 100.1%
| 0.5 | 1035.867 | 1035.2 | 100.1%
| 1 | 1035.904 | 1035.2 | 100.1%
| 1.5 | 1035.93 | 1035.15 | 100.1%
| 2 | 1035.96 | 1035.15 | 100.1%
| 2.5 | 1035.997 | 1035.2 | 100.1%
| 3 | 1036.035 | 1035.2 | 100.1%
| 3.5 | 1036.075 | 1035.2 | 100.1%
| 4 | 1036.106 | 1035.25 | 100.1%
| 4.5 | 1036.091 | 1035.2 | 100.1%
| 5 | 1035.994 | 1034.95 | 100.1%
100.1%

## X-Transverse Data
! begin File Header
DL001002.stp | ! data file
2014-08-21 | ! date
10:33 | ! start time
5.1 | ! stepper code version
0 | ! number Z of currents
21 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset |  |  | 16.93
Note: X Transverse, I=5.5A
! end File Header
! begin Step Header
0 | ! set current
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header |  |  |  |  | 11.93 | 12.428 | 12.928 | 13.429 | 13.929 | 14.429 | 14.929 | 15.429 | 15.929 | 16.429 | 16.929 | 17.43 | 17.93 | 18.43 | 18.93 | 19.43 | 19.93 | 20.431 | 20.931 | 21.431 | 21.931
Position (Z, X) |  | Hall Probe #1 | Voltmeter |  | -5 | -4.502 | -4.002 | -3.501 | -3.001 | -2.501 | -2.001 | -1.501 | -1.001 | -0.501 | -0.001 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.501 | 4.001 | 4.501 | 5.001
102.172 | 11.93 | 1035.3 | 0.1836 |  | 1035.3 | 1035.5 | 1035.45 | 1035.45 | 1035.4 | 1035.35 | 1035.3 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.15 | 1035.15 | 1035.2 | 1035.2 | 1035.2 | 1035.25 | 1035.2 | 1034.95
102.172 | 12.428 | 1035.5 | 0.18359 |  | 0.009659% | 0.02897% | 0.02414% | 0.02414% | 0.01932% | 0.01449% | 0.009659% | 0% | 0% | 0% | 0% | 0% | 0% | -0.00483% | -0.00483% | 0% | 0% | 0% | 0.00483% | 0% | -0.02416%
102.172 | 12.928 | 1035.45 | 0.1836
102.172 | 13.429 | 1035.45 | 0.1836
102.172 | 13.929 | 1035.4 | 0.1836
102.172 | 14.429 | 1035.35 | 0.18359
102.172 | 14.929 | 1035.3 | 0.1836
102.172 | 15.429 | 1035.2 | 0.18359
102.172 | 15.929 | 1035.2 | 0.18359
102.172 | 16.429 | 1035.2 | 0.1836
102.172 | 16.929 | 1035.2 | 0.1836
102.172 | 17.43 | 1035.2 | 0.1836
102.172 | 17.93 | 1035.2 | 0.18359
102.172 | 18.43 | 1035.15 | 0.18359
102.172 | 18.93 | 1035.15 | 0.18359
102.172 | 19.43 | 1035.2 | 0.18359
102.172 | 19.93 | 1035.2 | 0.18359
102.172 | 20.431 | 1035.2 | 0.1836
102.172 | 20.931 | 1035.25 | 0.18359
102.172 | 21.431 | 1035.2 | 0.18359
102.172 | 21.931 | 1034.95 | 0.1836
! begin File Header
DL001003.stp | ! data file
2014-08-21 | ! date
10:41 | ! start time
5.1 | ! stepper code version
0 | ! number Z of currents
21 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset |  |  | 16.93
Note: X Transverse, I=5.5A
! end File Header
! begin Step Header
0 | ! set current
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header |  |  |  |  | 11.93 | 12.428 | 12.928 | 13.429 | 13.929 | 14.429 | 14.929 | 15.429 | 15.929 | 16.429 | 16.929 | 17.43 | 17.93 | 18.43 | 18.93 | 19.43 | 19.93 | 20.431 | 20.931 | 21.431 | 21.931
Position (Z, X) |  | Hall Probe #1 | Voltmeter |  | -5 | -4.502 | -4.002 | -3.501 | -3.001 | -2.501 | -2.001 | -1.501 | -1.001 | -0.501 | -0.001 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.501 | 4.001 | 4.501 | 5.001
102.172 | 11.93 | 1035.3 | 0.1836 |  | 1035.3 | 1035.5 | 1035.45 | 1035.45 | 1035.4 | 1035.35 | 1035.3 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.2 | 1035.15 | 1035.15 | 1035.2 | 1035.2 | 1035.2 | 1035.25 | 1035.2 | 1034.95
102.172 | 12.428 | 1035.45 | 0.1836 |  | 0.009659% | 0.02897% | 0.02414% | 0.02414% | 0.01932% | 0.01449% | 0.009659% | 0% | 0% | 0% | 0% | 0% | 0% | -0.00483% | -0.00483% | 0% | 0% | 0% | 0.00483% | 0% | -0.02416%
102.172 | 12.928 | 1035.5 | 0.1836
102.172 | 13.428 | 1035.45 | 0.1836
102.172 | 13.929 | 1035.4 | 0.18359
102.172 | 14.429 | 1035.35 | 0.1836
102.172 | 14.929 | 1035.25 | 0.1836
102.172 | 15.429 | 1035.2 | 0.1836
102.172 | 15.929 | 1035.2 | 0.1836
102.172 | 16.429 | 1035.2 | 0.1836
102.172 | 16.929 | 1035.2 | 0.1836
102.172 | 17.43 | 1035.2 | 0.18359
102.172 | 17.929 | 1035.15 | 0.1836
102.172 | 18.43 | 1035.2 | 0.18359
102.172 | 18.93 | 1035.2 | 0.1836
102.172 | 19.43 | 1035.2 | 0.18359
102.172 | 19.93 | 1035.2 | 0.18359
102.172 | 20.43 | 1035.2 | 0.18359
102.172 | 20.931 | 1035.2 | 0.18359
102.172 | 21.431 | 1035.2 | 0.18359
102.172 | 21.931 | 1035 | 0.18359

## Center BdL Data
! begin File Header |  |  |  | 103 | Rows between values
DL001009.stp | ! data file
2014-08-29 | ! date
10:22 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=10A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 10.0137
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter |  | Current (A) | Core Field (G)
16.782 | 73.77 | 0.8 | 0.33379 | 0 | 10.0137 | 1863.4
16.782 | 74.769 | 0.85 | 0.33379 | 103 | 9.0147 | 1682.7
16.782 | 75.768 | 0.9 | 0.33379 | 206 | 8.0132 | 1499.15
16.782 | 76.767 | 1 | 0.33379 | 309 | 7.0114 | 1314.35
16.782 | 77.767 | 1.1 | 0.33379 | 412 | 6.0102 | 1128.7
16.782 | 78.768 | 1.2 | 0.33379 | 515 | 5.0105 | 942.75
16.782 | 79.77 | 1.35 | 0.33379 | 618 | 4.0102 | 756.15
16.782 | 80.765 | 1.5 | 0.33379 | 721 | 3.0094 | 569.25
16.782 | 81.767 | 1.75 | 0.33379 | 824 | 2.0091 | 382.25
16.782 | 82.769 | 2 | 0.33379 | 927 | 1.0092 | 195.15
16.782 | 83.768 | 2.5 | 0.33379 | 1030 | 0.003348 | 7.4
16.782 | 84.767 | 3 | 0.33379 | 1133 | -0.99049 | -179.45
16.782 | 85.766 | 3.75 | 0.33379 | 1236 | -1.9895 | -366.65
16.782 | 86.768 | 4.85 | 0.33379 | 1339 | -2.9892 | -554.15
16.782 | 87.769 | 6.4 | 0.33379 | 1442 | -3.9938 | -742.5
16.782 | 88.766 | 8.75 | 0.33379 | 1545 | -4.9927 | -929.65
16.782 | 89.766 | 12.25 | 0.33379 | 1648 | -5.9903 | -1116.25
16.782 | 90.767 | 17.95 | 0.33379 | 1751 | -6.9895 | -1302.9
16.782 | 91.769 | 27.4 | 0.33379 | 1854 | -7.991 | -1489.25
16.782 | 92.268 | 34.4 | 0.33379 | 1957 | -8.9956 | -1675.6
16.782 | 92.766 | 43.85 | 0.33379 | 2060 | -9.9925 | -1860.05
16.782 | 93.267 | 56.75 | 0.33379
16.782 | 93.766 | 74.5 | 0.33379
16.782 | 94.267 | 99.6 | 0.33379
16.782 | 94.767 | 135.4 | 0.33379
16.782 | 95.265 | 186.8 | 0.33379
16.782 | 95.768 | 261.45 | 0.33379
16.782 | 96.266 | 367.05 | 0.33379
16.782 | 96.765 | 514.25 | 0.33379
16.782 | 97.268 | 715.5 | 0.33379
16.782 | 97.766 | 976.95 | 0.33379
16.782 | 98.266 | 1289 | 0.33379
16.782 | 98.767 | 1577.35 | 0.33379
16.782 | 99.265 | 1753.3 | 0.33379
16.782 | 99.768 | 1827.35 | 0.33379
16.782 | 100.266 | 1852.35 | 0.33379
16.782 | 100.765 | 1860.1 | 0.33379
16.782 | 101.267 | 1862.5 | 0.33379
16.782 | 101.765 | 1863.2 | 0.33379
16.782 | 102.267 | 1863.4 | 0.33379
16.782 | 102.767 | 1863.4 | 0.33379
16.782 | 103.264 | 1863.4 | 0.33379
16.782 | 103.767 | 1863.4 | 0.33379
16.782 | 104.266 | 1863.45 | 0.33379
16.782 | 104.764 | 1863.4 | 0.33379
16.782 | 105.267 | 1863.35 | 0.33379
16.782 | 105.765 | 1863.15 | 0.33379
16.782 | 106.266 | 1862.4 | 0.33379
16.782 | 106.767 | 1859.95 | 0.33379
16.782 | 107.264 | 1851.85 | 0.33379
16.782 | 107.767 | 1825.75 | 0.33379
16.782 | 108.265 | 1750.05 | 0.33379
16.782 | 108.764 | 1571.2 | 0.33379
16.782 | 109.267 | 1279.2 | 0.33379
16.782 | 109.765 | 969.9 | 0.33379
16.782 | 110.265 | 709.45 | 0.33379
16.782 | 110.766 | 510.75 | 0.33379
16.782 | 111.263 | 365.2 | 0.33379
16.782 | 111.767 | 259.2 | 0.33379
16.782 | 112.265 | 185.55 | 0.33379
16.782 | 112.763 | 134.35 | 0.33379
16.782 | 113.265 | 98.7 | 0.33379
16.782 | 113.763 | 73.9 | 0.33379
16.782 | 114.264 | 56.15 | 0.33379
16.782 | 114.765 | 43.3 | 0.33379
16.782 | 115.263 | 33.95 | 0.33379
16.782 | 115.765 | 27.05 | 0.33379
16.782 | 116.762 | 17.75 | 0.33379
16.782 | 117.763 | 12.15 | 0.33379
16.782 | 118.764 | 8.6 | 0.33379
16.782 | 119.765 | 6.3 | 0.33379
16.782 | 120.762 | 4.8 | 0.33379
16.782 | 121.763 | 3.7 | 0.33379
16.782 | 122.765 | 2.95 | 0.33379
16.782 | 123.765 | 2.45 | 0.33379
16.782 | 124.763 | 2 | 0.33379
16.782 | 125.763 | 1.75 | 0.33379
16.782 | 126.766 | 1.5 | 0.33379
16.782 | 127.765 | 1.3 | 0.33379
16.782 | 128.763 | 1.15 | 0.33379
16.782 | 129.764 | 1.05 | 0.33379
16.782 | 130.765 | 0.95 | 0.33379
16.782 | 131.766 | 0.9 | 0.33379
16.782 | 132.763 | 0.8 | 0.33379
16.782 | 133.765 | 0.8 | 0.33379
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
10:31 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=9A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 9.0147
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.77 | 0.8 | 0.30049
16.782 | 74.769 | 0.85 | 0.30049
16.782 | 75.768 | 0.9 | 0.30049
16.782 | 76.767 | 1 | 0.30049
16.782 | 77.766 | 1.05 | 0.30049
16.782 | 78.769 | 1.2 | 0.30049
16.782 | 79.769 | 1.3 | 0.30049
16.782 | 80.766 | 1.45 | 0.30049
16.782 | 81.767 | 1.7 | 0.30049
16.782 | 82.769 | 2 | 0.30049
16.782 | 83.77 | 2.35 | 0.30049
16.782 | 84.767 | 2.9 | 0.30049
16.782 | 85.766 | 3.6 | 0.30049
16.782 | 86.769 | 4.6 | 0.30049
16.782 | 87.768 | 6 | 0.30049
16.782 | 88.767 | 8.15 | 0.30049
16.782 | 89.767 | 11.35 | 0.30049
16.782 | 90.767 | 16.55 | 0.30049
16.782 | 91.769 | 25.15 | 0.30049
16.782 | 92.267 | 31.5 | 0.30049
16.782 | 92.765 | 40 | 0.30049
16.782 | 93.268 | 51.75 | 0.30049
16.782 | 93.765 | 67.8 | 0.30049
16.782 | 94.267 | 90.55 | 0.30049
16.782 | 94.768 | 123 | 0.30049
16.782 | 95.265 | 169.3 | 0.30049
16.782 | 95.767 | 236.65 | 0.30049
16.782 | 96.266 | 332.15 | 0.30049
16.782 | 96.766 | 465.35 | 0.30049
16.782 | 97.268 | 646.95 | 0.30049
16.782 | 97.765 | 882.65 | 0.30049
16.782 | 98.265 | 1164.15 | 0.30049
16.782 | 98.767 | 1424.95 | 0.30049
16.782 | 99.265 | 1583.5 | 0.30049
16.782 | 99.768 | 1650.3 | 0.30049
16.782 | 100.266 | 1672.75 | 0.30049
16.782 | 100.765 | 1679.8 | 0.30049
16.782 | 101.267 | 1681.95 | 0.30049
16.782 | 101.766 | 1682.55 | 0.30049
16.782 | 102.267 | 1682.65 | 0.30049
16.782 | 102.767 | 1682.7 | 0.30049
16.782 | 103.264 | 1682.7 | 0.30049
16.782 | 103.767 | 1682.7 | 0.30049
16.782 | 104.266 | 1682.7 | 0.30049
16.782 | 104.764 | 1682.7 | 0.30049
16.782 | 105.267 | 1682.65 | 0.30049
16.782 | 105.765 | 1682.45 | 0.30049
16.782 | 106.266 | 1681.85 | 0.30049
16.782 | 106.767 | 1679.55 | 0.30049
16.782 | 107.264 | 1672.3 | 0.30049
16.782 | 107.767 | 1648.75 | 0.30049
16.782 | 108.266 | 1580.25 | 0.30049
16.782 | 108.763 | 1418.9 | 0.30049
16.782 | 109.267 | 1155.45 | 0.30049
16.782 | 109.764 | 876.2 | 0.30049
16.782 | 110.265 | 640.95 | 0.30049
16.782 | 110.765 | 461.95 | 0.30049
16.782 | 111.263 | 330.45 | 0.30049
16.782 | 111.766 | 234.75 | 0.30049
16.782 | 112.265 | 168.15 | 0.30049
16.782 | 112.763 | 121.95 | 0.30049
16.782 | 113.265 | 89.6 | 0.30049
16.782 | 113.763 | 67.2 | 0.30049
16.782 | 114.263 | 51.15 | 0.30049
16.782 | 114.765 | 39.6 | 0.30049
16.782 | 115.262 | 31.15 | 0.30049
16.782 | 115.765 | 24.8 | 0.30049
16.782 | 116.762 | 16.35 | 0.30049
16.782 | 117.763 | 11.25 | 0.30049
16.782 | 118.764 | 8 | 0.30049
16.782 | 119.764 | 5.9 | 0.30049
16.782 | 120.762 | 4.5 | 0.30049
16.782 | 121.763 | 3.55 | 0.30049
16.782 | 122.764 | 2.8 | 0.30049
16.782 | 123.764 | 2.3 | 0.30049
16.782 | 124.762 | 1.95 | 0.30049
16.782 | 125.763 | 1.65 | 0.30049
16.782 | 126.765 | 1.45 | 0.30049
16.782 | 127.765 | 1.3 | 0.30049
16.782 | 128.762 | 1.2 | 0.30049
16.782 | 129.764 | 1.05 | 0.30049
16.782 | 130.765 | 0.95 | 0.30049
16.782 | 131.766 | 0.85 | 0.30049
16.782 | 132.763 | 0.8 | 0.30049
16.782 | 133.765 | 0.75 | 0.30049
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
10:40 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=8A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 8.0132
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.77 | 0.8 | 0.26711
16.782 | 74.769 | 0.85 | 0.2671
16.782 | 75.769 | 0.9 | 0.26711
16.782 | 76.767 | 0.95 | 0.26711
16.782 | 77.767 | 1.05 | 0.26711
16.782 | 78.769 | 1.15 | 0.26711
16.782 | 79.769 | 1.25 | 0.26711
16.782 | 80.765 | 1.45 | 0.26711
16.782 | 81.766 | 1.6 | 0.26711
16.782 | 82.769 | 1.9 | 0.26711
16.782 | 83.769 | 2.2 | 0.26711
16.782 | 84.767 | 2.65 | 0.26711
16.782 | 85.766 | 3.35 | 0.26711
16.782 | 86.769 | 4.25 | 0.26711
16.782 | 87.768 | 5.55 | 0.26711
16.782 | 88.766 | 7.45 | 0.26711
16.782 | 89.767 | 10.35 | 0.26711
16.782 | 90.768 | 14.95 | 0.2671
16.782 | 91.768 | 22.6 | 0.26711
16.782 | 92.268 | 28.3 | 0.26711
16.782 | 92.765 | 35.95 | 0.26711
16.782 | 93.267 | 46.4 | 0.26711
16.782 | 93.766 | 60.75 | 0.26711
16.782 | 94.266 | 81.05 | 0.26711
16.782 | 94.767 | 109.9 | 0.26711
16.782 | 95.265 | 151.3 | 0.26711
16.782 | 95.767 | 211.35 | 0.26711
16.782 | 96.266 | 296.3 | 0.26711
16.782 | 96.765 | 414.75 | 0.26711
16.782 | 97.268 | 576.7 | 0.26711
16.782 | 97.765 | 786.85 | 0.26711
16.782 | 98.266 | 1038.1 | 0.26711
16.782 | 98.767 | 1269.7 | 0.26711
16.782 | 99.265 | 1410.9 | 0.26711
16.782 | 99.768 | 1470.4 | 0.26711
16.782 | 100.266 | 1490.35 | 0.26711
16.782 | 100.764 | 1496.6 | 0.26711
16.782 | 101.267 | 1498.5 | 0.26711
16.782 | 101.765 | 1499.05 | 0.26711
16.782 | 102.266 | 1499.2 | 0.26711
16.782 | 102.767 | 1499.15 | 0.26711
16.782 | 103.264 | 1499.2 | 0.26711
16.782 | 103.767 | 1499.15 | 0.26711
16.782 | 104.266 | 1499.15 | 0.26711
16.782 | 104.764 | 1499.15 | 0.26711
16.782 | 105.267 | 1499.15 | 0.26711
16.782 | 105.765 | 1498.95 | 0.26711
16.782 | 106.265 | 1498.45 | 0.26711
16.782 | 106.767 | 1496.4 | 0.26711
16.782 | 107.264 | 1489.9 | 0.26711
16.782 | 107.767 | 1469 | 0.26711
16.782 | 108.265 | 1407.95 | 0.26711
16.782 | 108.763 | 1263.95 | 0.26711
16.782 | 109.267 | 1029.25 | 0.26711
16.782 | 109.764 | 780.65 | 0.26711
16.782 | 110.265 | 571.3 | 0.26711
16.782 | 110.766 | 411.6 | 0.26711
16.782 | 111.263 | 294.6 | 0.26711
16.782 | 111.766 | 209.35 | 0.26711
16.782 | 112.265 | 150.15 | 0.26711
16.782 | 112.763 | 108.95 | 0.26711
16.782 | 113.266 | 80.2 | 0.26711
16.782 | 113.763 | 60.2 | 0.26711
16.782 | 114.264 | 45.9 | 0.26711
16.782 | 114.765 | 35.55 | 0.26711
16.782 | 115.263 | 27.95 | 0.26711
16.782 | 115.765 | 22.35 | 0.26711
16.782 | 116.762 | 14.75 | 0.26711
16.782 | 117.763 | 10.15 | 0.26711
16.782 | 118.764 | 7.3 | 0.2671
16.782 | 119.765 | 5.45 | 0.26711
16.782 | 120.763 | 4.15 | 0.26711
16.782 | 121.763 | 3.25 | 0.2671
16.782 | 122.765 | 2.65 | 0.26711
16.782 | 123.765 | 2.15 | 0.26711
16.782 | 124.762 | 1.85 | 0.26711
16.782 | 125.763 | 1.6 | 0.2671
16.782 | 126.765 | 1.4 | 0.26711
16.782 | 127.765 | 1.25 | 0.26711
16.782 | 128.763 | 1.15 | 0.2671
16.782 | 129.764 | 1.05 | 0.2671
16.782 | 130.765 | 0.95 | 0.26711
16.782 | 131.766 | 0.85 | 0.26711
16.782 | 132.763 | 0.8 | 0.2671
16.782 | 133.764 | 0.8 | 0.26711
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
10:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=7A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 7.0114
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.77 | 0.8 | 0.23371
16.782 | 74.769 | 0.8 | 0.23372
16.782 | 75.768 | 0.85 | 0.23372
16.782 | 76.767 | 0.9 | 0.23372
16.782 | 77.766 | 1 | 0.23372
16.782 | 78.769 | 1.1 | 0.23372
16.782 | 79.769 | 1.2 | 0.23372
16.782 | 80.765 | 1.35 | 0.23372
16.782 | 81.767 | 1.5 | 0.23372
16.782 | 82.769 | 1.75 | 0.23371
16.782 | 83.769 | 2.05 | 0.23372
16.782 | 84.767 | 2.45 | 0.23372
16.782 | 85.766 | 3.05 | 0.23372
16.782 | 86.768 | 3.85 | 0.23372
16.782 | 87.769 | 5 | 0.23371
16.782 | 88.766 | 6.65 | 0.23372
16.782 | 89.766 | 9.2 | 0.23372
16.782 | 90.767 | 13.25 | 0.23372
16.782 | 91.769 | 20 | 0.23372
16.782 | 92.267 | 25.05 | 0.23372
16.782 | 92.765 | 31.75 | 0.23372
16.782 | 93.267 | 40.9 | 0.23372
16.782 | 93.766 | 53.5 | 0.23372
16.782 | 94.267 | 71.35 | 0.23372
16.782 | 94.767 | 96.65 | 0.23372
16.782 | 95.265 | 132.95 | 0.23372
16.782 | 95.767 | 185.5 | 0.23372
16.782 | 96.266 | 260 | 0.23372
16.782 | 96.765 | 363.9 | 0.23372
16.782 | 97.268 | 505.65 | 0.23372
16.782 | 97.765 | 689.7 | 0.23371
16.782 | 98.266 | 910.2 | 0.23371
16.782 | 98.767 | 1113.25 | 0.23372
16.782 | 99.265 | 1237.1 | 0.23372
16.782 | 99.768 | 1289.1 | 0.23372
16.782 | 100.266 | 1306.6 | 0.23372
16.782 | 100.765 | 1312.15 | 0.23371
16.782 | 101.267 | 1313.75 | 0.23371
16.782 | 101.766 | 1314.25 | 0.23371
16.782 | 102.266 | 1314.35 | 0.23371
16.782 | 102.767 | 1314.35 | 0.23372
16.782 | 103.264 | 1314.35 | 0.23372
16.782 | 103.767 | 1314.35 | 0.23371
16.782 | 104.266 | 1314.35 | 0.23371
16.782 | 104.764 | 1314.3 | 0.23371
16.782 | 105.267 | 1314.25 | 0.23371
16.782 | 105.765 | 1314.05 | 0.23372
16.782 | 106.265 | 1313.6 | 0.23371
16.782 | 106.766 | 1311.95 | 0.23372
16.782 | 107.264 | 1306.3 | 0.23371
16.782 | 107.767 | 1287.85 | 0.23371
16.782 | 108.265 | 1234.35 | 0.23371
16.782 | 108.763 | 1108.2 | 0.23371
16.782 | 109.267 | 902.5 | 0.23371
16.782 | 109.765 | 684.35 | 0.23371
16.782 | 110.265 | 500.9 | 0.23371
16.782 | 110.766 | 360.9 | 0.23371
16.782 | 111.263 | 258.4 | 0.23371
16.782 | 111.767 | 183.75 | 0.23371
16.782 | 112.265 | 131.8 | 0.23371
16.782 | 112.763 | 95.75 | 0.23371
16.782 | 113.266 | 70.55 | 0.23371
16.782 | 113.763 | 53.05 | 0.23372
16.782 | 114.264 | 40.45 | 0.23371
16.782 | 114.765 | 31.45 | 0.23371
16.782 | 115.263 | 24.75 | 0.23372
16.782 | 115.765 | 19.8 | 0.23372
16.782 | 116.762 | 13.15 | 0.23372
16.782 | 117.763 | 9.15 | 0.23372
16.782 | 118.764 | 6.6 | 0.23372
16.782 | 119.765 | 4.9 | 0.23372
16.782 | 120.762 | 3.8 | 0.23371
16.782 | 121.763 | 3.05 | 0.23372
16.782 | 122.765 | 2.45 | 0.23371
16.782 | 123.765 | 2.05 | 0.23372
16.782 | 124.763 | 1.75 | 0.23372
16.782 | 125.763 | 1.5 | 0.23371
16.782 | 126.765 | 1.35 | 0.23371
16.782 | 127.765 | 1.2 | 0.23372
16.782 | 128.763 | 1.15 | 0.23371
16.782 | 129.764 | 1 | 0.23372
16.782 | 130.765 | 0.9 | 0.23372
16.782 | 131.766 | 0.85 | 0.23371
16.782 | 132.763 | 0.8 | 0.23371
16.782 | 133.765 | 0.75 | 0.23371
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
10:59 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=6A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 6.0102
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.77 | 0.8 | 0.20034
16.782 | 74.769 | 0.8 | 0.20034
16.782 | 75.768 | 0.85 | 0.20034
16.782 | 76.767 | 0.9 | 0.20034
16.782 | 77.766 | 0.9 | 0.20034
16.782 | 78.768 | 1 | 0.20034
16.782 | 79.769 | 1.15 | 0.20034
16.782 | 80.766 | 1.25 | 0.20034
16.782 | 81.767 | 1.4 | 0.20034
16.782 | 82.768 | 1.6 | 0.20034
16.782 | 83.77 | 1.85 | 0.20034
16.782 | 84.766 | 2.25 | 0.20034
16.782 | 85.766 | 2.7 | 0.20034
16.782 | 86.769 | 3.4 | 0.20034
16.782 | 87.768 | 4.4 | 0.20034
16.782 | 88.766 | 5.85 | 0.20034
16.782 | 89.767 | 8.1 | 0.20034
16.782 | 90.767 | 11.6 | 0.20034
16.782 | 91.768 | 17.4 | 0.20034
16.782 | 92.267 | 21.75 | 0.20034
16.782 | 92.765 | 27.45 | 0.20034
16.782 | 93.267 | 35.4 | 0.20034
16.782 | 93.765 | 46.2 | 0.20034
16.782 | 94.267 | 61.5 | 0.20034
16.782 | 94.767 | 83.25 | 0.20034
16.782 | 95.265 | 114.45 | 0.20034
16.782 | 95.767 | 159.55 | 0.20034
16.782 | 96.266 | 223.5 | 0.20034
16.782 | 96.766 | 312.95 | 0.20034
16.782 | 97.268 | 434.55 | 0.20034
16.782 | 97.765 | 592.55 | 0.20034
16.782 | 98.266 | 781.75 | 0.20034
16.782 | 98.767 | 956.1 | 0.20034
16.782 | 99.265 | 1062.4 | 0.20034
16.782 | 99.768 | 1107.15 | 0.20034
16.782 | 100.266 | 1122.15 | 0.20034
16.782 | 100.765 | 1126.8 | 0.20034
16.782 | 101.267 | 1128.2 | 0.20034
16.782 | 101.765 | 1128.6 | 0.20034
16.782 | 102.267 | 1128.7 | 0.20034
16.782 | 102.767 | 1128.65 | 0.20034
16.782 | 103.264 | 1128.7 | 0.20034
16.782 | 103.767 | 1128.7 | 0.20034
16.782 | 104.266 | 1128.75 | 0.20034
16.782 | 104.764 | 1128.75 | 0.20034
16.782 | 105.267 | 1128.7 | 0.20034
16.782 | 105.765 | 1128.6 | 0.20034
16.782 | 106.266 | 1128.2 | 0.20034
16.782 | 106.767 | 1126.75 | 0.20034
16.782 | 107.264 | 1121.9 | 0.20034
16.782 | 107.767 | 1106.1 | 0.20034
16.782 | 108.265 | 1060.2 | 0.20034
16.782 | 108.763 | 951.85 | 0.20034
16.782 | 109.267 | 775.2 | 0.20034
16.782 | 109.764 | 588 | 0.20034
16.782 | 110.265 | 430.3 | 0.20034
16.782 | 110.766 | 310.15 | 0.20034
16.782 | 111.263 | 222.1 | 0.20034
16.782 | 111.766 | 158.05 | 0.20034
16.782 | 112.265 | 113.45 | 0.20034
16.782 | 112.763 | 82.45 | 0.20034
16.782 | 113.265 | 60.8 | 0.20034
16.782 | 113.763 | 45.8 | 0.20034
16.782 | 114.264 | 35 | 0.20034
16.782 | 114.765 | 27.2 | 0.20034
16.782 | 115.262 | 21.5 | 0.20034
16.782 | 115.765 | 17.2 | 0.20034
16.782 | 116.762 | 11.5 | 0.20034
16.782 | 117.763 | 7.95 | 0.20034
16.782 | 118.764 | 5.75 | 0.20034
16.782 | 119.765 | 4.35 | 0.20034
16.782 | 120.762 | 3.4 | 0.20034
16.782 | 121.763 | 2.65 | 0.20034
16.782 | 122.765 | 2.2 | 0.20034
16.782 | 123.765 | 1.85 | 0.20034
16.782 | 124.762 | 1.6 | 0.20034
16.782 | 125.763 | 1.4 | 0.20034
16.782 | 126.765 | 1.25 | 0.20034
16.782 | 127.765 | 1.1 | 0.20034
16.782 | 128.762 | 1 | 0.20034
16.782 | 129.763 | 0.9 | 0.20034
16.782 | 130.765 | 0.85 | 0.20034
16.782 | 131.766 | 0.8 | 0.20034
16.782 | 132.763 | 0.75 | 0.20034
16.782 | 133.764 | 0.7 | 0.20034
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:06 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=5A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 5.0105
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.65 | 0.16702
16.782 | 74.769 | 0.7 | 0.16702
16.782 | 75.768 | 0.75 | 0.16702
16.782 | 76.766 | 0.8 | 0.16702
16.782 | 77.767 | 0.85 | 0.16702
16.782 | 78.768 | 0.95 | 0.16702
16.782 | 79.769 | 1 | 0.16702
16.782 | 80.766 | 1.15 | 0.16702
16.782 | 81.767 | 1.3 | 0.16702
16.782 | 82.769 | 1.4 | 0.16702
16.782 | 83.769 | 1.65 | 0.16702
16.782 | 84.767 | 1.95 | 0.16702
16.782 | 85.766 | 2.35 | 0.16701
16.782 | 86.768 | 2.95 | 0.16702
16.782 | 87.768 | 3.8 | 0.16701
16.782 | 88.765 | 5 | 0.16702
16.782 | 89.766 | 6.9 | 0.16702
16.782 | 90.767 | 9.85 | 0.16701
16.782 | 91.768 | 14.7 | 0.16702
16.782 | 92.268 | 18.3 | 0.16702
16.782 | 92.765 | 23.15 | 0.16702
16.782 | 93.267 | 29.7 | 0.16702
16.782 | 93.766 | 38.8 | 0.16702
16.782 | 94.267 | 51.6 | 0.16702
16.782 | 94.767 | 69.75 | 0.16702
16.782 | 95.265 | 95.8 | 0.16702
16.782 | 95.768 | 133.5 | 0.16702
16.782 | 96.266 | 186.9 | 0.16702
16.782 | 96.765 | 261.25 | 0.16702
16.782 | 97.268 | 363.05 | 0.16702
16.782 | 97.765 | 495.3 | 0.16701
16.782 | 98.266 | 653.2 | 0.16702
16.782 | 98.767 | 798.7 | 0.16701
16.782 | 99.265 | 887.45 | 0.16702
16.782 | 99.767 | 924.8 | 0.16701
16.782 | 100.266 | 937.35 | 0.16701
16.782 | 100.765 | 941.25 | 0.16702
16.782 | 101.267 | 942.4 | 0.16702
16.782 | 101.765 | 942.7 | 0.16701
16.782 | 102.266 | 942.8 | 0.16701
16.782 | 102.766 | 942.8 | 0.16702
16.782 | 103.264 | 942.8 | 0.16702
16.782 | 103.767 | 942.75 | 0.16702
16.782 | 104.266 | 942.8 | 0.16702
16.782 | 104.764 | 942.8 | 0.16701
16.782 | 105.267 | 942.8 | 0.16702
16.782 | 105.765 | 942.7 | 0.16701
16.782 | 106.265 | 942.4 | 0.16701
16.782 | 106.767 | 941.2 | 0.16701
16.782 | 107.264 | 937.15 | 0.16702
16.782 | 107.766 | 924 | 0.16702
16.782 | 108.265 | 885.65 | 0.16701
16.782 | 108.763 | 795.2 | 0.16702
16.782 | 109.267 | 647.6 | 0.16701
16.782 | 109.764 | 491.2 | 0.16701
16.782 | 110.265 | 359.45 | 0.16701
16.782 | 110.766 | 258.95 | 0.16701
16.782 | 111.263 | 185.5 | 0.16702
16.782 | 111.766 | 132 | 0.16701
16.782 | 112.265 | 94.8 | 0.16701
16.782 | 112.763 | 69 | 0.16702
16.782 | 113.265 | 50.9 | 0.16701
16.782 | 113.763 | 38.35 | 0.16701
16.782 | 114.264 | 29.35 | 0.16701
16.782 | 114.765 | 22.85 | 0.16701
16.782 | 115.262 | 18.1 | 0.16702
16.782 | 115.765 | 14.5 | 0.16702
16.782 | 116.762 | 9.7 | 0.16701
16.782 | 117.763 | 6.75 | 0.16701
16.782 | 118.764 | 4.95 | 0.16702
16.782 | 119.765 | 3.75 | 0.16701
16.782 | 120.762 | 2.9 | 0.16701
16.782 | 121.763 | 2.35 | 0.16701
16.782 | 122.765 | 1.95 | 0.16702
16.782 | 123.765 | 1.65 | 0.16702
16.782 | 124.762 | 1.4 | 0.16702
16.782 | 125.763 | 1.25 | 0.16701
16.782 | 126.765 | 1.1 | 0.16702
16.782 | 127.765 | 1 | 0.16702
16.782 | 128.762 | 0.95 | 0.16701
16.782 | 129.763 | 0.85 | 0.16702
16.782 | 130.765 | 0.8 | 0.16701
16.782 | 131.766 | 0.75 | 0.16702
16.782 | 132.763 | 0.7 | 0.16702
16.782 | 133.764 | 0.7 | 0.16702
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:16 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=4A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 4.0102
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.7 | 0.13367
16.782 | 74.769 | 0.7 | 0.13367
16.782 | 75.768 | 0.75 | 0.13367
16.782 | 76.767 | 0.8 | 0.13367
16.782 | 77.766 | 0.8 | 0.13367
16.782 | 78.769 | 0.85 | 0.13367
16.782 | 79.769 | 0.95 | 0.13367
16.782 | 80.765 | 1 | 0.13367
16.782 | 81.766 | 1.15 | 0.13367
16.782 | 82.769 | 1.25 | 0.13367
16.782 | 83.769 | 1.45 | 0.13367
16.782 | 84.767 | 1.7 | 0.13367
16.782 | 85.766 | 2 | 0.13367
16.782 | 86.769 | 2.5 | 0.13367
16.782 | 87.768 | 3.2 | 0.13367
16.782 | 88.766 | 4.15 | 0.13367
16.782 | 89.767 | 5.65 | 0.13367
16.782 | 90.767 | 8.05 | 0.13367
16.782 | 91.768 | 11.95 | 0.13367
16.782 | 92.267 | 14.8 | 0.13367
16.782 | 92.765 | 18.75 | 0.13367
16.782 | 93.267 | 24 | 0.13367
16.782 | 93.765 | 31.3 | 0.13367
16.782 | 94.267 | 41.55 | 0.13367
16.782 | 94.767 | 56.15 | 0.13367
16.782 | 95.265 | 76.95 | 0.13367
16.782 | 95.766 | 107.15 | 0.13367
16.782 | 96.266 | 150.05 | 0.13367
16.782 | 96.765 | 209.75 | 0.13367
16.782 | 97.268 | 291.3 | 0.13367
16.782 | 97.765 | 397.25 | 0.13367
16.782 | 98.266 | 523.95 | 0.13367
16.782 | 98.767 | 640.7 | 0.13367
16.782 | 99.265 | 711.95 | 0.13367
16.782 | 99.767 | 741.85 | 0.13367
16.782 | 100.266 | 751.9 | 0.13367
16.782 | 100.765 | 755 | 0.13367
16.782 | 101.267 | 755.9 | 0.13367
16.782 | 101.765 | 756.2 | 0.13367
16.782 | 102.267 | 756.25 | 0.13367
16.782 | 102.766 | 756.25 | 0.13367
16.782 | 103.264 | 756.2 | 0.13367
16.782 | 103.767 | 756.15 | 0.13367
16.782 | 104.266 | 756.15 | 0.13367
16.782 | 104.764 | 756.2 | 0.13367
16.782 | 105.267 | 756.2 | 0.13367
16.782 | 105.765 | 756.15 | 0.13367
16.782 | 106.265 | 755.9 | 0.13367
16.782 | 106.767 | 754.9 | 0.13367
16.782 | 107.264 | 751.75 | 0.13367
16.782 | 107.767 | 741.15 | 0.13367
16.782 | 108.265 | 710.45 | 0.13367
16.782 | 108.763 | 637.9 | 0.13367
16.782 | 109.267 | 519.6 | 0.13367
16.782 | 109.764 | 394.1 | 0.13367
16.782 | 110.265 | 288.4 | 0.13367
16.782 | 110.766 | 207.95 | 0.13367
16.782 | 111.263 | 149 | 0.13367
16.782 | 111.766 | 106.1 | 0.13367
16.782 | 112.265 | 76.25 | 0.13367
16.782 | 112.762 | 55.55 | 0.13367
16.782 | 113.265 | 41.05 | 0.13367
16.782 | 113.762 | 30.95 | 0.13367
16.782 | 114.264 | 23.7 | 0.13367
16.782 | 114.765 | 18.45 | 0.13367
16.782 | 115.262 | 14.6 | 0.13367
16.782 | 115.765 | 11.75 | 0.13367
16.782 | 116.762 | 7.9 | 0.13367
16.782 | 117.763 | 5.55 | 0.13367
16.782 | 118.764 | 4.1 | 0.13367
16.782 | 119.764 | 3.15 | 0.13367
16.782 | 120.762 | 2.45 | 0.13367
16.782 | 121.763 | 2 | 0.13367
16.782 | 122.765 | 1.7 | 0.13367
16.782 | 123.765 | 1.5 | 0.13368
16.782 | 124.762 | 1.25 | 0.13367
16.782 | 125.763 | 1.1 | 0.13367
16.782 | 126.765 | 1 | 0.13367
16.782 | 127.765 | 0.9 | 0.13367
16.782 | 128.762 | 0.85 | 0.13367
16.782 | 129.763 | 0.8 | 0.13367
16.782 | 130.765 | 0.75 | 0.13367
16.782 | 131.766 | 0.7 | 0.13367
16.782 | 132.762 | 0.65 | 0.13367
16.782 | 133.764 | 0.6 | 0.13367
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:24 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=3A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 3.0094
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.6 | 0.10031
16.782 | 74.769 | 0.65 | 0.10031
16.782 | 75.768 | 0.65 | 0.10031
16.782 | 76.767 | 0.75 | 0.10032
16.782 | 77.766 | 0.75 | 0.10031
16.782 | 78.769 | 0.8 | 0.10031
16.782 | 79.769 | 0.8 | 0.10031
16.782 | 80.766 | 0.9 | 0.10031
16.782 | 81.766 | 1 | 0.10031
16.782 | 82.769 | 1.1 | 0.10032
16.782 | 83.769 | 1.25 | 0.10031
16.782 | 84.766 | 1.45 | 0.10031
16.782 | 85.766 | 1.7 | 0.10031
16.782 | 86.769 | 2.05 | 0.10031
16.782 | 87.768 | 2.55 | 0.10031
16.782 | 88.766 | 3.3 | 0.10031
16.782 | 89.766 | 4.45 | 0.10032
16.782 | 90.767 | 6.2 | 0.10031
16.782 | 91.768 | 9.2 | 0.10031
16.782 | 92.267 | 11.35 | 0.10031
16.782 | 92.765 | 14.3 | 0.10031
16.782 | 93.267 | 18.3 | 0.10031
16.782 | 93.765 | 23.8 | 0.10032
16.782 | 94.267 | 31.5 | 0.10031
16.782 | 94.768 | 42.5 | 0.10032
16.782 | 95.264 | 58.15 | 0.10031
16.782 | 95.766 | 80.9 | 0.10032
16.782 | 96.265 | 113.15 | 0.10031
16.782 | 96.765 | 158.05 | 0.10032
16.782 | 97.267 | 219.35 | 0.10032
16.782 | 97.765 | 299.1 | 0.10032
16.782 | 98.266 | 394.55 | 0.10032
16.782 | 98.767 | 482.45 | 0.10031
16.782 | 99.265 | 536.05 | 0.10032
16.782 | 99.767 | 558.5 | 0.10031
16.782 | 100.266 | 566 | 0.10031
16.782 | 100.765 | 568.35 | 0.10031
16.782 | 101.267 | 569 | 0.10031
16.782 | 101.765 | 569.2 | 0.10031
16.782 | 102.266 | 569.25 | 0.10031
16.782 | 102.766 | 569.25 | 0.10031
16.782 | 103.264 | 569.25 | 0.10031
16.782 | 103.766 | 569.25 | 0.10032
16.782 | 104.266 | 569.3 | 0.10031
16.782 | 104.764 | 569.3 | 0.10031
16.782 | 105.267 | 569.3 | 0.10032
16.782 | 105.764 | 569.25 | 0.10031
16.782 | 106.265 | 569.05 | 0.10031
16.782 | 106.766 | 568.35 | 0.10031
16.782 | 107.264 | 565.95 | 0.10032
16.782 | 107.767 | 558 | 0.10032
16.782 | 108.265 | 534.9 | 0.10031
16.782 | 108.763 | 480.3 | 0.10032
16.782 | 109.267 | 391.15 | 0.10032
16.782 | 109.764 | 296.65 | 0.10031
16.782 | 110.265 | 217.15 | 0.10031
16.782 | 110.766 | 156.6 | 0.10032
16.782 | 111.263 | 112.3 | 0.10031
16.782 | 111.766 | 80 | 0.10031
16.782 | 112.265 | 57.55 | 0.10031
16.782 | 112.763 | 42 | 0.10031
16.782 | 113.265 | 31.1 | 0.10032
16.782 | 113.763 | 23.5 | 0.10031
16.782 | 114.264 | 18.05 | 0.10032
16.782 | 114.764 | 14.15 | 0.10031
16.782 | 115.262 | 11.25 | 0.10031
16.782 | 115.765 | 9.05 | 0.10032
16.782 | 116.762 | 6.15 | 0.10031
16.782 | 117.763 | 4.4 | 0.10031
16.782 | 118.764 | 3.25 | 0.10031
16.782 | 119.765 | 2.5 | 0.10031
16.782 | 120.762 | 2 | 0.10031
16.782 | 121.763 | 1.65 | 0.10032
16.782 | 122.765 | 1.4 | 0.10031
16.782 | 123.765 | 1.25 | 0.10031
16.782 | 124.762 | 1.1 | 0.10031
16.782 | 125.763 | 1 | 0.10031
16.782 | 126.765 | 0.9 | 0.10032
16.782 | 127.765 | 0.85 | 0.10032
16.782 | 128.762 | 0.75 | 0.10032
16.782 | 129.764 | 0.75 | 0.10031
16.782 | 130.765 | 0.65 | 0.10031
16.782 | 131.766 | 0.7 | 0.10031
16.782 | 132.762 | 0.6 | 0.10032
16.782 | 133.764 | 0.6 | 0.10032
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:34 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=2A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 2.0091
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.6 | 0.066971
16.782 | 74.769 | 0.6 | 0.066968
16.782 | 75.768 | 0.6 | 0.06697
16.782 | 76.766 | 0.65 | 0.066971
16.782 | 77.766 | 0.65 | 0.06697
16.782 | 78.768 | 0.7 | 0.066973
16.782 | 79.769 | 0.7 | 0.066969
16.782 | 80.766 | 0.8 | 0.066969
16.782 | 81.766 | 0.85 | 0.066967
16.782 | 82.768 | 0.9 | 0.066971
16.782 | 83.769 | 1 | 0.066973
16.782 | 84.767 | 1.15 | 0.066975
16.782 | 85.766 | 1.35 | 0.066968
16.782 | 86.769 | 1.6 | 0.066968
16.782 | 87.767 | 1.9 | 0.066968
16.782 | 88.767 | 2.4 | 0.066968
16.782 | 89.766 | 3.2 | 0.066969
16.782 | 90.767 | 4.45 | 0.066968
16.782 | 91.768 | 6.4 | 0.066971
16.782 | 92.267 | 7.9 | 0.066969
16.782 | 92.765 | 9.9 | 0.06697
16.782 | 93.267 | 12.6 | 0.06697
16.782 | 93.765 | 16.25 | 0.06697
16.782 | 94.267 | 21.45 | 0.066969
16.782 | 94.768 | 28.8 | 0.066971
16.782 | 95.264 | 39.3 | 0.066971
16.782 | 95.767 | 54.6 | 0.066969
16.782 | 96.266 | 76.2 | 0.066967
16.782 | 96.765 | 106.2 | 0.066971
16.782 | 97.268 | 147.45 | 0.066967
16.782 | 97.765 | 200.95 | 0.066969
16.782 | 98.266 | 264.95 | 0.066971
16.782 | 98.767 | 324 | 0.066966
16.782 | 99.265 | 359.95 | 0.06697
16.782 | 99.767 | 375.05 | 0.066972
16.782 | 100.266 | 380.05 | 0.06697
16.782 | 100.765 | 381.65 | 0.066969
16.782 | 101.267 | 382.1 | 0.066969
16.782 | 101.765 | 382.2 | 0.066969
16.782 | 102.266 | 382.25 | 0.06697
16.782 | 102.766 | 382.2 | 0.06697
16.782 | 103.264 | 382.25 | 0.066967
16.782 | 103.767 | 382.25 | 0.066971
16.782 | 104.266 | 382.25 | 0.066972
16.782 | 104.764 | 382.25 | 0.066973
16.782 | 105.267 | 382.25 | 0.066968
16.782 | 105.764 | 382.2 | 0.06697
16.782 | 106.265 | 382.1 | 0.066968
16.782 | 106.767 | 381.65 | 0.066972
16.782 | 107.264 | 380 | 0.06697
16.782 | 107.767 | 374.7 | 0.066971
16.782 | 108.265 | 359.2 | 0.066969
16.782 | 108.763 | 322.5 | 0.066972
16.782 | 109.267 | 262.75 | 0.066971
16.782 | 109.764 | 199.35 | 0.06697
16.782 | 110.265 | 145.95 | 0.06697
16.782 | 110.766 | 105.3 | 0.066971
16.782 | 111.263 | 75.55 | 0.06697
16.782 | 111.766 | 53.95 | 0.06697
16.782 | 112.265 | 38.9 | 0.066971
16.782 | 112.763 | 28.4 | 0.066972
16.782 | 113.265 | 21.15 | 0.066971
16.782 | 113.762 | 16 | 0.06697
16.782 | 114.263 | 12.3 | 0.066973
16.782 | 114.765 | 9.7 | 0.06697
16.782 | 115.262 | 7.75 | 0.066969
16.782 | 115.765 | 6.3 | 0.066972
16.782 | 116.762 | 4.35 | 0.066971
16.782 | 117.763 | 3.15 | 0.066974
16.782 | 118.764 | 2.35 | 0.066972
16.782 | 119.764 | 1.85 | 0.06697
16.782 | 120.762 | 1.5 | 0.066972
16.782 | 121.763 | 1.3 | 0.06697
16.782 | 122.765 | 1.15 | 0.066971
16.782 | 123.765 | 1 | 0.066971
16.782 | 124.762 | 0.9 | 0.066971
16.782 | 125.763 | 0.8 | 0.066971
16.782 | 126.765 | 0.75 | 0.066971
16.782 | 127.765 | 0.7 | 0.066969
16.782 | 128.762 | 0.65 | 0.066971
16.782 | 129.763 | 0.6 | 0.06697
16.782 | 130.765 | 0.6 | 0.066974
16.782 | 131.766 | 0.55 | 0.066971
16.782 | 132.763 | 0.55 | 0.066972
16.782 | 133.764 | 0.55 | 0.066969
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:43 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=1A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 1.0092
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.6 | 0.033642
16.782 | 74.769 | 0.55 | 0.03364
16.782 | 75.768 | 0.6 | 0.033641
16.782 | 76.767 | 0.65 | 0.03364
16.782 | 77.766 | 0.6 | 0.03364
16.782 | 78.768 | 0.65 | 0.033636
16.782 | 79.769 | 0.65 | 0.033637
16.782 | 80.765 | 0.65 | 0.033638
16.782 | 81.766 | 0.7 | 0.033641
16.782 | 82.768 | 0.7 | 0.033643
16.782 | 83.769 | 0.8 | 0.033641
16.782 | 84.766 | 0.9 | 0.033639
16.782 | 85.766 | 0.95 | 0.033639
16.782 | 86.769 | 1.1 | 0.033641
16.782 | 87.768 | 1.3 | 0.033643
16.782 | 88.766 | 1.55 | 0.033643
16.782 | 89.766 | 2 | 0.033642
16.782 | 90.767 | 2.6 | 0.03364
16.782 | 91.768 | 3.65 | 0.033643
16.782 | 92.267 | 4.4 | 0.033641
16.782 | 92.765 | 5.45 | 0.033642
16.782 | 93.267 | 6.8 | 0.033641
16.782 | 93.765 | 8.7 | 0.03364
16.782 | 94.267 | 11.35 | 0.033642
16.782 | 94.767 | 15.15 | 0.03364
16.782 | 95.265 | 20.5 | 0.03364
16.782 | 95.766 | 28.25 | 0.033641
16.782 | 96.265 | 39.25 | 0.033639
16.782 | 96.765 | 54.6 | 0.033641
16.782 | 97.268 | 75.55 | 0.03364
16.782 | 97.765 | 102.8 | 0.03364
16.782 | 98.266 | 135.45 | 0.033642
16.782 | 98.766 | 165.55 | 0.033641
16.782 | 99.265 | 183.95 | 0.033641
16.782 | 99.767 | 191.6 | 0.033639
16.782 | 100.266 | 194.15 | 0.033643
16.782 | 100.765 | 194.95 | 0.033641
16.782 | 101.266 | 195.2 | 0.033641
16.782 | 101.765 | 195.25 | 0.033639
16.782 | 102.266 | 195.25 | 0.033643
16.782 | 102.766 | 195.2 | 0.03364
16.782 | 103.264 | 195.15 | 0.033641
16.782 | 103.766 | 195.15 | 0.033643
16.782 | 104.266 | 195.1 | 0.033639
16.782 | 104.764 | 195.15 | 0.033641
16.782 | 105.267 | 195.2 | 0.033642
16.782 | 105.764 | 195.15 | 0.033641
16.782 | 106.265 | 195.15 | 0.03364
16.782 | 106.766 | 194.85 | 0.03364
16.782 | 107.264 | 194.1 | 0.033638
16.782 | 107.767 | 191.4 | 0.033642
16.782 | 108.265 | 183.45 | 0.03364
16.782 | 108.763 | 164.8 | 0.033643
16.782 | 109.267 | 134.25 | 0.033642
16.782 | 109.764 | 101.9 | 0.033641
16.782 | 110.265 | 74.75 | 0.03364
16.782 | 110.766 | 54 | 0.03364
16.782 | 111.263 | 38.9 | 0.03364
16.782 | 111.766 | 27.9 | 0.033638
16.782 | 112.265 | 20.2 | 0.033642
16.782 | 112.763 | 14.9 | 0.033643
16.782 | 113.265 | 11.15 | 0.033644
16.782 | 113.762 | 8.55 | 0.033641
16.782 | 114.264 | 6.7 | 0.033639
16.782 | 114.764 | 5.35 | 0.033642
16.782 | 115.262 | 4.3 | 0.03364
16.782 | 115.765 | 3.55 | 0.033641
16.782 | 116.762 | 2.55 | 0.03364
16.782 | 117.763 | 1.9 | 0.033643
16.782 | 118.764 | 1.55 | 0.033643
16.782 | 119.765 | 1.25 | 0.033642
16.782 | 120.762 | 1.05 | 0.033642
16.782 | 121.763 | 0.95 | 0.033644
16.782 | 122.764 | 0.85 | 0.033642
16.782 | 123.765 | 0.8 | 0.03364
16.782 | 124.762 | 0.75 | 0.033647
16.782 | 125.763 | 0.7 | 0.033644
16.782 | 126.765 | 0.65 | 0.033644
16.782 | 127.764 | 0.65 | 0.033644
16.782 | 128.762 | 0.6 | 0.033643
16.782 | 129.763 | 0.55 | 0.033642
16.782 | 130.765 | 0.55 | 0.033639
16.782 | 131.766 | 0.55 | 0.033644
16.782 | 132.762 | 0.55 | 0.033641
16.782 | 133.764 | 0.55 | 0.033641
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
11:53 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=0A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | 0.003348
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.5 | 0.000111
16.782 | 74.769 | 0.5 | 0.000111
16.782 | 75.768 | 0.5 | 0.000112
16.782 | 76.766 | 0.5 | 0.000114
16.782 | 77.766 | 0.5 | 0.000112
16.782 | 78.767 | 0.5 | 0.000111
16.782 | 79.769 | 0.55 | 0.000112
16.782 | 80.765 | 0.55 | 0.000112
16.782 | 81.765 | 0.55 | 0.000111
16.782 | 82.768 | 0.55 | 0.000111
16.782 | 83.769 | 0.55 | 0.000113
16.782 | 84.766 | 0.6 | 0.000111
16.782 | 85.766 | 0.6 | 0.000112
16.782 | 86.768 | 0.6 | 0.000113
16.782 | 87.768 | 0.65 | 0.000113
16.782 | 88.766 | 0.7 | 0.000113
16.782 | 89.766 | 0.7 | 0.000112
16.782 | 90.766 | 0.8 | 0.000112
16.782 | 91.768 | 0.9 | 0.000112
16.782 | 92.267 | 0.9 | 0.000112
16.782 | 92.765 | 1 | 0.000113
16.782 | 93.267 | 1.05 | 0.000112
16.782 | 93.765 | 1.2 | 0.000111
16.782 | 94.266 | 1.3 | 0.000112
16.782 | 94.767 | 1.45 | 0.000112
16.782 | 95.264 | 1.65 | 0.000111
16.782 | 95.766 | 1.9 | 0.000111
16.782 | 96.266 | 2.25 | 0.000112
16.782 | 96.765 | 2.7 | 0.000112
16.782 | 97.267 | 3.4 | 0.000113
16.782 | 97.765 | 4.35 | 0.000112
16.782 | 98.266 | 5.55 | 0.000111
16.782 | 98.766 | 6.6 | 0.000111
16.782 | 99.265 | 7.25 | 0.000111
16.782 | 99.767 | 7.5 | 0.000112
16.782 | 100.266 | 7.55 | 0.000111
16.782 | 100.764 | 7.5 | 0.000111
16.782 | 101.267 | 7.45 | 0.000112
16.782 | 101.765 | 7.45 | 0.000112
16.782 | 102.266 | 7.45 | 0.000111
16.782 | 102.766 | 7.4 | 0.000112
16.782 | 103.264 | 7.4 | 0.000112
16.782 | 103.766 | 7.4 | 0.000112
16.782 | 104.266 | 7.4 | 0.000111
16.782 | 104.764 | 7.4 | 0.000112
16.782 | 105.267 | 7.4 | 0.000113
16.782 | 105.764 | 7.45 | 0.000111
16.782 | 106.265 | 7.45 | 0.000112
16.782 | 106.766 | 7.5 | 0.000112
16.782 | 107.263 | 7.45 | 0.00011
16.782 | 107.767 | 7.4 | 0.000111
16.782 | 108.265 | 7.2 | 0.000111
16.782 | 108.763 | 6.5 | 0.000112
16.782 | 109.267 | 5.35 | 0.000112
16.782 | 109.764 | 4.2 | 0.000111
16.782 | 110.265 | 3.25 | 0.000111
16.782 | 110.766 | 2.55 | 0.000111
16.782 | 111.263 | 2.15 | 0.000111
16.782 | 111.766 | 1.75 | 0.000111
16.782 | 112.265 | 1.5 | 0.000111
16.782 | 112.763 | 1.35 | 0.000112
16.782 | 113.265 | 1.2 | 0.000111
16.782 | 113.762 | 1.1 | 0.00011
16.782 | 114.263 | 1 | 0.000111
16.782 | 114.765 | 0.9 | 0.000111
16.782 | 115.262 | 0.85 | 0.000112
16.782 | 115.764 | 0.8 | 0.000111
16.782 | 116.762 | 0.8 | 0.00011
16.782 | 117.762 | 0.7 | 0.000111
16.782 | 118.763 | 0.65 | 0.000112
16.782 | 119.764 | 0.6 | 0.000112
16.782 | 120.762 | 0.55 | 0.000111
16.782 | 121.763 | 0.55 | 0.000111
16.782 | 122.764 | 0.55 | 0.000111
16.782 | 123.764 | 0.5 | 0.000112
16.782 | 124.762 | 0.5 | 0.000112
16.782 | 125.763 | 0.5 | 0.000111
16.782 | 126.765 | 0.5 | 0.000111
16.782 | 127.764 | 0.5 | 0.000111
16.782 | 128.762 | 0.5 | 0.000112
16.782 | 129.763 | 0.5 | 0.000112
16.782 | 130.765 | 0.45 | 0.000112
16.782 | 131.766 | 0.5 | 0.000113
16.782 | 132.762 | 0.45 | 0.000111
16.782 | 133.764 | 0.5 | 0.000111
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
12:03 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-1A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -0.99049
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.769 | 0.5 | -0.033012
16.782 | 74.768 | 0.5 | -0.033015
16.782 | 75.768 | 0.45 | -0.033013
16.782 | 76.767 | 0.5 | -0.033016
16.782 | 77.766 | 0.45 | -0.033014
16.782 | 78.768 | 0.45 | -0.033012
16.782 | 79.769 | 0.45 | -0.033016
16.782 | 80.765 | 0.45 | -0.033014
16.782 | 81.767 | 0.4 | -0.033017
16.782 | 82.768 | 0.4 | -0.033014
16.782 | 83.769 | 0.35 | -0.033013
16.782 | 84.767 | 0.3 | -0.033014
16.782 | 85.765 | 0.25 | -0.033015
16.782 | 86.769 | 0.1 | -0.033014
16.782 | 87.768 | 0 | -0.033016
16.782 | 88.766 | -0.2 | -0.033017
16.782 | 89.766 | -0.55 | -0.033015
16.782 | 90.767 | -1.05 | -0.033016
16.782 | 91.768 | -1.95 | -0.033014
16.782 | 92.267 | -2.6 | -0.033014
16.782 | 92.765 | -3.45 | -0.033014
16.782 | 93.267 | -4.7 | -0.033015
16.782 | 93.765 | -6.4 | -0.033017
16.782 | 94.267 | -8.85 | -0.033014
16.782 | 94.767 | -12.3 | -0.033016
16.782 | 95.264 | -17.2 | -0.033017
16.782 | 95.767 | -24.4 | -0.033016
16.782 | 96.265 | -34.65 | -0.033015
16.782 | 96.765 | -48.95 | -0.033018
16.782 | 97.268 | -68.4 | -0.033016
16.782 | 97.765 | -93.65 | -0.033017
16.782 | 98.265 | -123.85 | -0.033017
16.782 | 98.766 | -151.6 | -0.033015
16.782 | 99.265 | -168.6 | -0.033017
16.782 | 99.767 | -175.75 | -0.033017
16.782 | 100.266 | -178.2 | -0.03302
16.782 | 100.765 | -179 | -0.033016
16.782 | 101.267 | -179.25 | -0.033016
16.782 | 101.765 | -179.35 | -0.03302
16.782 | 102.266 | -179.4 | -0.03302
16.782 | 102.766 | -179.4 | -0.033019
16.782 | 103.264 | -179.4 | -0.03302
16.782 | 103.766 | -179.45 | -0.03302
16.782 | 104.266 | -179.4 | -0.033018
16.782 | 104.763 | -179.4 | -0.033018
16.782 | 105.267 | -179.4 | -0.03302
16.782 | 105.764 | -179.35 | -0.03302
16.782 | 106.265 | -179.3 | -0.033016
16.782 | 106.766 | -179.05 | -0.033019
16.782 | 107.263 | -178.25 | -0.033018
16.782 | 107.766 | -175.65 | -0.033018
16.782 | 108.265 | -168.3 | -0.033015
16.782 | 108.763 | -151.05 | -0.033013
16.782 | 109.267 | -122.9 | -0.033016
16.782 | 109.764 | -93.05 | -0.033016
16.782 | 110.265 | -67.85 | -0.033017
16.782 | 110.766 | -48.6 | -0.033018
16.782 | 111.262 | -34.55 | -0.033015
16.782 | 111.766 | -24.3 | -0.033015
16.782 | 112.265 | -17.2 | -0.033016
16.782 | 112.763 | -12.25 | -0.033014
16.782 | 113.265 | -8.8 | -0.033015
16.782 | 113.762 | -6.45 | -0.033018
16.782 | 114.263 | -4.75 | -0.033015
16.782 | 114.764 | -3.5 | -0.033018
16.782 | 115.262 | -2.65 | -0.033014
16.782 | 115.764 | -2 | -0.033015
16.782 | 116.762 | -1.1 | -0.033016
16.782 | 117.762 | -0.6 | -0.033017
16.782 | 118.764 | -0.3 | -0.033016
16.782 | 119.764 | -0.05 | -0.033017
16.782 | 120.762 | 0.1 | -0.033017
16.782 | 121.763 | 0.2 | -0.033015
16.782 | 122.764 | 0.25 | -0.033015
16.782 | 123.764 | 0.35 | -0.033016
16.782 | 124.762 | 0.35 | -0.033017
16.782 | 125.763 | 0.4 | -0.033017
16.782 | 126.765 | 0.45 | -0.03302
16.782 | 127.764 | 0.45 | -0.033016
16.782 | 128.762 | 0.45 | -0.033018
16.782 | 129.763 | 0.4 | -0.033015
16.782 | 130.765 | 0.45 | -0.033018
16.782 | 131.765 | 0.45 | -0.033017
16.782 | 132.762 | 0.45 | -0.033017
16.782 | 133.764 | 0.4 | -0.033017
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
12:26 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-2A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -1.9895
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.768 | 0.45 | -0.066319
16.782 | 74.768 | 0.45 | -0.066319
16.782 | 75.768 | 0.45 | -0.066316
16.782 | 76.766 | 0.4 | -0.066313
16.782 | 77.766 | 0.35 | -0.066318
16.782 | 78.768 | 0.4 | -0.06632
16.782 | 79.769 | 0.35 | -0.066319
16.782 | 80.765 | 0.35 | -0.066317
16.782 | 81.765 | 0.25 | -0.066318
16.782 | 82.768 | 0.2 | -0.066319
16.782 | 83.768 | 0.15 | -0.066318
16.782 | 84.766 | 0 | -0.066318
16.782 | 85.765 | -0.15 | -0.066318
16.782 | 86.768 | -0.35 | -0.066316
16.782 | 87.768 | -0.65 | -0.066317
16.782 | 88.765 | -1.05 | -0.066315
16.782 | 89.766 | -1.8 | -0.066315
16.782 | 90.767 | -2.9 | -0.066314
16.782 | 91.768 | -4.7 | -0.066317
16.782 | 92.267 | -6.1 | -0.066318
16.782 | 92.765 | -7.95 | -0.066312
16.782 | 93.266 | -10.5 | -0.066315
16.782 | 93.765 | -13.95 | -0.066314
16.782 | 94.267 | -18.9 | -0.066316
16.782 | 94.767 | -25.95 | -0.066315
16.782 | 95.264 | -36.05 | -0.066314
16.782 | 95.766 | -50.75 | -0.066314
16.782 | 96.265 | -71.6 | -0.066318
16.782 | 96.764 | -100.65 | -0.066318
16.782 | 97.267 | -140.35 | -0.066316
16.782 | 97.765 | -191.95 | -0.066314
16.782 | 98.265 | -253.45 | -0.066319
16.782 | 98.766 | -310.15 | -0.066315
16.782 | 99.264 | -344.85 | -0.066315
16.782 | 99.767 | -359.45 | -0.066316
16.782 | 100.265 | -364.4 | -0.066316
16.782 | 100.764 | -366 | -0.066316
16.782 | 101.266 | -366.45 | -0.066314
16.782 | 101.765 | -366.6 | -0.066315
16.782 | 102.265 | -366.65 | -0.066317
16.782 | 102.766 | -366.65 | -0.066317
16.782 | 103.263 | -366.65 | -0.066318
16.782 | 103.765 | -366.65 | -0.066317
16.782 | 104.265 | -366.7 | -0.066317
16.782 | 104.763 | -366.7 | -0.066315
16.782 | 105.266 | -366.65 | -0.066314
16.782 | 105.764 | -366.65 | -0.066314
16.782 | 106.265 | -366.5 | -0.066318
16.782 | 106.766 | -366 | -0.066316
16.782 | 107.263 | -364.35 | -0.066317
16.782 | 107.766 | -359.2 | -0.066319
16.782 | 108.265 | -344.15 | -0.066316
16.782 | 108.762 | -308.95 | -0.066319
16.782 | 109.266 | -251.45 | -0.06632
16.782 | 109.763 | -190.45 | -0.066319
16.782 | 110.264 | -139.05 | -0.066318
16.782 | 110.765 | -99.85 | -0.066319
16.782 | 111.262 | -71.2 | -0.066319
16.782 | 111.766 | -50.35 | -0.066317
16.782 | 112.265 | -35.9 | -0.066315
16.782 | 112.762 | -25.8 | -0.066314
16.782 | 113.265 | -18.8 | -0.066315
16.782 | 113.762 | -13.95 | -0.066317
16.782 | 114.263 | -10.45 | -0.066317
16.782 | 114.764 | -7.95 | -0.066319
16.782 | 115.262 | -6.15 | -0.066318
16.782 | 115.764 | -4.75 | -0.066317
16.782 | 116.761 | -2.9 | -0.066316
16.782 | 117.762 | -1.85 | -0.066318
16.782 | 118.763 | -1.1 | -0.066315
16.782 | 119.764 | -0.7 | -0.066314
16.782 | 120.762 | -0.35 | -0.066317
16.782 | 121.762 | -0.15 | -0.066318
16.782 | 122.764 | 0 | -0.066318
16.782 | 123.764 | 0.1 | -0.066316
16.782 | 124.762 | 0.2 | -0.066315
16.782 | 125.762 | 0.25 | -0.066315
16.782 | 126.765 | 0.3 | -0.066318
16.782 | 127.764 | 0.35 | -0.066317
16.782 | 128.762 | 0.35 | -0.066318
16.782 | 129.763 | 0.4 | -0.066317
16.782 | 130.764 | 0.4 | -0.066318
16.782 | 131.765 | 0.4 | -0.066314
16.782 | 132.762 | 0.4 | -0.066321
16.782 | 133.764 | 0.4 | -0.066317
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
12:34 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-3A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -2.9892
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.768 | 0.4 | -0.099638
16.782 | 74.768 | 0.4 | -0.099638
16.782 | 75.767 | 0.4 | -0.099641
16.782 | 76.766 | 0.35 | -0.099639
16.782 | 77.766 | 0.3 | -0.099643
16.782 | 78.768 | 0.25 | -0.099642
16.782 | 79.768 | 0.25 | -0.099641
16.782 | 80.765 | 0.2 | -0.099639
16.782 | 81.766 | 0.15 | -0.09964
16.782 | 82.767 | 0.05 | -0.099639
16.782 | 83.769 | -0.05 | -0.099637
16.782 | 84.766 | -0.25 | -0.099641
16.782 | 85.765 | -0.5 | -0.099637
16.782 | 86.768 | -0.85 | -0.099637
16.782 | 87.767 | -1.25 | -0.099639
16.782 | 88.765 | -2 | -0.099638
16.782 | 89.766 | -3.05 | -0.099638
16.782 | 90.766 | -4.7 | -0.099639
16.782 | 91.768 | -7.5 | -0.099637
16.782 | 92.266 | -9.6 | -0.099639
16.782 | 92.765 | -12.4 | -0.099638
16.782 | 93.267 | -16.25 | -0.099638
16.782 | 93.764 | -21.55 | -0.099638
16.782 | 94.266 | -29 | -0.099637
16.782 | 94.767 | -39.65 | -0.099641
16.782 | 95.264 | -54.9 | -0.099639
16.782 | 95.766 | -77.05 | -0.099642
16.782 | 96.265 | -108.55 | -0.09964
16.782 | 96.764 | -152.45 | -0.099641
16.782 | 97.267 | -212.35 | -0.099641
16.782 | 97.765 | -290.2 | -0.099642
16.782 | 98.265 | -383.1 | -0.099637
16.782 | 98.766 | -468.9 | -0.09964
16.782 | 99.264 | -521.2 | -0.09964
16.782 | 99.767 | -543.3 | -0.099638
16.782 | 100.265 | -550.75 | -0.099641
16.782 | 100.764 | -553.15 | -0.09964
16.782 | 101.266 | -553.9 | -0.099639
16.782 | 101.765 | -554.1 | -0.099638
16.782 | 102.266 | -554.15 | -0.099639
16.782 | 102.766 | -554.1 | -0.09964
16.782 | 103.263 | -554.1 | -0.099638
16.782 | 103.766 | -554.15 | -0.099639
16.782 | 104.265 | -554.15 | -0.099638
16.782 | 104.763 | -554.2 | -0.099641
16.782 | 105.266 | -554.2 | -0.099638
16.782 | 105.764 | -554.15 | -0.099639
16.782 | 106.265 | -553.95 | -0.099638
16.782 | 106.766 | -553.2 | -0.099639
16.782 | 107.263 | -550.7 | -0.099642
16.782 | 107.766 | -542.95 | -0.09964
16.782 | 108.264 | -520.35 | -0.099639
16.782 | 108.762 | -467.05 | -0.099637
16.782 | 109.266 | -380.25 | -0.099639
16.782 | 109.763 | -288.05 | -0.099637
16.782 | 110.264 | -210.45 | -0.099639
16.782 | 110.765 | -151.25 | -0.099639
16.782 | 111.262 | -108 | -0.099638
16.782 | 111.766 | -76.5 | -0.099637
16.782 | 112.264 | -54.6 | -0.099638
16.782 | 112.762 | -39.4 | -0.09964
16.782 | 113.265 | -28.85 | -0.09964
16.782 | 113.762 | -21.45 | -0.099641
16.782 | 114.263 | -16.2 | -0.099641
16.782 | 114.764 | -12.4 | -0.09964
16.782 | 115.262 | -9.65 | -0.09964
16.782 | 115.764 | -7.5 | -0.099637
16.782 | 116.761 | -4.75 | -0.09964
16.782 | 117.762 | -3.05 | -0.099641
16.782 | 118.763 | -2 | -0.099639
16.782 | 119.764 | -1.3 | -0.099637
16.782 | 120.762 | -0.85 | -0.099639
16.782 | 121.762 | -0.5 | -0.099639
16.782 | 122.764 | -0.3 | -0.09964
16.782 | 123.764 | -0.1 | -0.09964
16.782 | 124.762 | 0.05 | -0.099638
16.782 | 125.763 | 0.1 | -0.099639
16.782 | 126.765 | 0.15 | -0.09964
16.782 | 127.764 | 0.2 | -0.099641
16.782 | 128.762 | 0.3 | -0.099642
16.782 | 129.763 | 0.3 | -0.09964
16.782 | 130.764 | 0.3 | -0.099642
16.782 | 131.765 | 0.35 | -0.099637
16.782 | 132.762 | 0.35 | -0.099642
16.782 | 133.764 | 0.35 | -0.099639
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
18:21 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-4A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -3.9938
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.765 | 0.35 | -0.13313
16.782 | 74.765 | 0.35 | -0.13313
16.782 | 75.764 | 0.35 | -0.13313
16.782 | 76.763 | 0.3 | -0.13313
16.782 | 77.762 | 0.25 | -0.13313
16.782 | 78.765 | 0.2 | -0.13313
16.782 | 79.765 | 0.2 | -0.13312
16.782 | 80.762 | 0.1 | -0.13312
16.782 | 81.762 | 0 | -0.13313
16.782 | 82.765 | -0.15 | -0.13313
16.782 | 83.765 | -0.3 | -0.13313
16.782 | 84.763 | -0.55 | -0.13313
16.782 | 85.763 | -0.85 | -0.13313
16.782 | 86.764 | -1.25 | -0.13313
16.782 | 87.765 | -1.95 | -0.13313
16.782 | 88.762 | -2.9 | -0.13313
16.782 | 89.763 | -4.3 | -0.13313
16.782 | 90.764 | -6.55 | -0.13313
16.782 | 91.765 | -10.3 | -0.13313
16.782 | 92.264 | -13.1 | -0.13313
16.782 | 92.762 | -16.85 | -0.13313
16.782 | 93.264 | -22.05 | -0.13313
16.782 | 93.762 | -29.1 | -0.13313
16.782 | 94.264 | -39.1 | -0.13313
16.782 | 94.764 | -53.3 | -0.13313
16.782 | 95.261 | -73.7 | -0.13313
16.782 | 95.764 | -103.45 | -0.13313
16.782 | 96.262 | -145.55 | -0.13313
16.782 | 96.762 | -204.25 | -0.13313
16.782 | 97.264 | -284.5 | -0.13313
16.782 | 97.762 | -388.65 | -0.13313
16.782 | 98.262 | -513.15 | -0.13313
16.782 | 98.763 | -628.1 | -0.13313
16.782 | 99.262 | -698.4 | -0.13313
16.782 | 99.764 | -728 | -0.13313
16.782 | 100.262 | -737.95 | -0.13313
16.782 | 100.762 | -741.15 | -0.13313
16.782 | 101.263 | -742.1 | -0.13313
16.782 | 101.762 | -742.4 | -0.13313
16.782 | 102.263 | -742.45 | -0.13313
16.782 | 102.764 | -742.45 | -0.13313
16.782 | 103.261 | -742.5 | -0.13313
16.782 | 103.763 | -742.5 | -0.13313
16.782 | 104.263 | -742.5 | -0.13313
16.782 | 104.761 | -742.5 | -0.13313
16.782 | 105.264 | -742.5 | -0.13312
16.782 | 105.761 | -742.45 | -0.13312
16.782 | 106.262 | -742.15 | -0.13312
16.782 | 106.763 | -741.15 | -0.13313
16.782 | 107.261 | -737.85 | -0.13313
16.782 | 107.764 | -727.5 | -0.13313
16.782 | 108.262 | -697.15 | -0.13313
16.782 | 108.76 | -625.75 | -0.13312
16.782 | 109.263 | -509.4 | -0.13312
16.782 | 109.761 | -386 | -0.13313
16.782 | 110.262 | -282 | -0.13313
16.782 | 110.763 | -202.8 | -0.13313
16.782 | 111.26 | -144.8 | -0.13312
16.782 | 111.763 | -102.6 | -0.13313
16.782 | 112.262 | -73.3 | -0.13313
16.782 | 112.76 | -52.95 | -0.13313
16.782 | 113.262 | -38.8 | -0.13313
16.782 | 113.759 | -28.95 | -0.13313
16.782 | 114.26 | -21.9 | -0.13313
16.782 | 114.761 | -16.8 | -0.13313
16.782 | 115.259 | -13.1 | -0.13313
16.782 | 115.762 | -10.3 | -0.13313
16.782 | 116.759 | -6.55 | -0.13313
16.782 | 117.76 | -4.3 | -0.13313
16.782 | 118.761 | -2.85 | -0.13313
16.782 | 119.762 | -1.95 | -0.13313
16.782 | 120.759 | -1.3 | -0.13313
16.782 | 121.76 | -0.9 | -0.13313
16.782 | 122.761 | -0.6 | -0.13313
16.782 | 123.762 | -0.35 | -0.13313
16.782 | 124.759 | -0.1 | -0.13313
16.782 | 125.76 | -0.05 | -0.13313
16.782 | 126.762 | 0.05 | -0.13313
16.782 | 127.762 | 0.15 | -0.13313
16.782 | 128.76 | 0.2 | -0.13313
16.782 | 129.76 | 0.25 | -0.13313
16.782 | 130.762 | 0.25 | -0.13312
16.782 | 131.763 | 0.25 | -0.13313
16.782 | 132.76 | 0.3 | -0.13313
16.782 | 133.761 | 0.3 | -0.13313
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
18:29 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-5A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -4.9927
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.765 | 0.35 | -0.16642
16.782 | 74.765 | 0.3 | -0.16642
16.782 | 75.764 | 0.3 | -0.16642
16.782 | 76.763 | 0.25 | -0.16642
16.782 | 77.762 | 0.2 | -0.16643
16.782 | 78.765 | 0.15 | -0.16643
16.782 | 79.765 | 0.1 | -0.16643
16.782 | 80.762 | -0.05 | -0.16643
16.782 | 81.762 | -0.1 | -0.16642
16.782 | 82.765 | -0.3 | -0.16642
16.782 | 83.766 | -0.55 | -0.16642
16.782 | 84.763 | -0.85 | -0.16642
16.782 | 85.762 | -1.2 | -0.16642
16.782 | 86.765 | -1.75 | -0.16643
16.782 | 87.765 | -2.55 | -0.16642
16.782 | 88.762 | -3.7 | -0.16642
16.782 | 89.763 | -5.5 | -0.16642
16.782 | 90.764 | -8.35 | -0.16643
16.782 | 91.765 | -13.05 | -0.16642
16.782 | 92.264 | -16.6 | -0.16642
16.782 | 92.762 | -21.3 | -0.16642
16.782 | 93.264 | -27.75 | -0.16643
16.782 | 93.762 | -36.6 | -0.16643
16.782 | 94.263 | -49.15 | -0.16642
16.782 | 94.763 | -66.95 | -0.16642
16.782 | 95.262 | -92.55 | -0.16643
16.782 | 95.763 | -129.7 | -0.16642
16.782 | 96.262 | -182.35 | -0.16642
16.782 | 96.762 | -255.85 | -0.16642
16.782 | 97.264 | -356.05 | -0.16642
16.782 | 97.762 | -486.7 | -0.16642
16.782 | 98.263 | -642.4 | -0.16642
16.782 | 98.763 | -786.25 | -0.16643
16.782 | 99.262 | -874.3 | -0.16642
16.782 | 99.764 | -911.35 | -0.16642
16.782 | 100.262 | -923.95 | -0.16642
16.782 | 100.762 | -927.95 | -0.16642
16.782 | 101.264 | -929.1 | -0.16642
16.782 | 101.762 | -929.45 | -0.16642
16.782 | 102.263 | -929.55 | -0.16642
16.782 | 102.763 | -929.6 | -0.16642
16.782 | 103.261 | -929.6 | -0.16643
16.782 | 103.763 | -929.65 | -0.16642
16.782 | 104.263 | -929.65 | -0.16642
16.782 | 104.761 | -929.65 | -0.16643
16.782 | 105.264 | -929.6 | -0.16643
16.782 | 105.761 | -929.5 | -0.16643
16.782 | 106.262 | -929.15 | -0.16642
16.782 | 106.763 | -927.9 | -0.16643
16.782 | 107.261 | -923.85 | -0.16643
16.782 | 107.764 | -910.8 | -0.16643
16.782 | 108.262 | -872.8 | -0.16643
16.782 | 108.76 | -783.5 | -0.16642
16.782 | 109.264 | -637.8 | -0.16642
16.782 | 109.761 | -483.35 | -0.16642
16.782 | 110.262 | -353.25 | -0.16642
16.782 | 110.763 | -254.1 | -0.16642
16.782 | 111.26 | -181.5 | -0.16643
16.782 | 111.763 | -128.75 | -0.16642
16.782 | 112.262 | -92 | -0.16642
16.782 | 112.76 | -66.6 | -0.16642
16.782 | 113.262 | -48.8 | -0.16642
16.782 | 113.759 | -36.45 | -0.16642
16.782 | 114.261 | -27.65 | -0.16642
16.782 | 114.762 | -21.2 | -0.16642
16.782 | 115.259 | -16.55 | -0.16642
16.782 | 115.762 | -13.05 | -0.16642
16.782 | 116.759 | -8.4 | -0.16643
16.782 | 117.76 | -5.55 | -0.16643
16.782 | 118.761 | -3.75 | -0.16643
16.782 | 119.762 | -2.6 | -0.16642
16.782 | 120.759 | -1.8 | -0.16642
16.782 | 121.76 | -1.2 | -0.16643
16.782 | 122.762 | -0.85 | -0.16643
16.782 | 123.762 | -0.55 | -0.16642
16.782 | 124.759 | -0.35 | -0.16642
16.782 | 125.76 | -0.2 | -0.16643
16.782 | 126.762 | -0.05 | -0.16643
16.782 | 127.762 | 0 | -0.16643
16.782 | 128.76 | 0.1 | -0.16643
16.782 | 129.76 | 0.15 | -0.16643
16.782 | 130.762 | 0.2 | -0.16643
16.782 | 131.763 | 0.25 | -0.16643
16.782 | 132.76 | 0.25 | -0.16643
16.782 | 133.762 | 0.25 | -0.16643
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
18:38 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-6A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -5.9903
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.766 | 0.3 | -0.19968
16.782 | 74.765 | 0.25 | -0.19968
16.782 | 75.765 | 0.2 | -0.19968
16.782 | 76.763 | 0.2 | -0.19968
16.782 | 77.763 | 0.15 | -0.19968
16.782 | 78.765 | 0.05 | -0.19968
16.782 | 79.765 | 0 | -0.19968
16.782 | 80.763 | -0.1 | -0.19968
16.782 | 81.763 | -0.3 | -0.19968
16.782 | 82.764 | -0.5 | -0.19968
16.782 | 83.765 | -0.75 | -0.19968
16.782 | 84.763 | -1.05 | -0.19968
16.782 | 85.763 | -1.55 | -0.19967
16.782 | 86.765 | -2.25 | -0.19967
16.782 | 87.765 | -3.2 | -0.19968
16.782 | 88.763 | -4.6 | -0.19968
16.782 | 89.763 | -6.75 | -0.19968
16.782 | 90.764 | -10.2 | -0.19967
16.782 | 91.765 | -15.85 | -0.19968
16.782 | 92.264 | -20.05 | -0.19968
16.782 | 92.762 | -25.7 | -0.19968
16.782 | 93.264 | -33.45 | -0.19968
16.782 | 93.762 | -44.1 | -0.19968
16.782 | 94.264 | -59.15 | -0.19968
16.782 | 94.764 | -80.55 | -0.19967
16.782 | 95.262 | -111.3 | -0.19968
16.782 | 95.764 | -156 | -0.19967
16.782 | 96.263 | -219.2 | -0.19967
16.782 | 96.762 | -307.3 | -0.19968
16.782 | 97.264 | -427.75 | -0.19968
16.782 | 97.762 | -584.35 | -0.19968
16.782 | 98.263 | -771.45 | -0.19968
16.782 | 98.763 | -944.2 | -0.19968
16.782 | 99.262 | -1049.85 | -0.19968
16.782 | 99.764 | -1094.4 | -0.19967
16.782 | 100.263 | -1109.45 | -0.19968
16.782 | 100.762 | -1114.15 | -0.19968
16.782 | 101.264 | -1115.65 | -0.19968
16.782 | 101.762 | -1116.05 | -0.19968
16.782 | 102.263 | -1116.15 | -0.19968
16.782 | 102.764 | -1116.25 | -0.19968
16.782 | 103.261 | -1116.2 | -0.19968
16.782 | 103.763 | -1116.25 | -0.19968
16.782 | 104.263 | -1116.25 | -0.19968
16.782 | 104.761 | -1116.25 | -0.19967
16.782 | 105.264 | -1116.25 | -0.19968
16.782 | 105.761 | -1116.1 | -0.19967
16.782 | 106.262 | -1115.7 | -0.19967
16.782 | 106.764 | -1114.25 | -0.19968
16.782 | 107.261 | -1109.35 | -0.19967
16.782 | 107.764 | -1093.7 | -0.19967
16.782 | 108.262 | -1048.1 | -0.19967
16.782 | 108.76 | -940.9 | -0.19967
16.782 | 109.264 | -765.95 | -0.19967
16.782 | 109.761 | -580.6 | -0.19968
16.782 | 110.262 | -424.4 | -0.19967
16.782 | 110.763 | -305.25 | -0.19967
16.782 | 111.26 | -218.1 | -0.19968
16.782 | 111.763 | -154.75 | -0.19968
16.782 | 112.262 | -110.7 | -0.19968
16.782 | 112.76 | -80.1 | -0.19967
16.782 | 113.262 | -58.8 | -0.19968
16.782 | 113.759 | -43.95 | -0.19968
16.782 | 114.261 | -33.3 | -0.19968
16.782 | 114.762 | -25.65 | -0.19968
16.782 | 115.259 | -20.05 | -0.19968
16.782 | 115.762 | -15.75 | -0.19968
16.782 | 116.759 | -10.15 | -0.19968
16.782 | 117.76 | -6.75 | -0.19968
16.782 | 118.761 | -4.6 | -0.19968
16.782 | 119.762 | -3.2 | -0.19968
16.782 | 120.759 | -2.25 | -0.19968
16.782 | 121.76 | -1.55 | -0.19968
16.782 | 122.762 | -1.05 | -0.19968
16.782 | 123.762 | -0.75 | -0.19968
16.782 | 124.76 | -0.5 | -0.19967
16.782 | 125.76 | -0.3 | -0.19968
16.782 | 126.763 | -0.2 | -0.19968
16.782 | 127.762 | -0.1 | -0.19967
16.782 | 128.76 | 0 | -0.19968
16.782 | 129.761 | 0.1 | -0.19968
16.782 | 130.762 | 0.15 | -0.19967
16.782 | 131.763 | 0.2 | -0.19967
16.782 | 132.76 | 0.25 | -0.19968
16.782 | 133.762 | 0.25 | -0.19967
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
18:47 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-7A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -6.9895
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.766 | 0.25 | -0.23298
16.782 | 74.765 | 0.2 | -0.23298
16.782 | 75.765 | 0.2 | -0.23298
16.782 | 76.764 | 0.15 | -0.23298
16.782 | 77.763 | 0.05 | -0.23298
16.782 | 78.765 | 0 | -0.23298
16.782 | 79.766 | -0.1 | -0.23298
16.782 | 80.762 | -0.25 | -0.23298
16.782 | 81.764 | -0.45 | -0.23298
16.782 | 82.765 | -0.65 | -0.23298
16.782 | 83.766 | -0.95 | -0.23298
16.782 | 84.764 | -1.35 | -0.23298
16.782 | 85.762 | -1.9 | -0.23298
16.782 | 86.766 | -2.7 | -0.23298
16.782 | 87.765 | -3.85 | -0.23298
16.782 | 88.763 | -5.45 | -0.23298
16.782 | 89.764 | -8 | -0.23298
16.782 | 90.764 | -12 | -0.23298
16.782 | 91.765 | -18.65 | -0.23298
16.782 | 92.264 | -23.6 | -0.23299
16.782 | 92.762 | -30.2 | -0.23298
16.782 | 93.264 | -39.25 | -0.23299
16.782 | 93.762 | -51.65 | -0.23298
16.782 | 94.264 | -69.3 | -0.23298
16.782 | 94.764 | -94.3 | -0.23298
16.782 | 95.262 | -130.3 | -0.23298
16.782 | 95.764 | -182.3 | -0.23298
16.782 | 96.263 | -256.1 | -0.23298
16.782 | 96.762 | -359.3 | -0.23298
16.782 | 97.265 | -499.65 | -0.23298
16.782 | 97.762 | -682.25 | -0.23298
16.782 | 98.263 | -900.7 | -0.23298
16.782 | 98.764 | -1102.15 | -0.23298
16.782 | 99.262 | -1225.5 | -0.23298
16.782 | 99.764 | -1277.4 | -0.23298
16.782 | 100.263 | -1295 | -0.23298
16.782 | 100.762 | -1300.5 | -0.23299
16.782 | 101.264 | -1302.2 | -0.23298
16.782 | 101.762 | -1302.7 | -0.23298
16.782 | 102.264 | -1302.85 | -0.23298
16.782 | 102.764 | -1302.9 | -0.23298
16.782 | 103.261 | -1302.9 | -0.23298
16.782 | 103.764 | -1302.9 | -0.23299
16.782 | 104.263 | -1302.95 | -0.23298
16.782 | 104.761 | -1303 | -0.23298
16.782 | 105.264 | -1302.95 | -0.23298
16.782 | 105.762 | -1302.85 | -0.23298
16.782 | 106.263 | -1302.35 | -0.23298
16.782 | 106.764 | -1300.6 | -0.23298
16.782 | 107.261 | -1294.95 | -0.23298
16.782 | 107.764 | -1276.6 | -0.23298
16.782 | 108.263 | -1223.45 | -0.23298
16.782 | 108.76 | -1098.4 | -0.23298
16.782 | 109.264 | -894.25 | -0.23298
16.782 | 109.761 | -677.8 | -0.23298
16.782 | 110.262 | -495.7 | -0.23298
16.782 | 110.763 | -356.75 | -0.23298
16.782 | 111.26 | -254.95 | -0.23299
16.782 | 111.763 | -181.05 | -0.23298
16.782 | 112.262 | -129.55 | -0.23298
16.782 | 112.76 | -93.75 | -0.23299
16.782 | 113.263 | -68.85 | -0.23298
16.782 | 113.76 | -51.5 | -0.23298
16.782 | 114.261 | -39.05 | -0.23298
16.782 | 114.762 | -30.05 | -0.23299
16.782 | 115.26 | -23.5 | -0.23299
16.782 | 115.762 | -18.55 | -0.23298
16.782 | 116.759 | -11.95 | -0.23299
16.782 | 117.76 | -8 | -0.23298
16.782 | 118.761 | -5.5 | -0.23298
16.782 | 119.762 | -3.85 | -0.23298
16.782 | 120.76 | -2.7 | -0.23298
16.782 | 121.76 | -1.95 | -0.23298
16.782 | 122.762 | -1.4 | -0.23299
16.782 | 123.762 | -0.95 | -0.23298
16.782 | 124.76 | -0.65 | -0.23298
16.782 | 125.761 | -0.45 | -0.23299
16.782 | 126.763 | -0.3 | -0.23298
16.782 | 127.762 | -0.15 | -0.23298
16.782 | 128.76 | -0.05 | -0.23299
16.782 | 129.761 | 0.05 | -0.23299
16.782 | 130.762 | 0.1 | -0.23299
16.782 | 131.763 | 0.15 | -0.23298
16.782 | 132.76 | 0.2 | -0.23298
16.782 | 133.762 | 0.2 | -0.23298
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
18:56 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-8A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -7.991
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.766 | 0.25 | -0.26637
16.782 | 74.765 | 0.25 | -0.26637
16.782 | 75.765 | 0.15 | -0.26637
16.782 | 76.764 | 0.1 | -0.26637
16.782 | 77.763 | 0.05 | -0.26637
16.782 | 78.765 | -0.05 | -0.26637
16.782 | 79.766 | -0.15 | -0.26637
16.782 | 80.762 | -0.35 | -0.26637
16.782 | 81.764 | -0.55 | -0.26637
16.782 | 82.765 | -0.8 | -0.26637
16.782 | 83.766 | -1.1 | -0.26637
16.782 | 84.764 | -1.6 | -0.26637
16.782 | 85.763 | -2.2 | -0.26637
16.782 | 86.765 | -3.1 | -0.26637
16.782 | 87.766 | -4.35 | -0.26637
16.782 | 88.763 | -6.25 | -0.26637
16.782 | 89.763 | -9.15 | -0.26637
16.782 | 90.764 | -13.7 | -0.26637
16.782 | 91.765 | -21.35 | -0.26637
16.782 | 92.264 | -27 | -0.26637
16.782 | 92.763 | -34.5 | -0.26637
16.782 | 93.264 | -44.9 | -0.26637
16.782 | 93.763 | -59.1 | -0.26637
16.782 | 94.265 | -79.25 | -0.26637
16.782 | 94.764 | -107.85 | -0.26637
16.782 | 95.261 | -148.8 | -0.26637
16.782 | 95.764 | -208.5 | -0.26637
16.782 | 96.263 | -292.9 | -0.26637
16.782 | 96.762 | -410.65 | -0.26637
16.782 | 97.265 | -571.4 | -0.26637
16.782 | 97.763 | -780.3 | -0.26637
16.782 | 98.263 | -1030 | -0.26637
16.782 | 98.764 | -1260.1 | -0.26637
16.782 | 99.262 | -1400.95 | -0.26637
16.782 | 99.765 | -1460.25 | -0.26637
16.782 | 100.263 | -1480.25 | -0.26637
16.782 | 100.762 | -1486.55 | -0.26637
16.782 | 101.264 | -1488.45 | -0.26637
16.782 | 101.763 | -1489 | -0.26637
16.782 | 102.264 | -1489.15 | -0.26637
16.782 | 102.764 | -1489.2 | -0.26637
16.782 | 103.261 | -1489.25 | -0.26637
16.782 | 103.764 | -1489.25 | -0.26637
16.782 | 104.263 | -1489.35 | -0.26637
16.782 | 104.761 | -1489.35 | -0.26637
16.782 | 105.264 | -1489.3 | -0.26637
16.782 | 105.762 | -1489.2 | -0.26637
16.782 | 106.263 | -1488.65 | -0.26637
16.782 | 106.764 | -1486.65 | -0.26637
16.782 | 107.261 | -1480.2 | -0.26637
16.782 | 107.765 | -1459.2 | -0.26636
16.782 | 108.263 | -1398.35 | -0.26637
16.782 | 108.76 | -1255.15 | -0.26637
16.782 | 109.264 | -1021.65 | -0.26636
16.782 | 109.762 | -774.4 | -0.26637
16.782 | 110.262 | -566.3 | -0.26637
16.782 | 110.763 | -407.6 | -0.26637
16.782 | 111.26 | -291.4 | -0.26636
16.782 | 111.764 | -206.9 | -0.26637
16.782 | 112.263 | -148.05 | -0.26637
16.782 | 112.76 | -107.15 | -0.26636
16.782 | 113.263 | -78.65 | -0.26637
16.782 | 113.76 | -58.85 | -0.26637
16.782 | 114.261 | -44.6 | -0.26637
16.782 | 114.762 | -34.3 | -0.26637
16.782 | 115.26 | -26.85 | -0.26637
16.782 | 115.763 | -21.15 | -0.26637
16.782 | 116.759 | -13.65 | -0.26637
16.782 | 117.76 | -9.1 | -0.26637
16.782 | 118.761 | -6.25 | -0.26637
16.782 | 119.762 | -4.35 | -0.26637
16.782 | 120.76 | -3.1 | -0.26637
16.782 | 121.761 | -2.2 | -0.26637
16.782 | 122.762 | -1.6 | -0.26637
16.782 | 123.762 | -1.1 | -0.26637
16.782 | 124.76 | -0.8 | -0.26637
16.782 | 125.761 | -0.55 | -0.26637
16.782 | 126.763 | -0.4 | -0.26637
16.782 | 127.762 | -0.2 | -0.26637
16.782 | 128.76 | -0.1 | -0.26637
16.782 | 129.761 | 0 | -0.26637
16.782 | 130.762 | 0 | -0.26637
16.782 | 131.764 | 0.1 | -0.26637
16.782 | 132.76 | 0.15 | -0.26637
16.782 | 133.762 | 0.2 | -0.26637
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
19:04 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-9A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -8.9956
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.766 | 0.2 | -0.29985
16.782 | 74.765 | 0.15 | -0.29985
16.782 | 75.765 | 0.1 | -0.29985
16.782 | 76.764 | 0.05 | -0.29985
16.782 | 77.763 | -0.05 | -0.29985
16.782 | 78.765 | -0.15 | -0.29985
16.782 | 79.766 | -0.3 | -0.29985
16.782 | 80.762 | -0.45 | -0.29985
16.782 | 81.764 | -0.7 | -0.29985
16.782 | 82.765 | -0.95 | -0.29985
16.782 | 83.766 | -1.3 | -0.29985
16.782 | 84.763 | -1.85 | -0.29985
16.782 | 85.762 | -2.5 | -0.29985
16.782 | 86.765 | -3.55 | -0.29985
16.782 | 87.765 | -4.95 | -0.29985
16.782 | 88.763 | -7.05 | -0.29985
16.782 | 89.763 | -10.3 | -0.29985
16.782 | 90.764 | -15.45 | -0.29985
16.782 | 91.765 | -24 | -0.29985
16.782 | 92.264 | -30.35 | -0.29985
16.782 | 92.762 | -38.85 | -0.29985
16.782 | 93.264 | -50.5 | -0.29985
16.782 | 93.763 | -66.55 | -0.29985
16.782 | 94.264 | -89.2 | -0.29985
16.782 | 94.764 | -121.45 | -0.29985
16.782 | 95.262 | -167.7 | -0.29985
16.782 | 95.764 | -234.85 | -0.29985
16.782 | 96.263 | -330 | -0.29985
16.782 | 96.762 | -462.65 | -0.29985
16.782 | 97.265 | -643.75 | -0.29985
16.782 | 97.762 | -878.85 | -0.29985
16.782 | 98.263 | -1159.85 | -0.29985
16.782 | 98.764 | -1418.65 | -0.29985
16.782 | 99.262 | -1576.65 | -0.29985
16.782 | 99.764 | -1643.15 | -0.29985
16.782 | 100.263 | -1665.55 | -0.29985
16.782 | 100.762 | -1672.6 | -0.29985
16.782 | 101.264 | -1674.75 | -0.29985
16.782 | 101.763 | -1675.35 | -0.29985
16.782 | 102.263 | -1675.5 | -0.29985
16.782 | 102.764 | -1675.55 | -0.29985
16.782 | 103.262 | -1675.55 | -0.29985
16.782 | 103.763 | -1675.6 | -0.29985
16.782 | 104.263 | -1675.65 | -0.29985
16.782 | 104.761 | -1675.7 | -0.29985
16.782 | 105.264 | -1675.7 | -0.29985
16.782 | 105.762 | -1675.5 | -0.29985
16.782 | 106.263 | -1674.85 | -0.29985
16.782 | 106.764 | -1672.65 | -0.29985
16.782 | 107.261 | -1665.3 | -0.29985
16.782 | 107.764 | -1641.6 | -0.29985
16.782 | 108.263 | -1572.85 | -0.29985
16.782 | 108.76 | -1411.15 | -0.29985
16.782 | 109.264 | -1148.65 | -0.29985
16.782 | 109.762 | -870.35 | -0.29985
16.782 | 110.262 | -636.5 | -0.29985
16.782 | 110.763 | -458.1 | -0.29985
16.782 | 111.26 | -327.55 | -0.29985
16.782 | 111.763 | -232.55 | -0.29985
16.782 | 112.262 | -166.4 | -0.29985
16.782 | 112.76 | -120.45 | -0.29985
16.782 | 113.263 | -88.4 | -0.29985
16.782 | 113.76 | -66.15 | -0.29985
16.782 | 114.261 | -50.15 | -0.29985
16.782 | 114.762 | -38.6 | -0.29985
16.782 | 115.26 | -30.15 | -0.29985
16.782 | 115.763 | -23.8 | -0.29985
16.782 | 116.759 | -15.35 | -0.29985
16.782 | 117.76 | -10.25 | -0.29985
16.782 | 118.761 | -7.05 | -0.29985
16.782 | 119.762 | -4.95 | -0.29985
16.782 | 120.76 | -3.55 | -0.29985
16.782 | 121.76 | -2.55 | -0.29985
16.782 | 122.762 | -1.85 | -0.29985
16.782 | 123.762 | -1.35 | -0.29985
16.782 | 124.76 | -0.95 | -0.29985
16.782 | 125.76 | -0.7 | -0.29985
16.782 | 126.763 | -0.5 | -0.29985
16.782 | 127.762 | -0.3 | -0.29985
16.782 | 128.76 | -0.15 | -0.29985
16.782 | 129.761 | -0.05 | -0.29985
16.782 | 130.762 | 0 | -0.29985
16.782 | 131.763 | 0.1 | -0.29985
16.782 | 132.76 | 0.15 | -0.29985
16.782 | 133.762 | 0.15 | -0.29985
! begin File Header
DL001009.stp | ! data file
2014-08-29 | ! date
19:13 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: BdL, I=-10A w/ hyst
! end File Header
! begin Step Header
0 | ! set current |  |  | -9.9925
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
16.782 | 73.766 | 0.2 | -0.33309
16.782 | 74.765 | 0.15 | -0.33309
16.782 | 75.765 | 0.05 | -0.33308
16.782 | 76.763 | 0 | -0.33308
16.782 | 77.763 | -0.05 | -0.33308
16.782 | 78.764 | -0.2 | -0.33308
16.782 | 79.766 | -0.35 | -0.33308
16.782 | 80.762 | -0.55 | -0.33308
16.782 | 81.763 | -0.8 | -0.33308
16.782 | 82.765 | -1.1 | -0.33308
16.782 | 83.766 | -1.5 | -0.33308
16.782 | 84.764 | -2.05 | -0.33308
16.782 | 85.763 | -2.85 | -0.33308
16.782 | 86.765 | -3.95 | -0.33308
16.782 | 87.765 | -5.55 | -0.33308
16.782 | 88.763 | -7.85 | -0.33308
16.782 | 89.763 | -11.45 | -0.33308
16.782 | 90.764 | -17.15 | -0.33308
16.782 | 91.765 | -26.65 | -0.33308
16.782 | 92.264 | -33.65 | -0.33308
16.782 | 92.762 | -43.1 | -0.33308
16.782 | 93.264 | -56.05 | -0.33308
16.782 | 93.762 | -73.85 | -0.33308
16.782 | 94.264 | -99.05 | -0.33308
16.782 | 94.764 | -134.9 | -0.33309
16.782 | 95.261 | -186.2 | -0.33308
16.782 | 95.764 | -260.95 | -0.33309
16.782 | 96.263 | -366.7 | -0.33308
16.782 | 96.762 | -514.05 | -0.33308
16.782 | 97.264 | -715.05 | -0.33308
16.782 | 97.762 | -976.65 | -0.33308
16.782 | 98.263 | -1288.8 | -0.33308
16.782 | 98.763 | -1575.4 | -0.33309
16.782 | 99.262 | -1750.45 | -0.33308
16.782 | 99.764 | -1824.1 | -0.33308
16.782 | 100.263 | -1848.95 | -0.33308
16.782 | 100.761 | -1856.75 | -0.33309
16.782 | 101.264 | -1859.1 | -0.33308
16.782 | 101.762 | -1859.8 | -0.33308
16.782 | 102.263 | -1860 | -0.33308
16.782 | 102.763 | -1860.05 | -0.33308
16.782 | 103.261 | -1860.05 | -0.33308
16.782 | 103.763 | -1860.05 | -0.33308
16.782 | 104.263 | -1860.1 | -0.33308
16.782 | 104.761 | -1860.15 | -0.33308
16.782 | 105.264 | -1860.1 | -0.33308
16.782 | 105.762 | -1859.95 | -0.33308
16.782 | 106.262 | -1859.25 | -0.33308
16.782 | 106.763 | -1856.75 | -0.33308
16.782 | 107.261 | -1848.55 | -0.33308
16.782 | 107.764 | -1822.15 | -0.33308
16.782 | 108.262 | -1745.65 | -0.33308
16.782 | 108.76 | -1566 | -0.33308
16.782 | 109.264 | -1274.1 | -0.33308
16.782 | 109.761 | -965.65 | -0.33308
16.782 | 110.262 | -706.1 | -0.33308
16.782 | 110.763 | -508.25 | -0.33308
16.782 | 111.26 | -363.45 | -0.33308
16.782 | 111.763 | -257.9 | -0.33308
16.782 | 112.262 | -184.5 | -0.33308
16.782 | 112.76 | -133.55 | -0.33308
16.782 | 113.262 | -98 | -0.33308
16.782 | 113.759 | -73.25 | -0.33308
16.782 | 114.261 | -55.55 | -0.33308
16.782 | 114.762 | -42.7 | -0.33308
16.782 | 115.259 | -33.4 | -0.33308
16.782 | 115.762 | -26.35 | -0.33308
16.782 | 116.759 | -16.95 | -0.33308
16.782 | 117.76 | -11.3 | -0.33308
16.782 | 118.761 | -7.8 | -0.33308
16.782 | 119.762 | -5.5 | -0.33308
16.782 | 120.759 | -3.9 | -0.33308
16.782 | 121.76 | -2.8 | -0.33308
16.782 | 122.762 | -2.05 | -0.33308
16.782 | 123.762 | -1.5 | -0.33308
16.782 | 124.76 | -1.05 | -0.33309
16.782 | 125.76 | -0.7 | -0.33308
16.782 | 126.763 | -0.5 | -0.33308
16.782 | 127.762 | -0.3 | -0.33309
16.782 | 128.76 | -0.2 | -0.33308
16.782 | 129.761 | -0.05 | -0.33308
16.782 | 130.762 | 0 | -0.33308
16.782 | 131.763 | 0.1 | -0.33308
16.782 | 132.76 | 0.15 | -0.33308
16.782 | 133.762 | 0.2 | -0.33308

## Full Grid 5.5A Data
! begin File Header
DL001004.stp | ! data file
2014-08-21 | ! date
10:47 | ! start time
5.1 | ! stepper code version
0 | ! number Z of currents
21 | ! X number of points
85 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Full Grid, I=5.5A
! end File Header
! begin Step Header
0 | ! set current
0 | ! actual current at start of step
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
11.93 | 72.172 | 0.45 | 0.1836
11.93 | 73.166 | 0.5 | 0.1836
11.93 | 74.17 | 0.5 | 0.1836
11.93 | 75.17 | 0.55 | 0.1836
11.93 | 76.166 | 0.55 | 0.1836
11.93 | 77.168 | 0.65 | 0.1836
11.93 | 78.167 | 0.75 | 0.1836
11.93 | 79.168 | 0.8 | 0.18359
11.93 | 80.166 | 0.95 | 0.1836
11.93 | 81.166 | 1.1 | 0.18359
11.93 | 82.169 | 1.35 | 0.1836
11.93 | 83.168 | 1.6 | 0.18359
11.93 | 84.166 | 1.95 | 0.1836
11.93 | 85.167 | 2.5 | 0.18359
11.93 | 86.169 | 3.25 | 0.18359
11.93 | 87.169 | 4.4 | 0.18359
11.93 | 88.165 | 6.05 | 0.18359
11.93 | 89.167 | 8.8 | 0.18359
11.93 | 90.168 | 13.3 | 0.18359
11.93 | 90.665 | 16.7 | 0.18359
11.93 | 91.167 | 21.35 | 0.18359
11.93 | 91.668 | 27.75 | 0.18359
11.93 | 92.165 | 36.5 | 0.1836
11.93 | 92.668 | 49.15 | 0.18359
11.93 | 93.165 | 67.15 | 0.18359
11.93 | 93.665 | 93.4 | 0.18359
11.93 | 94.167 | 131.95 | 0.18359
11.93 | 94.665 | 187.15 | 0.18359
11.93 | 95.167 | 265.75 | 0.18359
11.93 | 95.666 | 373.35 | 0.18359
11.93 | 96.164 | 515.9 | 0.18359
11.93 | 96.667 | 690.25 | 0.18359
11.93 | 97.165 | 855.85 | 0.18359
11.93 | 97.666 | 963.85 | 0.18359
11.93 | 98.167 | 1011.4 | 0.18359
11.93 | 98.665 | 1027.75 | 0.1836
11.93 | 99.169 | 1032.95 | 0.1836
11.93 | 99.667 | 1034.6 | 0.18359
11.93 | 100.166 | 1035.15 | 0.18359
11.93 | 100.668 | 1035.25 | 0.18359
11.93 | 101.166 | 1035.25 | 0.18359
11.93 | 101.667 | 1035.25 | 0.18359
11.93 | 102.168 | 1035.3 | 0.18359
11.93 | 102.665 | 1035.3 | 0.18359
11.93 | 103.167 | 1035.25 | 0.18359
11.93 | 103.666 | 1035.25 | 0.18359
11.93 | 104.166 | 1035.15 | 0.18359
11.93 | 104.668 | 1034.8 | 0.18359
11.93 | 105.165 | 1033.55 | 0.1836
11.93 | 105.666 | 1029.5 | 0.1836
11.93 | 106.166 | 1016.55 | 0.1836
11.93 | 106.664 | 978.4 | 0.18359
11.93 | 107.167 | 885.1 | 0.18359
11.93 | 107.666 | 729.05 | 0.1836
11.93 | 108.164 | 554.75 | 0.1836
11.93 | 108.665 | 404.35 | 0.1836
11.93 | 109.165 | 289.75 | 0.1836
11.93 | 109.664 | 205.25 | 0.1836
11.93 | 110.166 | 144.55 | 0.1836
11.93 | 110.664 | 102.35 | 0.1836
11.93 | 111.165 | 73.25 | 0.1836
11.93 | 111.665 | 53.2 | 0.1836
11.93 | 112.163 | 39.4 | 0.1836
11.93 | 112.666 | 29.55 | 0.1836
11.93 | 113.165 | 22.7 | 0.1836
11.93 | 113.663 | 17.8 | 0.1836
11.93 | 114.165 | 14 | 0.18359
11.93 | 115.165 | 9.2 | 0.18359
11.93 | 116.162 | 6.25 | 0.1836
11.93 | 117.163 | 4.45 | 0.18359
11.93 | 118.166 | 3.3 | 0.1836
11.93 | 119.165 | 2.45 | 0.1836
11.93 | 120.163 | 1.95 | 0.18359
11.93 | 121.163 | 1.55 | 0.18359
11.93 | 122.166 | 1.3 | 0.1836
11.93 | 123.165 | 1.15 | 0.18359
11.93 | 124.163 | 0.85 | 0.1836
11.93 | 125.164 | 0.75 | 0.1836
11.93 | 126.166 | 0.6 | 0.1836
11.93 | 127.166 | 0.55 | 0.1836
11.93 | 128.163 | 0.5 | 0.1836
11.93 | 129.165 | 0.45 | 0.18359
11.93 | 130.166 | 0.4 | 0.18359
11.93 | 131.166 | 0.35 | 0.1836
11.93 | 132.164 | 0.35 | 0.1836
12.428 | 72.172 | 0.4 | 0.1836
12.428 | 73.166 | 0.45 | 0.1836
12.428 | 74.169 | 0.5 | 0.1836
12.428 | 75.169 | 0.5 | 0.1836
12.428 | 76.167 | 0.55 | 0.1836
12.428 | 77.168 | 0.6 | 0.18359
12.428 | 78.168 | 0.7 | 0.1836
12.428 | 79.17 | 0.8 | 0.1836
12.428 | 80.168 | 0.9 | 0.1836
12.428 | 81.167 | 1.15 | 0.18359
12.428 | 82.17 | 1.35 | 0.18359
12.428 | 83.168 | 1.65 | 0.1836
12.428 | 84.168 | 2 | 0.1836
12.428 | 85.167 | 2.55 | 0.18359
12.428 | 86.17 | 3.3 | 0.18359
12.428 | 87.169 | 4.5 | 0.1836
12.428 | 88.166 | 6.2 | 0.18359
12.428 | 89.167 | 9 | 0.18359
12.428 | 90.169 | 13.75 | 0.1836
12.428 | 90.665 | 17.25 | 0.1836
12.428 | 91.167 | 21.95 | 0.18359
12.428 | 91.667 | 28.4 | 0.1836
12.428 | 92.165 | 37.5 | 0.1836
12.428 | 92.67 | 50.35 | 0.1836
12.428 | 93.167 | 68.55 | 0.1836
12.428 | 93.666 | 95.2 | 0.1836
12.428 | 94.167 | 133.9 | 0.1836
12.428 | 94.665 | 189.3 | 0.18359
12.428 | 95.167 | 268.05 | 0.1836
12.428 | 95.666 | 375.25 | 0.1836
12.428 | 96.164 | 517.2 | 0.1836
12.428 | 96.667 | 690.55 | 0.1836
12.428 | 97.165 | 855.75 | 0.1836
12.428 | 97.666 | 963.8 | 0.18359
12.428 | 98.167 | 1011.5 | 0.18359
12.428 | 98.665 | 1027.85 | 0.18359
12.428 | 99.168 | 1033.15 | 0.18359
12.428 | 99.667 | 1034.75 | 0.1836
12.428 | 100.165 | 1035.2 | 0.1836
12.428 | 100.668 | 1035.3 | 0.18359
12.428 | 101.165 | 1035.35 | 0.1836
12.428 | 101.667 | 1035.4 | 0.1836
12.428 | 102.168 | 1035.4 | 0.1836
12.428 | 102.665 | 1035.35 | 0.1836
12.428 | 103.167 | 1035.4 | 0.1836
12.428 | 103.666 | 1035.3 | 0.1836
12.428 | 104.165 | 1035.2 | 0.1836
12.428 | 104.668 | 1034.85 | 0.1836
12.428 | 105.165 | 1033.6 | 0.1836
12.428 | 105.666 | 1029.65 | 0.1836
12.428 | 106.167 | 1016.95 | 0.1836
12.428 | 106.665 | 979.3 | 0.1836
12.428 | 107.167 | 886.9 | 0.18359
12.428 | 107.667 | 731.85 | 0.1836
12.428 | 108.164 | 558.15 | 0.1836
12.428 | 108.665 | 407.95 | 0.1836
12.428 | 109.165 | 293.2 | 0.1836
12.428 | 109.665 | 208.2 | 0.1836
12.428 | 110.167 | 147.15 | 0.1836
12.428 | 110.665 | 104.45 | 0.1836
12.428 | 111.165 | 74.95 | 0.1836
12.428 | 111.666 | 54.55 | 0.1836
12.428 | 112.164 | 40.5 | 0.1836
12.428 | 112.666 | 30.45 | 0.1836
12.428 | 113.164 | 23.35 | 0.1836
12.428 | 113.663 | 18.25 | 0.1836
12.428 | 114.165 | 14.45 | 0.1836
12.428 | 115.165 | 9.45 | 0.1836
12.428 | 116.161 | 6.45 | 0.1836
12.428 | 117.163 | 4.6 | 0.1836
12.428 | 118.165 | 3.35 | 0.1836
12.428 | 119.164 | 2.55 | 0.1836
12.428 | 120.162 | 1.95 | 0.1836
12.428 | 121.163 | 1.6 | 0.1836
12.428 | 122.165 | 1.3 | 0.1836
12.428 | 123.165 | 1.15 | 0.1836
12.428 | 124.163 | 0.85 | 0.1836
12.428 | 125.164 | 0.7 | 0.1836
12.428 | 126.166 | 0.6 | 0.1836
12.428 | 127.166 | 0.55 | 0.1836
12.428 | 128.162 | 0.5 | 0.1836
12.428 | 129.164 | 0.45 | 0.1836
12.428 | 130.166 | 0.4 | 0.1836
12.428 | 131.166 | 0.35 | 0.1836
12.428 | 132.164 | 0.35 | 0.1836
12.928 | 72.172 | 0.4 | 0.1836
12.928 | 73.166 | 0.4 | 0.1836
12.928 | 74.169 | 0.45 | 0.1836
12.928 | 75.169 | 0.5 | 0.1836
12.928 | 76.167 | 0.6 | 0.1836
12.928 | 77.168 | 0.6 | 0.1836
12.928 | 78.168 | 0.7 | 0.1836
12.928 | 79.17 | 0.75 | 0.1836
12.928 | 80.168 | 0.9 | 0.18361
12.928 | 81.167 | 1.2 | 0.1836
12.928 | 82.169 | 1.4 | 0.1836
12.928 | 83.169 | 1.7 | 0.1836
12.928 | 84.168 | 2.05 | 0.1836
12.928 | 85.168 | 2.6 | 0.1836
12.928 | 86.168 | 3.4 | 0.1836
12.928 | 87.169 | 4.55 | 0.1836
12.928 | 88.165 | 6.4 | 0.18361
12.928 | 89.166 | 9.2 | 0.1836
12.928 | 90.169 | 14 | 0.1836
12.928 | 90.666 | 17.65 | 0.1836
12.928 | 91.167 | 22.45 | 0.1836
12.928 | 91.667 | 28.95 | 0.1836
12.928 | 92.165 | 38.15 | 0.1836
12.928 | 92.67 | 51.25 | 0.1836
12.928 | 93.165 | 69.55 | 0.1836
12.928 | 93.665 | 96.25 | 0.1836
12.928 | 94.169 | 135.35 | 0.1836
12.928 | 94.666 | 190.95 | 0.1836
12.928 | 95.168 | 269.4 | 0.1836
12.928 | 95.666 | 376.25 | 0.1836
12.928 | 96.165 | 517.75 | 0.1836
12.928 | 96.667 | 690.2 | 0.1836
12.928 | 97.166 | 855.2 | 0.1836
12.928 | 97.667 | 963.65 | 0.1836
12.928 | 98.167 | 1011.3 | 0.1836
12.928 | 98.664 | 1027.75 | 0.1836
12.928 | 99.168 | 1033 | 0.1836
12.928 | 99.667 | 1034.6 | 0.1836
12.928 | 100.165 | 1035.15 | 0.1836
12.928 | 100.667 | 1035.3 | 0.1836
12.928 | 101.165 | 1035.35 | 0.1836
12.928 | 101.667 | 1035.4 | 0.1836
12.928 | 102.168 | 1035.35 | 0.1836
12.928 | 102.665 | 1035.4 | 0.1836
12.928 | 103.168 | 1035.35 | 0.1836
12.928 | 103.666 | 1035.3 | 0.1836
12.928 | 104.165 | 1035.2 | 0.1836
12.928 | 104.668 | 1034.85 | 0.1836
12.928 | 105.165 | 1033.75 | 0.1836
12.928 | 105.665 | 1029.75 | 0.1836
12.928 | 106.166 | 1017.3 | 0.1836
12.928 | 106.665 | 979.9 | 0.1836
12.928 | 107.167 | 888.3 | 0.1836
12.928 | 107.666 | 734.2 | 0.1836
12.928 | 108.164 | 560.6 | 0.1836
12.928 | 108.666 | 410.6 | 0.1836
12.928 | 109.164 | 295.8 | 0.1836
12.928 | 109.665 | 210.55 | 0.1836
12.928 | 110.167 | 149 | 0.1836
12.928 | 110.664 | 106.15 | 0.1836
12.928 | 111.165 | 76.3 | 0.1836
12.928 | 111.665 | 55.65 | 0.1836
12.928 | 112.163 | 41.45 | 0.1836
12.928 | 112.665 | 31.2 | 0.1836
12.928 | 113.164 | 23.95 | 0.1836
12.928 | 113.662 | 18.75 | 0.1836
12.928 | 114.165 | 14.8 | 0.1836
12.928 | 115.164 | 9.75 | 0.1836
12.928 | 116.161 | 6.65 | 0.1836
12.928 | 117.163 | 4.75 | 0.1836
12.928 | 118.165 | 3.5 | 0.1836
12.928 | 119.164 | 2.6 | 0.1836
12.928 | 120.162 | 2 | 0.1836
12.928 | 121.163 | 1.65 | 0.1836
12.928 | 122.165 | 1.35 | 0.1836
12.928 | 123.165 | 1.15 | 0.1836
12.928 | 124.162 | 0.9 | 0.1836
12.928 | 125.163 | 0.8 | 0.1836
12.928 | 126.165 | 0.65 | 0.1836
12.928 | 127.166 | 0.55 | 0.1836
12.928 | 128.163 | 0.5 | 0.1836
12.928 | 129.164 | 0.45 | 0.1836
12.928 | 130.166 | 0.4 | 0.1836
12.928 | 131.166 | 0.35 | 0.1836
12.928 | 132.164 | 0.3 | 0.1836
13.428 | 72.171 | 0.4 | 0.1836
13.428 | 73.165 | 0.4 | 0.1836
13.428 | 74.169 | 0.45 | 0.1836
13.428 | 75.169 | 0.5 | 0.1836
13.428 | 76.167 | 0.55 | 0.1836
13.428 | 77.168 | 0.6 | 0.1836
13.428 | 78.167 | 0.7 | 0.1836
13.428 | 79.17 | 0.8 | 0.1836
13.428 | 80.166 | 0.9 | 0.1836
13.428 | 81.166 | 1.2 | 0.1836
13.428 | 82.17 | 1.4 | 0.1836
13.428 | 83.167 | 1.75 | 0.1836
13.428 | 84.168 | 2.05 | 0.1836
13.428 | 85.167 | 2.65 | 0.1836
13.428 | 86.169 | 3.5 | 0.18361
13.428 | 87.169 | 4.7 | 0.18361
13.428 | 88.166 | 6.5 | 0.1836
13.428 | 89.167 | 9.5 | 0.18361
13.428 | 90.168 | 14.35 | 0.1836
13.428 | 90.665 | 17.95 | 0.1836
13.428 | 91.168 | 22.8 | 0.1836
13.428 | 91.668 | 29.5 | 0.1836
13.428 | 92.166 | 38.75 | 0.1836
13.428 | 92.668 | 51.85 | 0.1836
13.428 | 93.166 | 70.35 | 0.1836
13.428 | 93.666 | 97.15 | 0.18361
13.428 | 94.167 | 136 | 0.1836
13.428 | 94.665 | 191.5 | 0.1836
13.428 | 95.169 | 270.25 | 0.18361
13.428 | 95.667 | 376.85 | 0.1836
13.428 | 96.164 | 517.25 | 0.18361
13.428 | 96.667 | 689.35 | 0.18361
13.428 | 97.166 | 854.65 | 0.1836
13.428 | 97.667 | 963.25 | 0.1836
13.428 | 98.167 | 1011.2 | 0.1836
13.428 | 98.664 | 1027.75 | 0.1836
13.428 | 99.168 | 1033.05 | 0.1836
13.428 | 99.667 | 1034.65 | 0.1836
13.428 | 100.165 | 1035.15 | 0.1836
13.428 | 100.667 | 1035.25 | 0.1836
13.428 | 101.165 | 1035.3 | 0.1836
13.428 | 101.666 | 1035.25 | 0.1836
13.428 | 102.167 | 1035.25 | 0.1836
13.428 | 102.665 | 1035.3 | 0.1836
13.428 | 103.166 | 1035.3 | 0.1836
13.428 | 103.665 | 1035.3 | 0.1836
13.428 | 104.164 | 1035.2 | 0.1836
13.428 | 104.668 | 1034.85 | 0.1836
13.428 | 105.166 | 1033.75 | 0.1836
13.428 | 105.664 | 1029.85 | 0.1836
13.428 | 106.165 | 1017.5 | 0.1836
13.428 | 106.664 | 980.45 | 0.1836
13.428 | 107.167 | 889.5 | 0.1836
13.428 | 107.666 | 735.8 | 0.1836
13.428 | 108.164 | 562.6 | 0.1836
13.428 | 108.666 | 412.4 | 0.1836
13.428 | 109.165 | 297.4 | 0.1836
13.428 | 109.665 | 212.2 | 0.1836
13.428 | 110.166 | 150.4 | 0.1836
13.428 | 110.664 | 107.3 | 0.1836
13.428 | 111.166 | 77.3 | 0.1836
13.428 | 111.666 | 56.45 | 0.1836
13.428 | 112.164 | 42.05 | 0.1836
13.428 | 112.666 | 31.75 | 0.1836
13.428 | 113.164 | 24.45 | 0.1836
13.428 | 113.662 | 19.15 | 0.1836
13.428 | 114.164 | 15.2 | 0.1836
13.428 | 115.165 | 9.9 | 0.1836
13.428 | 116.161 | 6.8 | 0.1836
13.428 | 117.162 | 4.8 | 0.1836
13.428 | 118.165 | 3.55 | 0.1836
13.428 | 119.164 | 2.65 | 0.1836
13.428 | 120.162 | 2.1 | 0.1836
13.428 | 121.163 | 1.65 | 0.1836
13.428 | 122.165 | 1.35 | 0.1836
13.428 | 123.165 | 1.15 | 0.1836
13.428 | 124.162 | 0.85 | 0.1836
13.428 | 125.164 | 0.8 | 0.1836
13.428 | 126.165 | 0.65 | 0.18359
13.428 | 127.166 | 0.55 | 0.1836
13.428 | 128.162 | 0.5 | 0.1836
13.428 | 129.164 | 0.45 | 0.1836
13.428 | 130.165 | 0.4 | 0.1836
13.428 | 131.166 | 0.35 | 0.1836
13.428 | 132.164 | 0.3 | 0.1836
13.929 | 72.171 | 0.4 | 0.1836
13.929 | 73.166 | 0.45 | 0.1836
13.929 | 74.169 | 0.45 | 0.1836
13.929 | 75.169 | 0.5 | 0.1836
13.929 | 76.165 | 0.6 | 0.1836
13.929 | 77.168 | 0.6 | 0.1836
13.929 | 78.168 | 0.75 | 0.1836
13.929 | 79.169 | 0.8 | 0.1836
13.929 | 80.168 | 0.95 | 0.1836
13.929 | 81.166 | 1.2 | 0.1836
13.929 | 82.169 | 1.4 | 0.1836
13.929 | 83.169 | 1.75 | 0.1836
13.929 | 84.166 | 2.1 | 0.1836
13.929 | 85.167 | 2.65 | 0.1836
13.929 | 86.167 | 3.5 | 0.1836
13.929 | 87.17 | 4.8 | 0.1836
13.929 | 88.165 | 6.6 | 0.1836
13.929 | 89.166 | 9.6 | 0.1836
13.929 | 90.168 | 14.55 | 0.1836
13.929 | 90.665 | 18.2 | 0.1836
13.929 | 91.168 | 23.15 | 0.1836
13.929 | 91.668 | 29.85 | 0.1836
13.929 | 92.166 | 39.15 | 0.1836
13.929 | 92.668 | 52.3 | 0.1836
13.929 | 93.166 | 70.8 | 0.1836
13.929 | 93.666 | 97.65 | 0.1836
13.929 | 94.167 | 136.45 | 0.1836
13.929 | 94.664 | 191.75 | 0.1836
13.929 | 95.167 | 270.05 | 0.1836
13.929 | 95.666 | 376.5 | 0.1836
13.929 | 96.165 | 516.8 | 0.1836
13.929 | 96.668 | 688.55 | 0.1836
13.929 | 97.166 | 853.75 | 0.1836
13.929 | 97.667 | 962.55 | 0.1836
13.929 | 98.166 | 1010.85 | 0.1836
13.929 | 98.665 | 1027.65 | 0.1836
13.929 | 99.168 | 1032.95 | 0.1836
13.929 | 99.667 | 1034.6 | 0.1836
13.929 | 100.164 | 1035.05 | 0.1836
13.929 | 100.667 | 1035.2 | 0.1836
13.929 | 101.165 | 1035.25 | 0.1836
13.929 | 101.666 | 1035.25 | 0.1836
13.929 | 102.167 | 1035.3 | 0.1836
13.929 | 102.664 | 1035.3 | 0.1836
13.929 | 103.167 | 1035.3 | 0.1836
13.929 | 103.665 | 1035.25 | 0.1836
13.929 | 104.165 | 1035.2 | 0.1836
13.929 | 104.668 | 1034.85 | 0.1836
13.929 | 105.165 | 1033.75 | 0.1836
13.929 | 105.665 | 1029.85 | 0.1836
13.929 | 106.166 | 1017.6 | 0.1836
13.929 | 106.664 | 981 | 0.1836
13.929 | 107.167 | 890.8 | 0.1836
13.929 | 107.666 | 737.6 | 0.1836
13.929 | 108.164 | 564.5 | 0.1836
13.929 | 108.666 | 414.25 | 0.1836
13.929 | 109.164 | 299.2 | 0.1836
13.929 | 109.664 | 213.6 | 0.1836
13.929 | 110.166 | 151.7 | 0.1836
13.929 | 110.664 | 108.2 | 0.1836
13.929 | 111.165 | 78 | 0.1836
13.929 | 111.665 | 57.2 | 0.1836
13.929 | 112.163 | 42.6 | 0.1836
13.929 | 112.665 | 32.25 | 0.1836
13.929 | 113.164 | 24.8 | 0.1836
13.929 | 113.662 | 19.45 | 0.1836
13.929 | 114.164 | 15.45 | 0.1836
13.929 | 115.165 | 10.1 | 0.1836
13.929 | 116.161 | 6.9 | 0.1836
13.929 | 117.162 | 4.9 | 0.1836
13.929 | 118.165 | 3.6 | 0.1836
13.929 | 119.164 | 2.75 | 0.1836
13.929 | 120.162 | 2.15 | 0.1836
13.929 | 121.163 | 1.75 | 0.1836
13.929 | 122.165 | 1.4 | 0.1836
13.929 | 123.164 | 1.15 | 0.1836
13.929 | 124.162 | 0.9 | 0.1836
13.929 | 125.164 | 0.75 | 0.1836
13.929 | 126.165 | 0.65 | 0.1836
13.929 | 127.166 | 0.55 | 0.1836
13.929 | 128.162 | 0.5 | 0.1836
13.929 | 129.164 | 0.45 | 0.1836
13.929 | 130.165 | 0.4 | 0.1836
13.929 | 131.165 | 0.35 | 0.1836
13.929 | 132.164 | 0.3 | 0.1836
14.429 | 72.171 | 0.4 | 0.1836
14.429 | 73.165 | 0.45 | 0.1836
14.429 | 74.169 | 0.5 | 0.1836
14.429 | 75.168 | 0.5 | 0.1836
14.429 | 76.166 | 0.55 | 0.1836
14.429 | 77.168 | 0.6 | 0.1836
14.429 | 78.167 | 0.7 | 0.1836
14.429 | 79.168 | 0.8 | 0.1836
14.429 | 80.166 | 0.95 | 0.1836
14.429 | 81.167 | 1.15 | 0.1836
14.429 | 82.169 | 1.45 | 0.1836
14.429 | 83.168 | 1.75 | 0.1836
14.429 | 84.167 | 2.15 | 0.1836
14.429 | 85.167 | 2.7 | 0.1836
14.429 | 86.169 | 3.55 | 0.1836
14.429 | 87.168 | 4.75 | 0.1836
14.429 | 88.165 | 6.7 | 0.1836
14.429 | 89.166 | 9.8 | 0.1836
14.429 | 90.168 | 14.75 | 0.1836
14.429 | 90.665 | 18.4 | 0.1836
14.429 | 91.168 | 23.4 | 0.1836
14.429 | 91.667 | 30.15 | 0.1836
14.429 | 92.166 | 39.45 | 0.1836
14.429 | 92.668 | 52.6 | 0.1836
14.429 | 93.165 | 71.15 | 0.1836
14.429 | 93.666 | 97.95 | 0.1836
14.429 | 94.168 | 136.75 | 0.1836
14.429 | 94.665 | 192 | 0.1836
14.429 | 95.167 | 269.8 | 0.1836
14.429 | 95.667 | 376.05 | 0.1836
14.429 | 96.164 | 515.6 | 0.1836
14.429 | 96.668 | 687.4 | 0.1836
14.429 | 97.166 | 852.65 | 0.1836
14.429 | 97.667 | 962 | 0.1836
14.429 | 98.167 | 1010.6 | 0.1836
14.429 | 98.664 | 1027.55 | 0.1836
14.429 | 99.168 | 1032.95 | 0.1836
14.429 | 99.667 | 1034.55 | 0.1836
14.429 | 100.165 | 1035 | 0.1836
14.429 | 100.667 | 1035.2 | 0.1836
14.429 | 101.165 | 1035.2 | 0.1836
14.429 | 101.666 | 1035.2 | 0.1836
14.429 | 102.167 | 1035.15 | 0.1836
14.429 | 102.664 | 1035.2 | 0.1836
14.429 | 103.167 | 1035.2 | 0.1836
14.429 | 103.665 | 1035.25 | 0.1836
14.429 | 104.165 | 1035.15 | 0.1836
14.429 | 104.667 | 1034.75 | 0.1836
14.429 | 105.165 | 1033.75 | 0.1836
14.429 | 105.665 | 1029.9 | 0.1836
14.429 | 106.165 | 1017.8 | 0.1836
14.429 | 106.664 | 981.6 | 0.1836
14.429 | 107.167 | 891.9 | 0.1836
14.429 | 107.666 | 739.2 | 0.1836
14.429 | 108.164 | 566.05 | 0.1836
14.429 | 108.666 | 415.45 | 0.1836
14.429 | 109.165 | 300.1 | 0.1836
14.429 | 109.665 | 214.55 | 0.1836
14.429 | 110.166 | 152.4 | 0.1836
14.429 | 110.664 | 108.95 | 0.1836
14.429 | 111.164 | 78.75 | 0.1836
14.429 | 111.665 | 57.7 | 0.18359
14.429 | 112.163 | 43.05 | 0.1836
14.429 | 112.665 | 32.55 | 0.1836
14.429 | 113.163 | 25.15 | 0.1836
14.429 | 113.662 | 19.7 | 0.1836
14.429 | 114.164 | 15.6 | 0.1836
14.429 | 115.164 | 10.15 | 0.1836
14.429 | 116.161 | 7.05 | 0.1836
14.429 | 117.162 | 4.95 | 0.1836
14.429 | 118.165 | 3.65 | 0.1836
14.429 | 119.164 | 2.8 | 0.1836
14.429 | 120.162 | 2.15 | 0.18359
14.429 | 121.163 | 1.75 | 0.1836
14.429 | 122.165 | 1.4 | 0.1836
14.429 | 123.165 | 1.2 | 0.1836
14.429 | 124.162 | 0.9 | 0.1836
14.429 | 125.163 | 0.8 | 0.1836
14.429 | 126.165 | 0.7 | 0.1836
14.429 | 127.165 | 0.6 | 0.1836
14.429 | 128.162 | 0.5 | 0.1836
14.429 | 129.164 | 0.45 | 0.1836
14.429 | 130.165 | 0.45 | 0.1836
14.429 | 131.165 | 0.4 | 0.1836
14.429 | 132.164 | 0.35 | 0.1836
14.929 | 72.171 | 0.35 | 0.1836
14.929 | 73.165 | 0.4 | 0.1836
14.929 | 74.169 | 0.45 | 0.1836
14.929 | 75.169 | 0.5 | 0.1836
14.929 | 76.165 | 0.55 | 0.1836
14.929 | 77.167 | 0.6 | 0.1836
14.929 | 78.168 | 0.7 | 0.1836
14.929 | 79.168 | 0.8 | 0.1836
14.929 | 80.167 | 0.95 | 0.1836
14.929 | 81.166 | 1.2 | 0.1836
14.929 | 82.168 | 1.4 | 0.1836
14.929 | 83.168 | 1.75 | 0.1836
14.929 | 84.167 | 2.15 | 0.1836
14.929 | 85.167 | 2.75 | 0.1836
14.929 | 86.169 | 3.6 | 0.1836
14.929 | 87.17 | 4.9 | 0.1836
14.929 | 88.165 | 6.75 | 0.1836
14.929 | 89.167 | 9.8 | 0.1836
14.929 | 90.167 | 14.8 | 0.1836
14.929 | 90.666 | 18.55 | 0.1836
14.929 | 91.167 | 23.5 | 0.1836
14.929 | 91.667 | 30.3 | 0.1836
14.929 | 92.166 | 39.65 | 0.1836
14.929 | 92.668 | 52.8 | 0.1836
14.929 | 93.165 | 71.35 | 0.1836
14.929 | 93.666 | 98.15 | 0.1836
14.929 | 94.167 | 136.8 | 0.1836
14.929 | 94.664 | 191.85 | 0.1836
14.929 | 95.168 | 269.85 | 0.1836
14.929 | 95.665 | 375.2 | 0.1836
14.929 | 96.163 | 514.55 | 0.1836
14.929 | 96.668 | 686.3 | 0.1836
14.929 | 97.166 | 851.55 | 0.1836
14.929 | 97.666 | 961.35 | 0.1836
14.929 | 98.166 | 1010.4 | 0.1836
14.929 | 98.664 | 1027.5 | 0.1836
14.929 | 99.167 | 1032.85 | 0.1836
14.929 | 99.666 | 1034.5 | 0.1836
14.929 | 100.165 | 1034.9 | 0.1836
14.929 | 100.667 | 1035.05 | 0.1836
14.929 | 101.165 | 1035.15 | 0.1836
14.929 | 101.666 | 1035.2 | 0.1836
14.929 | 102.167 | 1035.15 | 0.1836
14.929 | 102.664 | 1035.2 | 0.1836
14.929 | 103.166 | 1035.25 | 0.1836
14.929 | 103.665 | 1035.2 | 0.1836
14.929 | 104.165 | 1035.15 | 0.1836
14.929 | 104.667 | 1034.8 | 0.1836
14.929 | 105.165 | 1033.8 | 0.1836
14.929 | 105.665 | 1029.95 | 0.1836
14.929 | 106.165 | 1018 | 0.1836
14.929 | 106.664 | 982.05 | 0.1836
14.929 | 107.167 | 892.85 | 0.1836
14.929 | 107.666 | 740.7 | 0.1836
14.929 | 108.163 | 567.7 | 0.1836
14.929 | 108.665 | 417.05 | 0.1836
14.929 | 109.164 | 301.35 | 0.1836
14.929 | 109.664 | 215.5 | 0.1836
14.929 | 110.166 | 153.25 | 0.1836
14.929 | 110.663 | 109.65 | 0.1836
14.929 | 111.165 | 79.2 | 0.1836
14.929 | 111.665 | 58.15 | 0.1836
14.929 | 112.163 | 43.4 | 0.1836
14.929 | 112.665 | 32.85 | 0.1836
14.929 | 113.163 | 25.4 | 0.1836
14.929 | 113.661 | 19.9 | 0.1836
14.929 | 114.165 | 15.75 | 0.1836
14.929 | 115.165 | 10.3 | 0.1836
14.929 | 116.161 | 7.15 | 0.1836
14.929 | 117.162 | 5.1 | 0.1836
14.929 | 118.164 | 3.75 | 0.1836
14.929 | 119.164 | 2.85 | 0.1836
14.929 | 120.161 | 2.15 | 0.1836
14.929 | 121.162 | 1.75 | 0.1836
14.929 | 122.165 | 1.45 | 0.1836
14.929 | 123.164 | 1.15 | 0.1836
14.929 | 124.162 | 0.9 | 0.1836
14.929 | 125.163 | 0.8 | 0.1836
14.929 | 126.165 | 0.7 | 0.1836
14.929 | 127.165 | 0.6 | 0.1836
14.929 | 128.162 | 0.5 | 0.1836
14.929 | 129.164 | 0.5 | 0.1836
14.929 | 130.165 | 0.5 | 0.1836
14.929 | 131.165 | 0.35 | 0.1836
14.929 | 132.163 | 0.3 | 0.1836
15.429 | 72.171 | 0.4 | 0.1836
15.429 | 73.165 | 0.45 | 0.1836
15.429 | 74.169 | 0.45 | 0.1836
15.429 | 75.168 | 0.45 | 0.1836
15.429 | 76.167 | 0.55 | 0.1836
15.429 | 77.167 | 0.6 | 0.1836
15.429 | 78.166 | 0.7 | 0.1836
15.429 | 79.169 | 0.8 | 0.1836
15.429 | 80.167 | 0.95 | 0.1836
15.429 | 81.166 | 1.2 | 0.1836
15.429 | 82.169 | 1.45 | 0.1836
15.429 | 83.168 | 1.75 | 0.1836
15.429 | 84.167 | 2.2 | 0.1836
15.429 | 85.166 | 2.75 | 0.1836
15.429 | 86.169 | 3.65 | 0.1836
15.429 | 87.169 | 4.9 | 0.1836
15.429 | 88.164 | 6.8 | 0.1836
15.429 | 89.166 | 9.85 | 0.1836
15.429 | 90.168 | 14.9 | 0.1836
15.429 | 90.665 | 18.6 | 0.1836
15.429 | 91.167 | 23.6 | 0.1836
15.429 | 91.666 | 30.4 | 0.1836
15.429 | 92.166 | 39.75 | 0.1836
15.429 | 92.668 | 52.85 | 0.1836
15.429 | 93.164 | 71.45 | 0.1836
15.429 | 93.666 | 98.15 | 0.1836
15.429 | 94.168 | 136.9 | 0.1836
15.429 | 94.664 | 191.8 | 0.1836
15.429 | 95.168 | 269.55 | 0.1836
15.429 | 95.666 | 374.9 | 0.1836
15.429 | 96.164 | 513.95 | 0.1836
15.429 | 96.668 | 685.15 | 0.1836
15.429 | 97.166 | 850.5 | 0.1836
15.429 | 97.666 | 960.8 | 0.1836
15.429 | 98.166 | 1010.1 | 0.1836
15.429 | 98.663 | 1027.35 | 0.1836
15.429 | 99.167 | 1032.75 | 0.1836
15.429 | 99.667 | 1034.45 | 0.1836
15.429 | 100.165 | 1034.95 | 0.1836
15.429 | 100.666 | 1035 | 0.1836
15.429 | 101.164 | 1035.15 | 0.1836
15.429 | 101.666 | 1035.1 | 0.1836
15.429 | 102.167 | 1035.2 | 0.1836
15.429 | 102.665 | 1035.15 | 0.1836
15.429 | 103.166 | 1035.2 | 0.1836
15.429 | 103.665 | 1035.15 | 0.1836
15.429 | 104.164 | 1035.1 | 0.1836
15.429 | 104.667 | 1034.75 | 0.1836
15.429 | 105.164 | 1033.65 | 0.1836
15.429 | 105.664 | 1029.95 | 0.1836
15.429 | 106.166 | 1018 | 0.1836
15.429 | 106.664 | 982.35 | 0.1836
15.429 | 107.166 | 893.75 | 0.1836
15.429 | 107.665 | 742.15 | 0.1836
15.429 | 108.164 | 568.7 | 0.1836
15.429 | 108.665 | 418 | 0.1836
15.429 | 109.164 | 302.25 | 0.1836
15.429 | 109.664 | 216.2 | 0.1836
15.429 | 110.166 | 153.8 | 0.1836
15.429 | 110.664 | 110.15 | 0.18359
15.429 | 111.165 | 79.55 | 0.1836
15.429 | 111.665 | 58.4 | 0.1836
15.429 | 112.163 | 43.65 | 0.1836
15.429 | 112.665 | 33.1 | 0.1836
15.429 | 113.163 | 25.5 | 0.1836
15.429 | 113.661 | 20 | 0.1836
15.429 | 114.165 | 15.85 | 0.1836
15.429 | 115.165 | 10.35 | 0.1836
15.429 | 116.161 | 7.15 | 0.1836
15.429 | 117.162 | 5.05 | 0.1836
15.429 | 118.164 | 3.75 | 0.1836
15.429 | 119.164 | 2.8 | 0.1836
15.429 | 120.162 | 2.15 | 0.1836
15.429 | 121.162 | 1.75 | 0.1836
15.429 | 122.164 | 1.4 | 0.1836
15.429 | 123.164 | 1.2 | 0.1836
15.429 | 124.162 | 0.9 | 0.1836
15.429 | 125.163 | 0.75 | 0.1836
15.429 | 126.165 | 0.6 | 0.1836
15.429 | 127.165 | 0.55 | 0.1836
15.429 | 128.162 | 0.5 | 0.1836
15.429 | 129.163 | 0.4 | 0.1836
15.429 | 130.165 | 0.35 | 0.1836
15.429 | 131.165 | 0.3 | 0.1836
15.429 | 132.163 | 0.3 | 0.1836
15.929 | 72.171 | 0.3 | 0.1836
15.929 | 73.165 | 0.4 | 0.1836
15.929 | 74.169 | 0.45 | 0.1836
15.929 | 75.168 | 0.5 | 0.1836
15.929 | 76.166 | 0.5 | 0.1836
15.929 | 77.167 | 0.55 | 0.1836
15.929 | 78.167 | 0.65 | 0.1836
15.929 | 79.169 | 0.8 | 0.1836
15.929 | 80.165 | 0.9 | 0.1836
15.929 | 81.166 | 1.2 | 0.1836
15.929 | 82.169 | 1.4 | 0.1836
15.929 | 83.168 | 1.75 | 0.1836
15.929 | 84.167 | 2.15 | 0.1836
15.929 | 85.166 | 2.75 | 0.1836
15.929 | 86.169 | 3.55 | 0.1836
15.929 | 87.168 | 4.9 | 0.1836
15.929 | 88.165 | 6.8 | 0.1836
15.929 | 89.166 | 9.9 | 0.1836
15.929 | 90.167 | 14.95 | 0.1836
15.929 | 90.664 | 18.65 | 0.18359
15.929 | 91.167 | 23.7 | 0.1836
15.929 | 91.666 | 30.45 | 0.1836
15.929 | 92.165 | 39.8 | 0.1836
15.929 | 92.668 | 52.9 | 0.1836
15.929 | 93.164 | 71.45 | 0.1836
15.929 | 93.665 | 97.95 | 0.1836
15.929 | 94.168 | 136.7 | 0.1836
15.929 | 94.665 | 191.6 | 0.1836
15.929 | 95.166 | 268.8 | 0.1836
15.929 | 95.666 | 374.15 | 0.1836
15.929 | 96.164 | 512.95 | 0.1836
15.929 | 96.667 | 683.65 | 0.1836
15.929 | 97.165 | 849.3 | 0.1836
15.929 | 97.666 | 960.15 | 0.1836
15.929 | 98.166 | 1009.8 | 0.1836
15.929 | 98.664 | 1027.2 | 0.1836
15.929 | 99.167 | 1032.75 | 0.1836
15.929 | 99.666 | 1034.4 | 0.1836
15.929 | 100.164 | 1034.85 | 0.1836
15.929 | 100.666 | 1035 | 0.1836
15.929 | 101.164 | 1035 | 0.1836
15.929 | 101.665 | 1035 | 0.1836
15.929 | 102.167 | 1035 | 0.1836
15.929 | 102.663 | 1035.05 | 0.1836
15.929 | 103.167 | 1035.05 | 0.1836
15.929 | 103.665 | 1035.15 | 0.1836
15.929 | 104.165 | 1034.95 | 0.1836
15.929 | 104.667 | 1034.75 | 0.1836
15.929 | 105.164 | 1033.65 | 0.1836
15.929 | 105.665 | 1030 | 0.1836
15.929 | 106.165 | 1018.2 | 0.1836
15.929 | 106.664 | 982.75 | 0.1836
15.929 | 107.166 | 894.75 | 0.1836
15.929 | 107.665 | 743.3 | 0.1836
15.929 | 108.163 | 570.15 | 0.1836
15.929 | 108.665 | 418.9 | 0.1836
15.929 | 109.164 | 302.95 | 0.1836
15.929 | 109.664 | 216.75 | 0.1836
15.929 | 110.165 | 154.35 | 0.1836
15.929 | 110.663 | 110.5 | 0.1836
15.929 | 111.165 | 79.8 | 0.1836
15.929 | 111.665 | 58.6 | 0.1836
15.929 | 112.163 | 43.8 | 0.1836
15.929 | 112.665 | 33.2 | 0.1836
15.929 | 113.163 | 25.6 | 0.1836
15.929 | 113.661 | 20.15 | 0.1836
15.929 | 114.164 | 15.95 | 0.1836
15.929 | 115.164 | 10.45 | 0.1836
15.929 | 116.16 | 7.25 | 0.1836
15.929 | 117.162 | 5.1 | 0.1836
15.929 | 118.164 | 3.75 | 0.1836
15.929 | 119.163 | 2.75 | 0.1836
15.929 | 120.161 | 2.15 | 0.1836
15.929 | 121.162 | 1.75 | 0.1836
15.929 | 122.164 | 1.4 | 0.1836
15.929 | 123.164 | 1.15 | 0.1836
15.929 | 124.161 | 0.9 | 0.1836
15.929 | 125.163 | 0.75 | 0.1836
15.929 | 126.164 | 0.65 | 0.1836
15.929 | 127.165 | 0.55 | 0.1836
15.929 | 128.162 | 0.5 | 0.1836
15.929 | 129.163 | 0.45 | 0.1836
15.929 | 130.165 | 0.35 | 0.1836
15.929 | 131.165 | 0.35 | 0.1836
15.929 | 132.163 | 0.3 | 0.1836
16.429 | 72.17 | 0.3 | 0.1836
16.429 | 73.165 | 0.3 | 0.1836
16.429 | 74.168 | 0.4 | 0.1836
16.429 | 75.168 | 0.45 | 0.1836
16.429 | 76.166 | 0.5 | 0.1836
16.429 | 77.166 | 0.55 | 0.1836
16.429 | 78.166 | 0.65 | 0.1836
16.429 | 79.169 | 0.75 | 0.1836
16.429 | 80.166 | 0.9 | 0.1836
16.429 | 81.165 | 1.2 | 0.1836
16.429 | 82.169 | 1.4 | 0.1836
16.429 | 83.168 | 1.75 | 0.1836
16.429 | 84.166 | 2.15 | 0.1836
16.429 | 85.166 | 2.7 | 0.18359
16.429 | 86.168 | 3.6 | 0.1836
16.429 | 87.168 | 4.9 | 0.1836
16.429 | 88.165 | 6.8 | 0.1836
16.429 | 89.166 | 9.9 | 0.1836
16.429 | 90.167 | 14.95 | 0.1836
16.429 | 90.665 | 18.65 | 0.1836
16.429 | 91.167 | 23.7 | 0.1836
16.429 | 91.665 | 30.45 | 0.1836
16.429 | 92.165 | 39.8 | 0.1836
16.429 | 92.668 | 52.85 | 0.1836
16.429 | 93.165 | 71.4 | 0.1836
16.429 | 93.664 | 97.85 | 0.1836
16.429 | 94.167 | 136.4 | 0.1836
16.429 | 94.665 | 191.35 | 0.1836
16.429 | 95.167 | 268.25 | 0.1836
16.429 | 95.664 | 372.8 | 0.1836
16.429 | 96.164 | 511.9 | 0.1836
16.429 | 96.667 | 682.45 | 0.1836
16.429 | 97.165 | 848 | 0.1836
16.429 | 97.665 | 959.35 | 0.1836
16.429 | 98.167 | 1009.6 | 0.1836
16.429 | 98.663 | 1027.15 | 0.1836
16.429 | 99.167 | 1032.65 | 0.1836
16.429 | 99.666 | 1034.35 | 0.1836
16.429 | 100.164 | 1034.8 | 0.1836
16.429 | 100.666 | 1034.95 | 0.1836
16.429 | 101.164 | 1035 | 0.1836
16.429 | 101.665 | 1035 | 0.1836
16.429 | 102.167 | 1034.95 | 0.1836
16.429 | 102.664 | 1035 | 0.1836
16.429 | 103.166 | 1035 | 0.1836
16.429 | 103.665 | 1035 | 0.1836
16.429 | 104.164 | 1034.95 | 0.1836
16.429 | 104.666 | 1034.75 | 0.1836
16.429 | 105.164 | 1033.6 | 0.1836
16.429 | 105.664 | 1029.95 | 0.1836
16.429 | 106.165 | 1018.25 | 0.1836
16.429 | 106.663 | 983.2 | 0.1836
16.429 | 107.166 | 895.65 | 0.1836
16.429 | 107.665 | 744.5 | 0.1836
16.429 | 108.163 | 571.4 | 0.1836
16.429 | 108.665 | 419.95 | 0.1836
16.429 | 109.163 | 303.85 | 0.1836
16.429 | 109.664 | 217.4 | 0.1836
16.429 | 110.165 | 154.8 | 0.1836
16.429 | 110.663 | 110.85 | 0.1836
16.429 | 111.165 | 80.1 | 0.1836
16.429 | 111.665 | 58.8 | 0.1836
16.429 | 112.163 | 43.95 | 0.1836
16.429 | 112.664 | 33.35 | 0.1836
16.429 | 113.163 | 25.75 | 0.1836
16.429 | 113.662 | 20.2 | 0.1836
16.429 | 114.165 | 16 | 0.1836
16.429 | 115.164 | 10.45 | 0.1836
16.429 | 116.16 | 7.25 | 0.1836
16.429 | 117.161 | 5.15 | 0.1836
16.429 | 118.164 | 3.75 | 0.1836
16.429 | 119.163 | 2.8 | 0.1836
16.429 | 120.161 | 2.2 | 0.1836
16.429 | 121.162 | 1.75 | 0.1836
16.429 | 122.164 | 1.4 | 0.1836
16.429 | 123.164 | 1.2 | 0.1836
16.429 | 124.161 | 0.9 | 0.1836
16.429 | 125.163 | 0.75 | 0.1836
16.429 | 126.165 | 0.65 | 0.1836
16.429 | 127.165 | 0.55 | 0.1836
16.429 | 128.162 | 0.45 | 0.1836
16.429 | 129.163 | 0.4 | 0.1836
16.429 | 130.164 | 0.35 | 0.1836
16.429 | 131.165 | 0.3 | 0.1836
16.429 | 132.163 | 0.25 | 0.1836
16.929 | 72.17 | 0.35 | 0.18359
16.929 | 73.165 | 0.35 | 0.1836
16.929 | 74.168 | 0.4 | 0.1836
16.929 | 75.167 | 0.4 | 0.1836
16.929 | 76.166 | 0.45 | 0.1836
16.929 | 77.167 | 0.55 | 0.1836
16.929 | 78.166 | 0.65 | 0.1836
16.929 | 79.168 | 0.75 | 0.18359
16.929 | 80.166 | 0.85 | 0.1836
16.929 | 81.165 | 1.15 | 0.1836
16.929 | 82.168 | 1.35 | 0.1836
16.929 | 83.168 | 1.75 | 0.1836
16.929 | 84.166 | 2.15 | 0.1836
16.929 | 85.167 | 2.75 | 0.1836
16.929 | 86.168 | 3.6 | 0.1836
16.929 | 87.169 | 4.9 | 0.1836
16.929 | 88.164 | 6.8 | 0.1836
16.929 | 89.166 | 9.9 | 0.1836
16.929 | 90.167 | 14.95 | 0.1836
16.929 | 90.665 | 18.65 | 0.1836
16.929 | 91.166 | 23.6 | 0.1836
16.929 | 91.666 | 30.4 | 0.1836
16.929 | 92.166 | 39.75 | 0.1836
16.929 | 92.667 | 52.75 | 0.1836
16.929 | 93.165 | 71.25 | 0.1836
16.929 | 93.665 | 97.75 | 0.1836
16.929 | 94.167 | 136.2 | 0.18359
16.929 | 94.665 | 190.75 | 0.1836
16.929 | 95.166 | 267.55 | 0.1836
16.929 | 95.664 | 371.85 | 0.1836
16.929 | 96.164 | 510.45 | 0.1836
16.929 | 96.666 | 680.4 | 0.1836
16.929 | 97.164 | 846.55 | 0.1836
16.929 | 97.666 | 958.8 | 0.1836
16.929 | 98.166 | 1009.2 | 0.1836
16.929 | 98.663 | 1026.9 | 0.1836
16.929 | 99.167 | 1032.6 | 0.1836
16.929 | 99.666 | 1034.3 | 0.1836
16.929 | 100.163 | 1034.8 | 0.18359
16.929 | 100.666 | 1034.9 | 0.18359
16.929 | 101.164 | 1034.9 | 0.18359
16.929 | 101.665 | 1034.9 | 0.1836
16.929 | 102.166 | 1034.95 | 0.18359
16.929 | 102.664 | 1034.95 | 0.1836
16.929 | 103.166 | 1034.95 | 0.1836
16.929 | 103.664 | 1035 | 0.1836
16.929 | 104.164 | 1034.95 | 0.1836
16.929 | 104.666 | 1034.7 | 0.1836
16.929 | 105.164 | 1033.6 | 0.1836
16.929 | 105.664 | 1030 | 0.1836
16.929 | 106.165 | 1018.4 | 0.1836
16.929 | 106.662 | 983.75 | 0.1836
16.929 | 107.166 | 896.7 | 0.1836
16.929 | 107.665 | 746.15 | 0.1836
16.929 | 108.163 | 572.85 | 0.1836
16.929 | 108.665 | 421.15 | 0.1836
16.929 | 109.164 | 304.6 | 0.1836
16.929 | 109.664 | 218.05 | 0.1836
16.929 | 110.165 | 155.2 | 0.1836
16.929 | 110.663 | 111.25 | 0.1836
16.929 | 111.165 | 80.25 | 0.1836
16.929 | 111.664 | 58.95 | 0.1836
16.929 | 112.163 | 44.15 | 0.1836
16.929 | 112.665 | 33.4 | 0.1836
16.929 | 113.162 | 25.8 | 0.1836
16.929 | 113.661 | 20.2 | 0.1836
16.929 | 114.164 | 16.1 | 0.1836
16.929 | 115.164 | 10.5 | 0.1836
16.929 | 116.16 | 7.3 | 0.1836
16.929 | 117.161 | 5.15 | 0.18359
16.929 | 118.164 | 3.7 | 0.1836
16.929 | 119.163 | 2.8 | 0.1836
16.929 | 120.161 | 2.15 | 0.1836
16.929 | 121.162 | 1.7 | 0.1836
16.929 | 122.164 | 1.35 | 0.1836
16.929 | 123.164 | 1.15 | 0.1836
16.929 | 124.161 | 0.85 | 0.1836
16.929 | 125.162 | 0.75 | 0.1836
16.929 | 126.164 | 0.6 | 0.1836
16.929 | 127.164 | 0.5 | 0.1836
16.929 | 128.161 | 0.45 | 0.1836
16.929 | 129.163 | 0.4 | 0.1836
16.929 | 130.164 | 0.3 | 0.1836
16.929 | 131.164 | 0.3 | 0.1836
16.929 | 132.163 | 0.25 | 0.1836
17.43 | 72.17 | 0.3 | 0.1836
17.43 | 73.164 | 0.3 | 0.1836
17.43 | 74.168 | 0.35 | 0.1836
17.43 | 75.167 | 0.4 | 0.1836
17.43 | 76.165 | 0.5 | 0.1836
17.43 | 77.167 | 0.55 | 0.1836
17.43 | 78.166 | 0.6 | 0.1836
17.43 | 79.168 | 0.75 | 0.1836
17.43 | 80.165 | 0.85 | 0.1836
17.43 | 81.165 | 1.2 | 0.1836
17.43 | 82.168 | 1.35 | 0.1836
17.43 | 83.166 | 1.7 | 0.1836
17.43 | 84.167 | 2.05 | 0.1836
17.43 | 85.165 | 2.7 | 0.1836
17.43 | 86.168 | 3.55 | 0.1836
17.43 | 87.167 | 4.8 | 0.1836
17.43 | 88.165 | 6.8 | 0.1836
17.43 | 89.166 | 9.8 | 0.1836
17.43 | 90.167 | 14.9 | 0.1836
17.43 | 90.664 | 18.6 | 0.1836
17.43 | 91.166 | 23.55 | 0.1836
17.43 | 91.665 | 30.25 | 0.1836
17.43 | 92.165 | 39.55 | 0.1836
17.43 | 92.668 | 52.6 | 0.1836
17.43 | 93.165 | 70.95 | 0.1836
17.43 | 93.663 | 97.25 | 0.1836
17.43 | 94.167 | 135.65 | 0.1836
17.43 | 94.664 | 190 | 0.1836
17.43 | 95.166 | 266.6 | 0.1836
17.43 | 95.666 | 371.2 | 0.1836
17.43 | 96.163 | 508.8 | 0.1836
17.43 | 96.667 | 678.95 | 0.1836
17.43 | 97.165 | 845.55 | 0.1836
17.43 | 97.666 | 958.1 | 0.1836
17.43 | 98.165 | 1008.85 | 0.1836
17.43 | 98.663 | 1026.8 | 0.1836
17.43 | 99.167 | 1032.5 | 0.1836
17.43 | 99.666 | 1034.25 | 0.1836
17.43 | 100.163 | 1034.75 | 0.1836
17.43 | 100.665 | 1034.85 | 0.1836
17.43 | 101.164 | 1034.85 | 0.1836
17.43 | 101.665 | 1034.9 | 0.1836
17.43 | 102.166 | 1034.95 | 0.1836
17.43 | 102.663 | 1034.95 | 0.1836
17.43 | 103.166 | 1034.95 | 0.1836
17.43 | 103.664 | 1035 | 0.18359
17.43 | 104.164 | 1034.9 | 0.1836
17.43 | 104.666 | 1034.65 | 0.1836
17.43 | 105.164 | 1033.6 | 0.1836
17.43 | 105.664 | 1030.05 | 0.1836
17.43 | 106.165 | 1018.55 | 0.1836
17.43 | 106.663 | 984.1 | 0.1836
17.43 | 107.165 | 897.85 | 0.1836
17.43 | 107.664 | 747.8 | 0.1836
17.43 | 108.163 | 574.4 | 0.1836
17.43 | 108.664 | 422.35 | 0.1836
17.43 | 109.163 | 305.5 | 0.1836
17.43 | 109.664 | 218.6 | 0.1836
17.43 | 110.165 | 155.6 | 0.1836
17.43 | 110.662 | 111.5 | 0.18359
17.43 | 111.163 | 80.5 | 0.1836
17.43 | 111.664 | 59.1 | 0.1836
17.43 | 112.162 | 44.15 | 0.1836
17.43 | 112.664 | 33.4 | 0.1836
17.43 | 113.163 | 25.75 | 0.1836
17.43 | 113.661 | 20.15 | 0.1836
17.43 | 114.163 | 16 | 0.1836
17.43 | 115.164 | 10.45 | 0.1836
17.43 | 116.16 | 7.2 | 0.1836
17.43 | 117.161 | 5.15 | 0.1836
17.43 | 118.164 | 3.7 | 0.1836
17.43 | 119.163 | 2.75 | 0.1836
17.43 | 120.161 | 2.15 | 0.1836
17.43 | 121.162 | 1.7 | 0.1836
17.43 | 122.164 | 1.4 | 0.1836
17.43 | 123.164 | 1.15 | 0.1836
17.43 | 124.161 | 0.85 | 0.1836
17.43 | 125.162 | 0.7 | 0.1836
17.43 | 126.164 | 0.6 | 0.1836
17.43 | 127.164 | 0.45 | 0.1836
17.43 | 128.161 | 0.45 | 0.1836
17.43 | 129.163 | 0.35 | 0.1836
17.43 | 130.164 | 0.3 | 0.1836
17.43 | 131.165 | 0.3 | 0.1836
17.43 | 132.162 | 0.25 | 0.1836
17.93 | 72.17 | 0.25 | 0.1836
17.93 | 73.164 | 0.3 | 0.1836
17.93 | 74.167 | 0.35 | 0.1836
17.93 | 75.168 | 0.4 | 0.1836
17.93 | 76.164 | 0.45 | 0.1836
17.93 | 77.167 | 0.55 | 0.1836
17.93 | 78.167 | 0.6 | 0.1836
17.93 | 79.168 | 0.75 | 0.1836
17.93 | 80.166 | 0.85 | 0.1836
17.93 | 81.164 | 1.15 | 0.1836
17.93 | 82.168 | 1.35 | 0.1836
17.93 | 83.167 | 1.7 | 0.1836
17.93 | 84.166 | 2.05 | 0.1836
17.93 | 85.166 | 2.7 | 0.1836
17.93 | 86.168 | 3.55 | 0.1836
17.93 | 87.169 | 4.8 | 0.1836
17.93 | 88.164 | 6.75 | 0.1836
17.93 | 89.165 | 9.8 | 0.1836
17.93 | 90.167 | 14.85 | 0.1836
17.93 | 90.663 | 18.45 | 0.1836
17.93 | 91.167 | 23.4 | 0.1836
17.93 | 91.666 | 30.15 | 0.1836
17.93 | 92.164 | 39.35 | 0.1836
17.93 | 92.668 | 52.4 | 0.1836
17.93 | 93.164 | 70.65 | 0.1836
17.93 | 93.664 | 96.95 | 0.1836
17.93 | 94.167 | 135.25 | 0.1836
17.93 | 94.664 | 189.55 | 0.1836
17.93 | 95.166 | 265.75 | 0.1836
17.93 | 95.665 | 369.75 | 0.1836
17.93 | 96.164 | 507.65 | 0.1836
17.93 | 96.665 | 677.15 | 0.1836
17.93 | 97.165 | 844.15 | 0.1836
17.93 | 97.666 | 957.45 | 0.1836
17.93 | 98.165 | 1008.6 | 0.1836
17.93 | 98.663 | 1026.7 | 0.1836
17.93 | 99.166 | 1032.5 | 0.1836
17.93 | 99.665 | 1034.25 | 0.1836
17.93 | 100.163 | 1034.75 | 0.1836
17.93 | 100.665 | 1034.85 | 0.1836
17.93 | 101.164 | 1034.85 | 0.1836
17.93 | 101.665 | 1034.9 | 0.1836
17.93 | 102.166 | 1034.9 | 0.1836
17.93 | 102.663 | 1034.9 | 0.1836
17.93 | 103.165 | 1034.95 | 0.1836
17.93 | 103.664 | 1034.9 | 0.1836
17.93 | 104.164 | 1034.85 | 0.1836
17.93 | 104.666 | 1034.6 | 0.1836
17.93 | 105.164 | 1033.55 | 0.1836
17.93 | 105.663 | 1030.15 | 0.1836
17.93 | 106.165 | 1018.7 | 0.1836
17.93 | 106.663 | 984.6 | 0.1836
17.93 | 107.165 | 898.8 | 0.1836
17.93 | 107.664 | 749.15 | 0.1836
17.93 | 108.162 | 575.7 | 0.1836
17.93 | 108.665 | 423.2 | 0.1836
17.93 | 109.164 | 306.1 | 0.1836
17.93 | 109.663 | 219.1 | 0.1836
17.93 | 110.165 | 155.85 | 0.1836
17.93 | 110.663 | 111.6 | 0.1836
17.93 | 111.163 | 80.55 | 0.1836
17.93 | 111.664 | 59.15 | 0.1836
17.93 | 112.162 | 44.2 | 0.1836
17.93 | 112.664 | 33.5 | 0.1836
17.93 | 113.162 | 25.8 | 0.1836
17.93 | 113.661 | 20.2 | 0.1836
17.93 | 114.164 | 16 | 0.1836
17.93 | 115.164 | 10.45 | 0.1836
17.93 | 116.161 | 7.2 | 0.1836
17.93 | 117.161 | 5.1 | 0.1836
17.93 | 118.163 | 3.7 | 0.1836
17.93 | 119.163 | 2.75 | 0.1836
17.93 | 120.161 | 2.15 | 0.1836
17.93 | 121.162 | 1.7 | 0.1836
17.93 | 122.164 | 1.35 | 0.1836
17.93 | 123.164 | 1.15 | 0.1836
17.93 | 124.161 | 0.8 | 0.1836
17.93 | 125.162 | 0.75 | 0.1836
17.93 | 126.164 | 0.55 | 0.1836
17.93 | 127.164 | 0.5 | 0.1836
17.93 | 128.161 | 0.45 | 0.1836
17.93 | 129.163 | 0.4 | 0.1836
17.93 | 130.164 | 0.3 | 0.1836
17.93 | 131.164 | 0.25 | 0.1836
17.93 | 132.162 | 0.25 | 0.1836
18.43 | 72.169 | 0.3 | 0.1836
18.43 | 73.164 | 0.3 | 0.1836
18.43 | 74.168 | 0.35 | 0.1836
18.43 | 75.167 | 0.4 | 0.1836
18.43 | 76.165 | 0.45 | 0.1836
18.43 | 77.165 | 0.5 | 0.1836
18.43 | 78.166 | 0.6 | 0.1836
18.43 | 79.167 | 0.75 | 0.1836
18.43 | 80.165 | 0.8 | 0.1836
18.43 | 81.164 | 1.15 | 0.1836
18.43 | 82.168 | 1.3 | 0.1836
18.43 | 83.166 | 1.65 | 0.1836
18.43 | 84.166 | 2 | 0.1836
18.43 | 85.165 | 2.6 | 0.1836
18.43 | 86.167 | 3.5 | 0.1836
18.43 | 87.169 | 4.8 | 0.1836
18.43 | 88.164 | 6.75 | 0.1836
18.43 | 89.164 | 9.75 | 0.1836
18.43 | 90.167 | 14.65 | 0.1836
18.43 | 90.664 | 18.35 | 0.1836
18.43 | 91.165 | 23.2 | 0.1836
18.43 | 91.666 | 29.95 | 0.1836
18.43 | 92.165 | 39.2 | 0.1836
18.43 | 92.667 | 52.2 | 0.1836
18.43 | 93.163 | 70.3 | 0.1836
18.43 | 93.665 | 96.65 | 0.1836
18.43 | 94.167 | 134.75 | 0.1836
18.43 | 94.663 | 188.55 | 0.1836
18.43 | 95.166 | 265 | 0.1836
18.43 | 95.665 | 368.75 | 0.1836
18.43 | 96.164 | 506.2 | 0.1836
18.43 | 96.666 | 675.65 | 0.1836
18.43 | 97.164 | 842.5 | 0.1836
18.43 | 97.666 | 956.65 | 0.1836
18.43 | 98.166 | 1008.3 | 0.1836
18.43 | 98.662 | 1026.55 | 0.1836
18.43 | 99.166 | 1032.4 | 0.1836
18.43 | 99.665 | 1034.2 | 0.1836
18.43 | 100.164 | 1034.7 | 0.1836
18.43 | 100.665 | 1034.85 | 0.1836
18.43 | 101.163 | 1034.8 | 0.1836
18.43 | 101.665 | 1034.85 | 0.1836
18.43 | 102.166 | 1034.85 | 0.1836
18.43 | 102.663 | 1034.9 | 0.1836
18.43 | 103.166 | 1034.95 | 0.1836
18.43 | 103.664 | 1034.95 | 0.1836
18.43 | 104.163 | 1034.85 | 0.1836
18.43 | 104.666 | 1034.6 | 0.1836
18.43 | 105.164 | 1033.6 | 0.1836
18.43 | 105.664 | 1030.2 | 0.1836
18.43 | 106.164 | 1018.85 | 0.1836
18.43 | 106.662 | 985.15 | 0.1836
18.43 | 107.165 | 899.7 | 0.1836
18.43 | 107.665 | 750.4 | 0.1836
18.43 | 108.162 | 577.15 | 0.1836
18.43 | 108.664 | 424.25 | 0.1836
18.43 | 109.163 | 306.8 | 0.1836
18.43 | 109.663 | 219.5 | 0.1836
18.43 | 110.164 | 156.15 | 0.1836
18.43 | 110.663 | 111.75 | 0.1836
18.43 | 111.164 | 80.55 | 0.1836
18.43 | 111.664 | 59.05 | 0.1836
18.43 | 112.162 | 44.15 | 0.1836
18.43 | 112.664 | 33.35 | 0.1836
18.43 | 113.162 | 25.75 | 0.1836
18.43 | 113.66 | 20.15 | 0.1836
18.43 | 114.163 | 15.95 | 0.1836
18.43 | 115.163 | 10.4 | 0.1836
18.43 | 116.16 | 7.15 | 0.1836
18.43 | 117.161 | 5 | 0.1836
18.43 | 118.163 | 3.6 | 0.1836
18.43 | 119.162 | 2.75 | 0.1836
18.43 | 120.161 | 2.15 | 0.1836
18.43 | 121.161 | 1.7 | 0.1836
18.43 | 122.164 | 1.3 | 0.1836
18.43 | 123.163 | 1.15 | 0.1836
18.43 | 124.161 | 0.8 | 0.1836
18.43 | 125.162 | 0.7 | 0.1836
18.43 | 126.164 | 0.6 | 0.1836
18.43 | 127.164 | 0.45 | 0.1836
18.43 | 128.161 | 0.4 | 0.1836
18.43 | 129.163 | 0.35 | 0.1836
18.43 | 130.164 | 0.25 | 0.1836
18.43 | 131.164 | 0.25 | 0.1836
18.43 | 132.162 | 0.25 | 0.1836
18.93 | 72.169 | 0.3 | 0.1836
18.93 | 73.164 | 0.3 | 0.1836
18.93 | 74.168 | 0.3 | 0.1836
18.93 | 75.166 | 0.3 | 0.1836
18.93 | 76.165 | 0.45 | 0.1836
18.93 | 77.165 | 0.45 | 0.1836
18.93 | 78.166 | 0.6 | 0.1836
18.93 | 79.168 | 0.75 | 0.1836
18.93 | 80.164 | 0.8 | 0.1836
18.93 | 81.164 | 1.15 | 0.1836
18.93 | 82.168 | 1.25 | 0.1836
18.93 | 83.166 | 1.6 | 0.1836
18.93 | 84.166 | 2 | 0.1836
18.93 | 85.166 | 2.55 | 0.1836
18.93 | 86.166 | 3.45 | 0.1836
18.93 | 87.168 | 4.8 | 0.1836
18.93 | 88.162 | 6.6 | 0.1836
18.93 | 89.165 | 9.6 | 0.1836
18.93 | 90.167 | 14.55 | 0.1836
18.93 | 90.664 | 18.15 | 0.1836
18.93 | 91.165 | 23.05 | 0.1836
18.93 | 91.666 | 29.75 | 0.1836
18.93 | 92.165 | 38.9 | 0.1836
18.93 | 92.667 | 51.8 | 0.1836
18.93 | 93.163 | 69.85 | 0.1836
18.93 | 93.664 | 96.15 | 0.1836
18.93 | 94.167 | 134.15 | 0.1836
18.93 | 94.663 | 187.85 | 0.1836
18.93 | 95.166 | 264 | 0.1836
18.93 | 95.665 | 367.55 | 0.1836
18.93 | 96.163 | 504.8 | 0.1836
18.93 | 96.666 | 674.3 | 0.1836
18.93 | 97.164 | 841.6 | 0.1836
18.93 | 97.665 | 956 | 0.1836
18.93 | 98.165 | 1008.05 | 0.1836
18.93 | 98.663 | 1026.5 | 0.1836
18.93 | 99.166 | 1032.35 | 0.1836
18.93 | 99.665 | 1034.2 | 0.1836
18.93 | 100.163 | 1034.7 | 0.1836
18.93 | 100.665 | 1034.8 | 0.1836
18.93 | 101.164 | 1034.85 | 0.1836
18.93 | 101.665 | 1034.85 | 0.1836
18.93 | 102.166 | 1034.85 | 0.1836
18.93 | 102.663 | 1034.9 | 0.1836
18.93 | 103.166 | 1034.9 | 0.1836
18.93 | 103.664 | 1034.9 | 0.1836
18.93 | 104.163 | 1034.85 | 0.1836
18.93 | 104.665 | 1034.6 | 0.1836
18.93 | 105.164 | 1033.55 | 0.1836
18.93 | 105.664 | 1030.15 | 0.1836
18.93 | 106.164 | 1018.95 | 0.1836
18.93 | 106.663 | 985.4 | 0.1836
18.93 | 107.165 | 900.7 | 0.1836
18.93 | 107.665 | 751.7 | 0.1836
18.93 | 108.162 | 578.3 | 0.1836
18.93 | 108.663 | 425.15 | 0.1836
18.93 | 109.163 | 307.25 | 0.1836
18.93 | 109.663 | 219.75 | 0.1836
18.93 | 110.165 | 156.2 | 0.1836
18.93 | 110.662 | 111.8 | 0.1836
18.93 | 111.164 | 80.5 | 0.1836
18.93 | 111.664 | 58.9 | 0.1836
18.93 | 112.162 | 43.95 | 0.1836
18.93 | 112.664 | 33.2 | 0.1836
18.93 | 113.162 | 25.55 | 0.1836
18.93 | 113.661 | 20 | 0.1836
18.93 | 114.163 | 15.8 | 0.1836
18.93 | 115.164 | 10.25 | 0.1836
18.93 | 116.16 | 7.15 | 0.1836
18.93 | 117.161 | 5 | 0.1836
18.93 | 118.163 | 3.55 | 0.1836
18.93 | 119.163 | 2.75 | 0.1836
18.93 | 120.161 | 2.1 | 0.1836
18.93 | 121.161 | 1.65 | 0.1836
18.93 | 122.164 | 1.3 | 0.1836
18.93 | 123.163 | 1.1 | 0.1836
18.93 | 124.161 | 0.8 | 0.1836
18.93 | 125.162 | 0.7 | 0.1836
18.93 | 126.164 | 0.55 | 0.1836
18.93 | 127.164 | 0.5 | 0.1836
18.93 | 128.161 | 0.45 | 0.1836
18.93 | 129.163 | 0.35 | 0.1836
18.93 | 130.164 | 0.3 | 0.1836
18.93 | 131.164 | 0.25 | 0.1836
18.93 | 132.162 | 0.25 | 0.1836
19.43 | 72.169 | 0.3 | 0.1836
19.43 | 73.163 | 0.3 | 0.1836
19.43 | 74.167 | 0.35 | 0.1836
19.43 | 75.166 | 0.4 | 0.1836
19.43 | 76.165 | 0.45 | 0.1836
19.43 | 77.166 | 0.5 | 0.1836
19.43 | 78.166 | 0.6 | 0.1836
19.43 | 79.167 | 0.7 | 0.1836
19.43 | 80.165 | 0.8 | 0.1836
19.43 | 81.164 | 1.05 | 0.1836
19.43 | 82.167 | 1.3 | 0.1836
19.43 | 83.167 | 1.6 | 0.1836
19.43 | 84.164 | 1.95 | 0.1836
19.43 | 85.165 | 2.55 | 0.1836
19.43 | 86.166 | 3.45 | 0.1836
19.43 | 87.168 | 4.7 | 0.1836
19.43 | 88.162 | 6.55 | 0.1836
19.43 | 89.165 | 9.5 | 0.1836
19.43 | 90.166 | 14.35 | 0.1836
19.43 | 90.664 | 17.95 | 0.1836
19.43 | 91.165 | 22.75 | 0.1836
19.43 | 91.665 | 29.4 | 0.1836
19.43 | 92.165 | 38.5 | 0.1836
19.43 | 92.667 | 51.4 | 0.1836
19.43 | 93.163 | 69.35 | 0.1836
19.43 | 93.664 | 95.55 | 0.1836
19.43 | 94.166 | 133.3 | 0.1836
19.43 | 94.663 | 186.95 | 0.1836
19.43 | 95.165 | 262.8 | 0.1836
19.43 | 95.665 | 366.35 | 0.1836
19.43 | 96.163 | 503.2 | 0.1836
19.43 | 96.666 | 672.6 | 0.1836
19.43 | 97.164 | 840.2 | 0.1836
19.43 | 97.665 | 955.3 | 0.1836
19.43 | 98.164 | 1007.7 | 0.1836
19.43 | 98.663 | 1026.4 | 0.1836
19.43 | 99.166 | 1032.4 | 0.1836
19.43 | 99.665 | 1034.25 | 0.1836
19.43 | 100.163 | 1034.7 | 0.1836
19.43 | 100.665 | 1034.85 | 0.1836
19.43 | 101.163 | 1034.85 | 0.1836
19.43 | 101.664 | 1034.9 | 0.1836
19.43 | 102.165 | 1034.95 | 0.1836
19.43 | 102.662 | 1035 | 0.1836
19.43 | 103.165 | 1034.95 | 0.1836
19.43 | 103.664 | 1034.95 | 0.1836
19.43 | 104.163 | 1034.9 | 0.1836
19.43 | 104.665 | 1034.65 | 0.1836
19.43 | 105.164 | 1033.65 | 0.1836
19.43 | 105.663 | 1030.25 | 0.1836
19.43 | 106.164 | 1019.2 | 0.1836
19.43 | 106.663 | 985.85 | 0.1836
19.43 | 107.165 | 901.8 | 0.1836
19.43 | 107.664 | 753.25 | 0.1836
19.43 | 108.162 | 579.45 | 0.1836
19.43 | 108.664 | 425.75 | 0.1836
19.43 | 109.163 | 307.75 | 0.1836
19.43 | 109.663 | 219.9 | 0.1836
19.43 | 110.164 | 156.25 | 0.1836
19.43 | 110.662 | 111.7 | 0.1836
19.43 | 111.163 | 80.35 | 0.1836
19.43 | 111.664 | 58.75 | 0.1836
19.43 | 112.162 | 43.75 | 0.1836
19.43 | 112.664 | 33.05 | 0.1836
19.43 | 113.162 | 25.4 | 0.1836
19.43 | 113.661 | 19.85 | 0.1836
19.43 | 114.163 | 15.75 | 0.1836
19.43 | 115.164 | 10.2 | 0.1836
19.43 | 116.16 | 6.95 | 0.1836
19.43 | 117.161 | 4.9 | 0.1836
19.43 | 118.163 | 3.55 | 0.1836
19.43 | 119.162 | 2.7 | 0.18361
19.43 | 120.161 | 2.1 | 0.1836
19.43 | 121.161 | 1.7 | 0.1836
19.43 | 122.164 | 1.35 | 0.1836
19.43 | 123.163 | 1.15 | 0.1836
19.43 | 124.161 | 0.8 | 0.1836
19.43 | 125.162 | 0.75 | 0.1836
19.43 | 126.164 | 0.6 | 0.1836
19.43 | 127.164 | 0.5 | 0.1836
19.43 | 128.161 | 0.45 | 0.1836
19.43 | 129.162 | 0.35 | 0.1836
19.43 | 130.164 | 0.3 | 0.1836
19.43 | 131.164 | 0.3 | 0.18361
19.43 | 132.162 | 0.25 | 0.1836
19.93 | 72.169 | 0.25 | 0.1836
19.93 | 73.163 | 0.3 | 0.1836
19.93 | 74.167 | 0.35 | 0.1836
19.93 | 75.166 | 0.4 | 0.1836
19.93 | 76.165 | 0.45 | 0.1836
19.93 | 77.166 | 0.5 | 0.1836
19.93 | 78.165 | 0.6 | 0.1836
19.93 | 79.168 | 0.75 | 0.1836
19.93 | 80.165 | 0.85 | 0.1836
19.93 | 81.164 | 1.1 | 0.1836
19.93 | 82.167 | 1.3 | 0.1836
19.93 | 83.167 | 1.6 | 0.1836
19.93 | 84.165 | 1.95 | 0.1836
19.93 | 85.166 | 2.55 | 0.1836
19.93 | 86.166 | 3.4 | 0.1836
19.93 | 87.168 | 4.6 | 0.1836
19.93 | 88.164 | 6.45 | 0.18361
19.93 | 89.163 | 9.35 | 0.1836
19.93 | 90.167 | 14.15 | 0.1836
19.93 | 90.663 | 17.75 | 0.1836
19.93 | 91.165 | 22.55 | 0.18361
19.93 | 91.666 | 29.1 | 0.1836
19.93 | 92.165 | 38.15 | 0.1836
19.93 | 92.666 | 50.8 | 0.1836
19.93 | 93.163 | 68.75 | 0.1836
19.93 | 93.664 | 94.75 | 0.1836
19.93 | 94.166 | 132.4 | 0.1836
19.93 | 94.663 | 185.85 | 0.1836
19.93 | 95.167 | 261.85 | 0.1836
19.93 | 95.663 | 364.5 | 0.1836
19.93 | 96.162 | 501.15 | 0.1836
19.93 | 96.666 | 670.9 | 0.1836
19.93 | 97.164 | 838.7 | 0.18361
19.93 | 97.664 | 954.45 | 0.1836
19.93 | 98.165 | 1007.55 | 0.1836
19.93 | 98.663 | 1026.35 | 0.18361
19.93 | 99.166 | 1032.4 | 0.18361
19.93 | 99.665 | 1034.25 | 0.1836
19.93 | 100.163 | 1034.75 | 0.1836
19.93 | 100.665 | 1034.9 | 0.18361
19.93 | 101.163 | 1034.95 | 0.1836
19.93 | 101.664 | 1035 | 0.1836
19.93 | 102.165 | 1035 | 0.1836
19.93 | 102.662 | 1035 | 0.1836
19.93 | 103.165 | 1035 | 0.1836
19.93 | 103.664 | 1035 | 0.1836
19.93 | 104.163 | 1034.9 | 0.1836
19.93 | 104.665 | 1034.7 | 0.1836
19.93 | 105.164 | 1033.65 | 0.1836
19.93 | 105.663 | 1030.3 | 0.18361
19.93 | 106.164 | 1019.4 | 0.18361
19.93 | 106.663 | 986.35 | 0.18361
19.93 | 107.165 | 902.75 | 0.18361
19.93 | 107.664 | 754.55 | 0.1836
19.93 | 108.162 | 580.55 | 0.18361
19.93 | 108.664 | 426.6 | 0.18361
19.93 | 109.163 | 308 | 0.18361
19.93 | 109.662 | 219.95 | 0.18361
19.93 | 110.164 | 156.1 | 0.1836
19.93 | 110.662 | 111.4 | 0.18361
19.93 | 111.163 | 80.1 | 0.18361
19.93 | 111.663 | 58.5 | 0.18361
19.93 | 112.162 | 43.5 | 0.18361
19.93 | 112.664 | 32.75 | 0.1836
19.93 | 113.162 | 25.25 | 0.1836
19.93 | 113.66 | 19.7 | 0.18361
19.93 | 114.163 | 15.55 | 0.18361
19.93 | 115.163 | 10.1 | 0.18361
19.93 | 116.16 | 6.85 | 0.1836
19.93 | 117.161 | 4.85 | 0.1836
19.93 | 118.163 | 3.55 | 0.1836
19.93 | 119.163 | 2.65 | 0.1836
19.93 | 120.16 | 2.05 | 0.1836
19.93 | 121.161 | 1.6 | 0.1836
19.93 | 122.164 | 1.35 | 0.1836
19.93 | 123.163 | 1.15 | 0.1836
19.93 | 124.161 | 0.8 | 0.1836
19.93 | 125.162 | 0.75 | 0.18361
19.93 | 126.163 | 0.6 | 0.1836
19.93 | 127.164 | 0.55 | 0.1836
19.93 | 128.161 | 0.45 | 0.1836
19.93 | 129.162 | 0.4 | 0.1836
19.93 | 130.164 | 0.35 | 0.1836
19.93 | 131.164 | 0.3 | 0.1836
19.93 | 132.162 | 0.25 | 0.1836
20.431 | 72.169 | 0.3 | 0.1836
20.431 | 73.163 | 0.35 | 0.1836
20.431 | 74.166 | 0.35 | 0.1836
20.431 | 75.167 | 0.45 | 0.1836
20.431 | 76.164 | 0.45 | 0.1836
20.431 | 77.166 | 0.55 | 0.1836
20.431 | 78.165 | 0.6 | 0.1836
20.431 | 79.167 | 0.7 | 0.1836
20.431 | 80.165 | 0.8 | 0.18361
20.431 | 81.164 | 1.1 | 0.1836
20.431 | 82.167 | 1.3 | 0.1836
20.431 | 83.167 | 1.6 | 0.1836
20.431 | 84.165 | 2 | 0.1836
20.431 | 85.166 | 2.55 | 0.1836
20.431 | 86.167 | 3.35 | 0.1836
20.431 | 87.168 | 4.5 | 0.18361
20.431 | 88.162 | 6.35 | 0.1836
20.431 | 89.165 | 9.2 | 0.1836
20.431 | 90.166 | 13.85 | 0.1836
20.431 | 90.663 | 17.45 | 0.1836
20.431 | 91.165 | 22.15 | 0.18361
20.431 | 91.666 | 28.6 | 0.1836
20.431 | 92.165 | 37.55 | 0.1836
20.431 | 92.667 | 50.15 | 0.18361
20.431 | 93.163 | 68 | 0.1836
20.431 | 93.664 | 93.8 | 0.18361
20.431 | 94.165 | 131.25 | 0.18361
20.431 | 94.662 | 184.45 | 0.1836
20.431 | 95.166 | 260.25 | 0.1836
20.431 | 95.664 | 362.8 | 0.1836
20.431 | 96.162 | 499.25 | 0.18361
20.431 | 96.666 | 668.9 | 0.18361
20.431 | 97.164 | 837.35 | 0.18361
20.431 | 97.665 | 953.75 | 0.18361
20.431 | 98.165 | 1007.25 | 0.18361
20.431 | 98.662 | 1026.2 | 0.18361
20.431 | 99.166 | 1032.35 | 0.1836
20.431 | 99.665 | 1034.2 | 0.1836
20.431 | 100.162 | 1034.75 | 0.1836
20.431 | 100.666 | 1034.9 | 0.1836
20.431 | 101.163 | 1034.9 | 0.1836
20.431 | 101.664 | 1034.9 | 0.1836
20.431 | 102.165 | 1034.95 | 0.18361
20.431 | 102.662 | 1034.95 | 0.1836
20.431 | 103.165 | 1035 | 0.1836
20.431 | 103.664 | 1035 | 0.1836
20.431 | 104.163 | 1034.95 | 0.1836
20.431 | 104.665 | 1034.75 | 0.1836
20.431 | 105.163 | 1033.75 | 0.1836
20.431 | 105.663 | 1030.35 | 0.18361
20.431 | 106.163 | 1019.55 | 0.1836
20.431 | 106.662 | 986.75 | 0.1836
20.431 | 107.165 | 903.5 | 0.18361
20.431 | 107.664 | 755.7 | 0.1836
20.431 | 108.161 | 581.55 | 0.18361
20.431 | 108.664 | 427.1 | 0.18361
20.431 | 109.163 | 307.95 | 0.18361
20.431 | 109.662 | 219.65 | 0.1836
20.431 | 110.164 | 155.6 | 0.1836
20.431 | 110.662 | 110.75 | 0.1836
20.431 | 111.163 | 79.55 | 0.18361
20.431 | 111.663 | 57.95 | 0.18361
20.431 | 112.162 | 43 | 0.18361
20.431 | 112.663 | 32.4 | 0.1836
20.431 | 113.161 | 24.8 | 0.1836
20.431 | 113.66 | 19.4 | 0.1836
20.431 | 114.163 | 15.3 | 0.1836
20.431 | 115.163 | 9.95 | 0.18361
20.431 | 116.16 | 6.75 | 0.18361
20.431 | 117.16 | 4.75 | 0.1836
20.431 | 118.163 | 3.5 | 0.18361
20.431 | 119.163 | 2.65 | 0.18361
20.431 | 120.16 | 2 | 0.18361
20.431 | 121.162 | 1.65 | 0.18361
20.431 | 122.163 | 1.3 | 0.18361
20.431 | 123.163 | 1.1 | 0.18361
20.431 | 124.161 | 0.85 | 0.18361
20.431 | 125.162 | 0.75 | 0.18361
20.431 | 126.163 | 0.6 | 0.18361
20.431 | 127.163 | 0.5 | 0.18361
20.431 | 128.161 | 0.45 | 0.18361
20.431 | 129.162 | 0.4 | 0.18361
20.431 | 130.163 | 0.35 | 0.18361
20.431 | 131.164 | 0.3 | 0.1836
20.431 | 132.162 | 0.25 | 0.18361
20.931 | 72.169 | 0.3 | 0.1836
20.931 | 73.163 | 0.35 | 0.1836
20.931 | 74.167 | 0.35 | 0.1836
20.931 | 75.166 | 0.35 | 0.1836
20.931 | 76.164 | 0.5 | 0.1836
20.931 | 77.165 | 0.55 | 0.1836
20.931 | 78.165 | 0.6 | 0.1836
20.931 | 79.167 | 0.75 | 0.1836
20.931 | 80.165 | 0.8 | 0.1836
20.931 | 81.164 | 1.05 | 0.1836
20.931 | 82.167 | 1.3 | 0.1836
20.931 | 83.166 | 1.55 | 0.1836
20.931 | 84.164 | 1.95 | 0.1836
20.931 | 85.165 | 2.45 | 0.1836
20.931 | 86.167 | 3.25 | 0.1836
20.931 | 87.167 | 4.45 | 0.1836
20.931 | 88.164 | 6.2 | 0.18361
20.931 | 89.164 | 8.95 | 0.1836
20.931 | 90.166 | 13.55 | 0.1836
20.931 | 90.663 | 17.1 | 0.1836
20.931 | 91.166 | 21.75 | 0.1836
20.931 | 91.665 | 28 | 0.1836
20.931 | 92.163 | 36.8 | 0.1836
20.931 | 92.667 | 49.3 | 0.1836
20.931 | 93.163 | 66.85 | 0.1836
20.931 | 93.664 | 92.5 | 0.18361
20.931 | 94.164 | 129.55 | 0.1836
20.931 | 94.663 | 182.8 | 0.1836
20.931 | 95.166 | 258.15 | 0.1836
20.931 | 95.663 | 360.45 | 0.1836
20.931 | 96.162 | 496.9 | 0.18361
20.931 | 96.666 | 666.75 | 0.1836
20.931 | 97.163 | 835.15 | 0.1836
20.931 | 97.665 | 952.75 | 0.1836
20.931 | 98.164 | 1006.8 | 0.18361
20.931 | 98.661 | 1026.1 | 0.1836
20.931 | 99.166 | 1032.3 | 0.1836
20.931 | 99.665 | 1034.2 | 0.1836
20.931 | 100.162 | 1034.75 | 0.1836
20.931 | 100.665 | 1034.9 | 0.1836
20.931 | 101.163 | 1034.95 | 0.1836
20.931 | 101.664 | 1035 | 0.1836
20.931 | 102.165 | 1035.05 | 0.1836
20.931 | 102.663 | 1035.15 | 0.1836
20.931 | 103.164 | 1035.15 | 0.18361
20.931 | 103.663 | 1035.15 | 0.1836
20.931 | 104.163 | 1035 | 0.1836
20.931 | 104.665 | 1034.75 | 0.1836
20.931 | 105.163 | 1033.75 | 0.1836
20.931 | 105.663 | 1030.45 | 0.1836
20.931 | 106.164 | 1019.7 | 0.1836
20.931 | 106.662 | 987.25 | 0.1836
20.931 | 107.164 | 904.55 | 0.1836
20.931 | 107.664 | 756.9 | 0.1836
20.931 | 108.161 | 582.3 | 0.18361
20.931 | 108.664 | 427.15 | 0.1836
20.931 | 109.163 | 307.6 | 0.1836
20.931 | 109.663 | 218.9 | 0.18361
20.931 | 110.164 | 154.8 | 0.1836
20.931 | 110.661 | 110.05 | 0.1836
20.931 | 111.164 | 78.8 | 0.1836
20.931 | 111.664 | 57.3 | 0.1836
20.931 | 112.161 | 42.45 | 0.1836
20.931 | 112.663 | 31.9 | 0.1836
20.931 | 113.162 | 24.45 | 0.1836
20.931 | 113.661 | 19.05 | 0.1836
20.931 | 114.163 | 15 | 0.1836
20.931 | 115.163 | 9.75 | 0.1836
20.931 | 116.159 | 6.65 | 0.1836
20.931 | 117.161 | 4.75 | 0.1836
20.931 | 118.163 | 3.45 | 0.1836
20.931 | 119.163 | 2.55 | 0.1836
20.931 | 120.161 | 2 | 0.1836
20.931 | 121.161 | 1.6 | 0.1836
20.931 | 122.163 | 1.35 | 0.1836
20.931 | 123.163 | 1.1 | 0.1836
20.931 | 124.161 | 0.8 | 0.1836
20.931 | 125.162 | 0.75 | 0.18361
20.931 | 126.163 | 0.6 | 0.1836
20.931 | 127.163 | 0.5 | 0.1836
20.931 | 128.16 | 0.45 | 0.1836
20.931 | 129.162 | 0.4 | 0.1836
20.931 | 130.163 | 0.35 | 0.18361
20.931 | 131.164 | 0.3 | 0.1836
20.931 | 132.162 | 0.3 | 0.1836
21.431 | 72.168 | 0.3 | 0.1836
21.431 | 73.163 | 0.35 | 0.1836
21.431 | 74.167 | 0.4 | 0.1836
21.431 | 75.166 | 0.45 | 0.18361
21.431 | 76.164 | 0.45 | 0.18361
21.431 | 77.165 | 0.55 | 0.18361
21.431 | 78.165 | 0.6 | 0.1836
21.431 | 79.166 | 0.75 | 0.1836
21.431 | 80.163 | 0.8 | 0.18361
21.431 | 81.164 | 1 | 0.18361
21.431 | 82.167 | 1.3 | 0.1836
21.431 | 83.166 | 1.55 | 0.18361
21.431 | 84.165 | 1.9 | 0.18361
21.431 | 85.165 | 2.45 | 0.18361
21.431 | 86.165 | 3.2 | 0.18361
21.431 | 87.167 | 4.35 | 0.18361
21.431 | 88.163 | 6 | 0.18361
21.431 | 89.164 | 8.75 | 0.1836
21.431 | 90.167 | 13.25 | 0.18361
21.431 | 90.663 | 16.6 | 0.18361
21.431 | 91.166 | 21.2 | 0.18361
21.431 | 91.665 | 27.4 | 0.18361
21.431 | 92.164 | 36 | 0.1836
21.431 | 92.666 | 48.2 | 0.18361
21.431 | 93.163 | 65.6 | 0.1836
21.431 | 93.664 | 90.85 | 0.18361
21.431 | 94.165 | 127.8 | 0.18361
21.431 | 94.663 | 180.6 | 0.18361
21.431 | 95.166 | 255.6 | 0.18361
21.431 | 95.663 | 357.75 | 0.1836
21.431 | 96.162 | 494.15 | 0.18361
21.431 | 96.666 | 664 | 0.1836
21.431 | 97.164 | 833.45 | 0.18361
21.431 | 97.664 | 951.6 | 0.18361
21.431 | 98.164 | 1006.3 | 0.18361
21.431 | 98.662 | 1025.8 | 0.18361
21.431 | 99.165 | 1032.15 | 0.18361
21.431 | 99.664 | 1034.05 | 0.18361
21.431 | 100.162 | 1034.6 | 0.18361
21.431 | 100.665 | 1034.8 | 0.18361
21.431 | 101.163 | 1034.9 | 0.18361
21.431 | 101.664 | 1034.95 | 0.18361
21.431 | 102.165 | 1035 | 0.18361
21.431 | 102.662 | 1035 | 0.18361
21.431 | 103.164 | 1035 | 0.18361
21.431 | 103.663 | 1035 | 0.18361
21.431 | 104.163 | 1034.95 | 0.18361
21.431 | 104.665 | 1034.65 | 0.18361
21.431 | 105.162 | 1033.65 | 0.18361
21.431 | 105.662 | 1030.35 | 0.18361
21.431 | 106.164 | 1019.7 | 0.18361
21.431 | 106.662 | 987.3 | 0.18361
21.431 | 107.164 | 904.85 | 0.18361
21.431 | 107.663 | 757.45 | 0.1836
21.431 | 108.162 | 581.85 | 0.18361
21.431 | 108.663 | 426.3 | 0.18361
21.431 | 109.163 | 306.3 | 0.18361
21.431 | 109.662 | 217.5 | 0.18361
21.431 | 110.164 | 153.35 | 0.18361
21.431 | 110.662 | 108.55 | 0.18361
21.431 | 111.163 | 77.6 | 0.18361
21.431 | 111.663 | 56.3 | 0.18361
21.431 | 112.161 | 41.6 | 0.18361
21.431 | 112.663 | 31.25 | 0.18361
21.431 | 113.161 | 23.8 | 0.18361
21.431 | 113.659 | 18.6 | 0.18361
21.431 | 114.162 | 14.65 | 0.18361
21.431 | 115.163 | 9.45 | 0.18361
21.431 | 116.159 | 6.45 | 0.1836
21.431 | 117.161 | 4.6 | 0.18361
21.431 | 118.162 | 3.3 | 0.18361
21.431 | 119.161 | 2.5 | 0.18361
21.431 | 120.16 | 2 | 0.18361
21.431 | 121.16 | 1.6 | 0.18361
21.431 | 122.163 | 1.3 | 0.18361
21.431 | 123.162 | 1.15 | 0.18361
21.431 | 124.16 | 0.8 | 0.18361
21.431 | 125.161 | 0.75 | 0.18361
21.431 | 126.163 | 0.6 | 0.18361
21.431 | 127.163 | 0.55 | 0.18361
21.431 | 128.16 | 0.5 | 0.18361
21.431 | 129.162 | 0.45 | 0.18361
21.431 | 130.163 | 0.35 | 0.18361
21.431 | 131.163 | 0.35 | 0.18361
21.431 | 132.162 | 0.3 | 0.18361
21.931 | 72.168 | 0.3 | 0.1836
21.931 | 73.163 | 0.35 | 0.1836
21.931 | 74.166 | 0.35 | 0.18361
21.931 | 75.166 | 0.4 | 0.18361
21.931 | 76.163 | 0.45 | 0.1836
21.931 | 77.165 | 0.6 | 0.1836
21.931 | 78.165 | 0.6 | 0.1836
21.931 | 79.166 | 0.75 | 0.18361
21.931 | 80.165 | 0.8 | 0.1836
21.931 | 81.164 | 1 | 0.1836
21.931 | 82.167 | 1.25 | 0.18361
21.931 | 83.165 | 1.5 | 0.1836
21.931 | 84.165 | 1.9 | 0.18361
21.931 | 85.164 | 2.35 | 0.1836
21.931 | 86.166 | 3.15 | 0.1836
21.931 | 87.168 | 4.2 | 0.18361
21.931 | 88.163 | 5.8 | 0.1836
21.931 | 89.163 | 8.45 | 0.1836
21.931 | 90.166 | 12.8 | 0.1836
21.931 | 90.662 | 16.05 | 0.1836
21.931 | 91.165 | 20.45 | 0.1836
21.931 | 91.665 | 26.5 | 0.1836
21.931 | 92.162 | 34.85 | 0.1836
21.931 | 92.667 | 46.9 | 0.1836
21.931 | 93.163 | 63.85 | 0.1836
21.931 | 93.663 | 88.65 | 0.18361
21.931 | 94.165 | 125 | 0.18361
21.931 | 94.663 | 177.45 | 0.1836
21.931 | 95.165 | 251.8 | 0.1836
21.931 | 95.663 | 353.65 | 0.18361
21.931 | 96.162 | 490.2 | 0.18361
21.931 | 96.664 | 659.9 | 0.18361
21.931 | 97.162 | 830.35 | 0.18361
21.931 | 97.665 | 950.2 | 0.1836
21.931 | 98.164 | 1005.6 | 0.18361
21.931 | 98.661 | 1025.45 | 0.1836
21.931 | 99.166 | 1031.95 | 0.18361
21.931 | 99.664 | 1033.85 | 0.18361
21.931 | 100.162 | 1034.5 | 0.18361
21.931 | 100.665 | 1034.7 | 0.18361
21.931 | 101.163 | 1034.8 | 0.1836
21.931 | 101.664 | 1034.8 | 0.18361
21.931 | 102.165 | 1034.85 | 0.18361
21.931 | 102.662 | 1034.8 | 0.1836
21.931 | 103.164 | 1034.85 | 0.18361
21.931 | 103.663 | 1034.85 | 0.1836
21.931 | 104.163 | 1034.8 | 0.18361
21.931 | 104.664 | 1034.5 | 0.18361
21.931 | 105.162 | 1033.5 | 0.18361
21.931 | 105.663 | 1030.15 | 0.18361
21.931 | 106.163 | 1019.45 | 0.18361
21.931 | 106.662 | 987.25 | 0.18361
21.931 | 107.164 | 904.85 | 0.18361
21.931 | 107.663 | 756.9 | 0.18361
21.931 | 108.161 | 580.9 | 0.18361
21.931 | 108.663 | 424.25 | 0.18361
21.931 | 109.162 | 303.8 | 0.18361
21.931 | 109.662 | 214.9 | 0.18361
21.931 | 110.163 | 151.05 | 0.18361
21.931 | 110.662 | 106.5 | 0.18361
21.931 | 111.163 | 75.9 | 0.18361
21.931 | 111.663 | 54.95 | 0.18361
21.931 | 112.161 | 40.5 | 0.18361
21.931 | 112.663 | 30.3 | 0.18361
21.931 | 113.161 | 23.15 | 0.18361
21.931 | 113.66 | 18 | 0.18361
21.931 | 114.162 | 14.2 | 0.18361
21.931 | 115.163 | 9.25 | 0.18361
21.931 | 116.158 | 6.25 | 0.18361
21.931 | 117.16 | 4.5 | 0.18361
21.931 | 118.162 | 3.25 | 0.18361
21.931 | 119.162 | 2.5 | 0.18361
21.931 | 120.16 | 1.95 | 0.18361
21.931 | 121.16 | 1.55 | 0.18361
21.931 | 122.163 | 1.3 | 0.18361
21.931 | 123.162 | 1.1 | 0.18361
21.931 | 124.16 | 0.8 | 0.18361
21.931 | 125.161 | 0.75 | 0.18361
21.931 | 126.163 | 0.6 | 0.18361
21.931 | 127.163 | 0.55 | 0.18361
21.931 | 128.16 | 0.45 | 0.18361
21.931 | 129.162 | 0.45 | 0.18361
21.931 | 130.163 | 0.35 | 0.18361
21.931 | 131.163 | 0.35 | 0.18361
21.931 | 132.161 | 0.3 | 0.18361