## Slide 1: Jefferson Lab Polarized Photogun
P. Adderley, J. Clark, S. Covert, J. Grames, J. Hansknecht,
R. Mammei, M. Poelker, M. Stutzman,
R. Suleiman, K. Surles-Law
TRIUMF, Vancouver, Canada, January 19, 2011

## Slide 2: Outline
Jefferson Lab
Polarized Electron Injector
Electron Spin Manipulation
Parity Violation (PV) Experiments
Polarized Electron Source:
  Photoemission from GaAs
  Fiber-based Drive Laser
  Load-Lock Photogun
      Key Features
      Vacuum and Ion Back-bombardment
Measure Lifetime vs. Laser Position and Active Area
Benefits of Higher Gun Voltage
Road to Higher Gun Voltage
Summary and Outlook

## Slide 3: Jefferson Lab Accelerator Site
CEBAF Injector
Injector
Test Cave
2 SRF Linacs (600 MeV)
3 Experimental Halls
Arcs
Arcs

## Slide 4
Continuous Electron Beam Accelerator Facility
A B C A B C
A
B
C
A
B
C
RF Synchronous
Lasers
499 MHz,
ΔФ = 120
Pockels Cell
Gun
2 Linacs
1497 MHz
RF Separators
499 MHz
B
A
C
111 ps

## Slide 5
Chopper
FC#1
FC#2
Buncher
Capture
Bunchlength
Cavity
Cryounit
Cryomodules
Synchrotron
Light Monitor
100/500 keV Mott
Polarimeter
5 MeV Mott
Polarimeter
Injection Chicane
45 MeV Dump/
Spectrometer
PreBuncher
Gun#2
Gun#3
V-Wien Filter
H-Wien Filter
Apertures
Polarized Electron Injector
Polarized Electron Source
(100 kV)
Electron Spin Manipulation
Synchronous Photoinjection
Bunching & Acceleration
500 keV
SRF Acceleration
(5 MeV)
SRF Acceleration
(65 MeV)
Spin Solenoids

## Slide 6: 4p Electron Spin Manipulation
R(z) – Two Solenoids
R(y) – Horizontal Wien Filter
R(x) – Vertical Wien Filter

## Slide 7: Parity Violation Experiments at CEBAF
Experiment | Energy (GeV) | I (µA) | Target | Apv (ppb) | Maximum Charge Asym (ppb) | Maximum Position Diff (nm) | Maximum Angle Diff (nrad) | Maximum Size Diff (δσ/σ)
HAPPEx-II (Achieved) | 3.0 | 55 | 1H (20 cm) | 1400 | 400 | 1 | 0.2 | Was not specified
HAPPEx-III (Achieved) | 3.484 | 100 | 1H (25 cm) | 16900 | 200±100 | 3±3 | 0.5±0.1 | 10-3
PREx | 1.063 | 70 | 208Pb (0.5 mm) | 500 | 100±10 | 2±1 | 0.3±0.1 | 10-4
QWeak | 1.162 | 180 | 1H (35 cm) | 234 | 100±10 | 2±1 | 30±3 | 10-4
Møller | 11.0 | 75 | 1H (150 cm) | 35.6 | 10±10 | 0.5±0.5 | 0.05±0.05 | 10-4
PV experiments motivate polarized e-source R&D

## Slide 8
Gain-switched Diode Laser and Fiber Amplifier
Attenuator
PC WP LP
Shutter
Rotatable GaAs
Photocathode
V-Wien Filter
Vacuum
Window
15° Dipole
PZT
Mirror
IHWP
RHWP
Pockels
Cell
Delayed
Helicity Fiber
HV Supply
(0 – 4 kV)
HV Supply
(0 – 90 V)
CEBAF
Hall
T-Settle
Fiber
Charge Feedback (PITA)
Electron
Beam
Helicity Fiber
Charge Feedback (IA)
LP HWP LP
IA
Target
BCM
BPMs
5 MeV
Helicity Magnets
Parity
DAQ
nHelicity
Fiber
Position
Feedback
Helicity Generator
H-Wien Filter
Spin Solenoids

## Slide 9: Electron Spin Reversal for PV
“Spin Reversal”
Vertical Wien = 90 deg
Two Solenoids = ±90 deg
“Longitudinal Polarization”
Horizontal Wien = {-90…+90}
LEFT
RIGHT
From Gun
Cancels all helicity-correlated beam systematic errors from laser and photocathode

## Slide 10: Polarized e-Source Requirements
Superlattice GaAs
QE ~ 1%
Pol ~ 85%
@ 780 nm
chekc
Fiber-based Laser
Good Photocathode
Good Laser
All three halls get polarized e-beam from the same gun and photocathode
~ 35 Wks continuous operation at high average current, high polarization
Inverted Guns
Good gun
Good Vacuum
No Field Emission
Long Lifetime

## Slide 11: Photoemission from GaAs
Bare GaAs surface
Large Work Function
Positive Electron Affinity (PEA)
PEA=4.07 eV
Egap
Conduction
Band
Valence
Band
GaAs
Vacuum
Dipole layers of
Cesium + NF3
Reduces Work Function
Negative Electron Affinity (NEA)
+
-
-
χNEA ~ 0.2 eV
GaAs
Vacuum
Light
NEA Surface Activation

## Slide 12: NEA Activation of GaAs
“Activate” GaAs photocathode by applying about one mono-layer of Cesium and NF3 to very clean surface
QE
0
NF3
Cs
R | GaAs Light Reflection Coefficient (~ 0.3)
d | GaAs layer thickness (~ 0.1 μm)
α(Һω) | Photo-absorption Coefficient (~ 5x103 cm-1)
B(χ) | Surface Tunneling Probability(~ 0.2)
Time

## Slide 13
Bulk GaAs
Laser excitation from P3/2 to S1/2 : Egap < E < Egap+Δ
Electron Polarization:
σ+: Right-handed circularly polarized light
σ- : Left-handed circularly polarized light
Reverse electron polarization by reversing light polarization
mj
E
-1/2
+1/2
J
P1/2
-3/2
-1/2
+3/2
+1/2
P3/2
-1/2
+1/2
S1/2
σ+
σ-
3
S1/2
P3/2
P1/2
Egap=1.42 eV
Δ=0.34 eV
E
k
3
1
EF

## Slide 14
Higher P: Breaking GaAs Degeneracy
Split degeneracy of P3/2 : Introduce strain on GaAs crystal by growing it on substrate (GaAsP) with different lattice constant
High polarization by laser excitation from P3/2 to S1/2 :
  Egap < E < Egap+δ
Higher QE: Alternating layers of GaAs and GaAsP – Superlattice GaAs
+3/2
Egap
δ=20–50
meV
mj = -1/2
+1/2
S1/2
P3/2
mj = -3/2
-1/2
+1/2
σ+
σ-

## Slide 15: GaAs Photocathode Evolution
QE ~ 20%, 120 μA/mW
Unpolarized @ 532 nm
QE ~ 5%, 30 μA/mW
Pol ~ 40% @ 780 nm
Bulk GaAs
QE ~ 0.2%, 1 μA/mW
Pol ~ 75% @ 850 nm
Strained GaAs:
GaAs on GaAsP
100 nm
Superlattice GaAs:
Layers of GaAs on GaAsP
No strain relaxation
QE ~ 1%, 6 μA/mW
Pol ~ 85% @ 780 nm
100 nm
14 Layers
E = hc/λ
hc = 1240 eV.nm
λ (nm) | E (eV) | Color
532 | 2.33 | Green
780 | 1.59 | NIR
850 | 1.46 | NIR
2 μm
350 μm
625 μm

## Slide 16: Fiber-based Drive Laser
1560 nm
780 nm
499 MHz
RF Locked Low-Power (1 mW)
1560 nm Fiber Diode
Frequency-doubler
1560 nm to 780 nm
30% Efficiency (2 W)
High Power (6 W) 1560 nm Fiber Amplifier

## Slide 17
Storage
Manipulators
Loading Chamber
Load-lock Photogun
Best vacuum inside HV Chamber, which is never vented except to change electrodes
Photocathode Heat and Activation takes place inside Preparation Chamber
Use “Suitcase” to replace photocathodes through a Loading Chamber
Preparation Chamber
HV Chamber
Activation Laser

## Slide 18
Manipulator Key (male)
Puck Stop (rolling)
Puck Key (female)
Puck (Mo)
Ring (Ta)
GaAs
Key Features:
5 pucks can be stored in Storage Manipulators
8 hours to heat and activate new sample
Mask to limit active area
Suitcase for installing new photocathodes (one day to replace all pucks)

## Slide 20: Imperfect Vacuum = Finite Lifetime
Ion back-bombardment – with characteristic QE “trench” from laser spot to electrostatic center of photocathode – damages the NEA of GaAs
High energy ions are focused to electrostatic center: create QE “hole”. We don’t run beam from electrostatic center
QE can be restored, but takes about 8 hours to heat and reactivate
Residual
Gas
Laser
Photocathode “QE scan”
  Active area = 5 mm
  Laser size = 0.4 mm
Can run beam from 6 locations (spots) before heating and reactivating

## Slide 21
Ion Back-bombardment

## Slide 22: Reduce Outgassing & Improve Pumping
Primary source of gas in Polarized Gun: hydrogen outgassing
To reduce outgassing:
  Reduce thick flange area
  HV Chamber 400°C bake for 10 days
Pump system combines:
  Non-Evaporable Getters (NEGs) (Zr-V-Fe)
  Ion Pumps
NEGs
Outgassing rate: 1.0×10-13 Torr.L/s.cm2
Extractor Gauge: 2.0×10-12 Torr (raw value), our lowest value
Ion Pump

## Slide 23
Injector Test Cave
SUITCASE

## Slide 25
Measure Lifetime vs. Laser Position and
Active Area
Activate with different Masks: 5 mm, 7 mm, and No Mask (12.8 mm)
Measure Lifetime from different spots on Bulk GaAs with 532 nm green laser
NF3
Leak Valve
Cs
Bias Ring
+200 V
Mask
Heater
Long Manipulator

## Slide 26
After
Activation
After Lifetime Measurements
Electrostatic Center
Spot
5 mm
7 mm
12.8 mm

## Slide 27
 | I=0 (HV = 0, 100 kV) | I = 2 mA (Bulk GaAs, 532 nm, 350 μm Laser Spot) (HV = 100 kV) |  | 
 |  | 5 mm Active Area | 7 mm Active Area | 12.8 mm Active Area
Anode Current (pA) | 0.0 ± 0.3 | 0.0 ± 0.3 | 0.0 ± 0.3 | -100 – -1000
X-ray Detector (E-2 mR/h) | 0.6 ± 0.3 | 1.5 ± 0.5 | 1.8 ± 0.5 | 3 – 7
Gun Vacuum (pA) | 0 | 0 | 0 | 0 – 30
Y-Chamber Vacuum (nA) | 3.0 | 3.0 | 3.0 | 4 – 20
This is the range of values at the start of the lifetime measurements. The anode current is proportional to the laser power.
Chart: Extractor Gauge Pressure vs. UHV Ion Pump Current
| 0.25 | 0.28000000000000008 | 0.6500000000000058 | 0.94000000000000061 | 1 | 1.1100000000000001 | 1.159999999999989 | 1.47 | 1.53 | 1.6400000000000001 | 1.9600000000000048 | 2.27 | 2.74 | 3.1 | 3.15 | 4.04 | 4.26 | 4.5 | 5.44 | 6.5 | 7.07 | 7.56 | 8.52 | 9.07 | 0.26 | 3 | 3.82 | 4.5999999999999996 | 6.1599999999999975 | 8 | 8.4 | 10.7 | 12.2 | 17.399999999999999 | 21.1 | 42.2 | 64 | 80 | 107 | 350 | 880 | 1250 | 1500 | 2620 | 3150 | 3580
 | 6.4e-12 | 6.7e-12 | 1.4e-11 | 1.76e-11 | 1.84e-11 | 2e-11 | 2.08e-11 | 2.46e-11 | 2.62e-11 | 2.78e-11 | 3.12e-11 | 3.68e-11 | 3.9e-11 | 4.24e-11 | 4.36e-11 | 5.24e-11 | 5.4e-11 | 5.54e-11 | 6.54e-11 | 7e-11 | 7.28e-11 | 7.6e-11 | 8.16e-11 | 8.5e-11 | 6.5e-12 | 3.68e-11 | 4.14e-11 | 5.22e-11 | 6.48e-11 | 7.7e-11 | 8.04e-11 | 9.5e-11 | 1.08e-10 | 1.3e-10 | 1.5e-10 | 2.46e-10 | 3.66e-10 | 6e-10 | 7.52e-10 | 1.74e-09 | 2.92e-09 | 4.28e-09 | 5.02e-09 | 7.82e-09 | 9.25e-09 | 1.02e-08

## Slide 28
Finite lifetime due to ion back-bombardment
Lesson: it is extremely important to manage ALL of the extracted beam. Beam from outside 5 mm active area can hit beam-pipe walls, degrade vacuum, reduce lifetime

## Slide 29: Benefits of Higher Gun Voltage
Reduce space-charge-induced emittance growth, maintain small transverse beam profile and short bunch-length. In other words, make a “stiff” beam right from the gun
Address current density limitation due to Child’s Law (Space Charge Limit)
Reduce problems associated with Surface Charge Limit (i.e., QE reduction at high laser power)
Prolong Charge Lifetime
Biggest Obstacle: Field Emission and HV Breakdown… which lead to bad vacuum and photocathode death

## Slide 30
Emittance and Brightness
Normalized Emittance from GaAs:
q | Bunch Charge (= 0.4 pC, 200 μA and 499 MHz)
Es | Electric Field at GaAs surface
Teff | Effective Temperature of GaAs (= 300 – 400 K, 780 nm)
kBTeff | Thermal Energy (= 34 meV)
Normalized Brightness:
Brightness is proportional to Gun HV
Gun HV (kV) | Es (MV/m) | εn (μm)
100 | 2.0 | 0.011
140 | 2.8 | 0.009
200 | 4.0 | 0.008

## Slide 31
At higher gun HV, bunchlength is shorter – when Chopper Slit is fully open, it passes 111 ps - and space charge emittance blow-up is suppressed
Transmission through Injector, improves at higher gun voltage
Better transmission prolongs lifetime:
  Less beam loss that degrades vacuum
  Lower current from gun creates less ions and thus slows NEA damage
Better transmission reduces current fluctuations and thus improves Hall measurement of helicity-correlated charge asymmetry
Chart: Electron Bunchlength vs. Gun Voltage
| 0.8 | 12 | 51 | 98 | 148 | 198
115kV | 43.771 | 60.6061 | 104.3771 | 138.0471 | 161.6162 | 188.5522
100kV | 43.8596 | 70.1754 | 96.4912 | 125.731 | 149.1228 | 152.0468 | 181.2865 | 198.8304 | 213.4503 | 233.9181
85kV | 53.8721 | 77.4411 | 107.7441 | 144.7811 | 158.2492 | 181.8182 | 208.7542 | 232.3232 | 255.8923 | 269.3603
70kV | 50.8547 | 77.7778 | 113.6752 | 152.5641 | 173.5043 | 203.4188 | 227.3504 | 245.2991 | 260.2564 | 278.2051
Chart: Transmission vs. Gun Voltage
| 11.7 | 26 | 51.4 | 74.8 | 98.2 | 122.2 | 148 | 170.5 | 198
115kV | 96.5812 | 92.6923 | 82.4903 | 73.262 | 67.0061 | 61.784 | 57.6351 | 54.8974 | 51.3131
100kV | 99.0099 | 92.7419 | 77.8431 | 66.7553 | 59.2 | 52.7911 | 47.451 | 44.2153 | 40.5941
85kV | 93.8931 | 83.274 | 67.2761 | 56.2421 | 47.6465 | 40.9161 | 35.0449 | 30.9244 | 27.2059
70kV | 77.2727 | 61.6327 | 45.2479 | 35.9712 | 29.0608 | 23.7197 | 19.8027 | 17.2973 | 14.5238
Bunchlength and Transmission

## Slide 32: Space Charge Limit at CEBAF
Child’s Law (1D):
Child’s Law (2D) (PRL 87, 278301) :
Short Pulse (PRL 98, 164802):
Maximum current density that can be transported across cathode-anode
gap is (for an infinite charge plane):
For electron emission from a finite circular spot on the cathode:
For CEBAF electron beam (499 MHz):
[A/cm2]

## Slide 33
V | Gun Voltage
D | Cathode-anode Gap (6.3 cm)
R | Laser Spot Size (0.5 mm = 2r)
Tb | Bunchlength (50 ps)
τCL | Gap Transit Time (0.96 ns at 100 kV)

## Slide 34: Surface Charge Limit
NEA of GaAs depends on Gun HV. QE increases with external Electric Field at GaAs surface, Es,
Where χ (~200 meV) is the zero-field NEA value (Physics Letters A 282, 309) and potential barrier lowering due to Electric Field is
Where εs (= 13.1) is GaAs relative permittivity.
Gun HV (kV) | Es (MV/m) | δU(Es) (meV)
100 | 2.0 | 50
140 | 2.8 | 59
200 | 4.0 | 70
χ
δU(Es)
Egap

## Slide 35
Surface Charge Limit, also known as Surface Photovoltage Effect, reduces NEA of GaAs: Photoelectrons trapped near GaAs surface produce opposing field that reduces NEA resulting in QE reduction at high laser power (LP),
Where U(LP) is up-shifting of potential barrier due to photovoltage.
For heavily Zn doped GaAs surface, U(LP) → 0
Higher Gun HV suppresses photovoltage
χ
Egap
δU(Es)
U(LP)

## Slide 36: Prolong Photocathode Charge Lifetime
Most ions created at low energy
Most ions created close to GaAs surface
Beam Current: 2.0 mA
Vacuum: 8.0 × 10-12 Torr

## Slide 37
Low energy ion column for 100 kV gun
Low energy ion column for 200 kV gun
Ion energy
At 200 kV, only 60% of ions are created compared to 100 kV, longer lifetime

## Slide 38: New Inverted Gun
e-
Medical x-ray technology
New Ceramic
Compact
$5k
Less metal at HV
No SF6 of N2
Moved away from “conventional” insulator used on most GaAs Photoguns today – expensive, months to build, prone to damage from field emission.
High gradient locations not related to beam optics, lots of metal to polish.
Less flanges material, less outgassing, better vacuum.
Old Gun
Road to Higher Gun Voltage
Old Ceramic
Exposed to field emission
Large area
Expensive ($50k)
Lots of metal at HV

## Slide 39
Inverted Gun #1 at CEBAF
HV Conditioned to 150 kV with Stainless Steel Electrode
100 kV Lifetime ~ 100 C at 180 μA with transmission of 85%
Now at 130 kV for QWeak. Lifetime ~ 100 C at 180 μA with transmission of 95%
Inverted Gun #2 at Test Cave
HV Conditioned to 225 kV, with large grain Niobium (Nb) Electrode
Performs well at 140 kV
Problematic at 200 kV. Field emission (~ nA) and vacuum bursts that ruin QE
Preparing for “Plan B”

## Slide 40
Field Emission (FE)
Electron tunneling through the Electrode surface barrier due to the presence of Electric Field is described by Fowler – Nordheim formula:
φ | Work Function of Electrode
E | Electric Field at Electrode surface
β | Field Enhancement Factor at the emission site (= 30 – 60)
ζ | Constant determined by Electrode surface properties
γ | Constant determined by Electrode surface properties
φ
EF
-eEx

## Slide 41
HV Conditioning and Field Emission
Floating Anode
connected to
Pico-ammeter
Macor glass-ceramic
spacer
100 MΩ
Conditioning
Resistor
(NICROM
100 kV, 75 W)
UNIVOLT
Transformer
Oil Tank
A Short for Beam
Delivery

## Slide 42: Anode won’t capture all FE, better to look for x-rays …
Ion Chamber Detector: Model IP 100
Energy threshold of 50 keV
Readout: CANBERRA ADM 616

## Slide 43: Charge Lifetime vs. Gun High Voltage
Poor lifetime at 200 kV due to field emission
And vacuum bursts often killed QE completely
 | I = 2 mA, Bulk GaAs (5 mm Active Area, 350 μm Laser Spot) |  | 
Gun HV | 100 kV | 140 kV | 200 kV
Anode Current (pA) | 0.0 ± 0.3 | -8 ± 1 | -520 ± 20
X-ray Detector (E-2 mR/h) | 1.5 ± 0.5 | 105 ± 5 | 5700 ± 300
Gun Vacuum (pA) | 0 | 0 | 400
Y-Chamber Vacuum (nA) | 3.0 | 3.0 – 3.2 | 3 – 4

## Slide 44: What’s Next? Plan B
Re-polished large grain Nb Electrode
HV conditioned to 225 kV; it looks very promising
x-ray stayed at background level (8 E-3 mR/h)
Added a leak valve to HV Chamber in case we need to Krypton-process the electrode to eliminate FE
Ready to run beam at 200 kV in January 2011

## Slide 45: Summary and Outlook
Jefferson Lab is a world-class polarized electron beam research facility
Group Research:
  Build 200 kV Electron Gun
  Produce 5 mA polarized beam. We hold the world record of 1 mA from Superlattice GaAs
  Electrode Research using HV Test Stand: SS and Nb electrodes with different surface treatments and coatings. FE Elimination by Krypton-processing
  Cryo-pumping to improve vacuum
  New Photocathode Materials:
    5% Polarized QE
    K2CsSb: less vacuum requirement
  Polarized Positron Source