## Slide 1: Evaluation of Bright Beams from a DC High Voltage Photogun with alkali-antimonide photocathode at high average current and high bunch charge
Yan Wang
Sept. 8, 2014

## Slide 2: Outline
The high voltage gun
  Design of the gun
  Field emission & HV conditioning
  CsK2Sb photocathode
  Preliminary HV tests
The beamline
  Design and Components
Evaluation of the beam
  Emittance, bunch length and energy spread
  Fast kick
  Low and high bunch charge beams
Study of the CsK2Sb photocathode
  Charge lifetime
  How different preparation of the photo cathode affects the QE and beam quality.
Simulations
  TOSCA, Poisson, Elegant and GPT

## Slide 3: High Voltage Gun
Design and build a DC high voltage photogun with the following characteristics:
Providing high brightness electron beams with high average current and high bunch charge.
Reliable operation at 350 – 500 kV
Minimal field emission
Inverted insulator design
Load - locked

## Slide 4: The Gun Design - Why high voltage?
Many nuclear physics experiments require electron beams with high average current (1 - 100mA), high bunch charge (~1nC) and high brightness.
High brightness beams mean low emittance.
The beam emittance degrades in the first few centimeters of acceleration due to space charge forces within the electron bunch.
The space charge effect decreases when the beam energy increases.
The higher the gun voltage the higher the initial beam energy.

## Slide 5: The Gun Design - Field Emission
Field emission is emission of electrons induced by an electrostatic field.
Occurs in high electric fields: MV/m
Originates from sharp tips and contamination of the cathode surface.
Is a primary source of vacuum breakdown and electrical discharge.
Limits the maximum achievable bias voltage

## Slide 6: The Gun Design - How to Reduce Field Emission
Proper design of the cathode and anode.
High voltage conditioning:
  Gas conditioning - apply voltage to the cathode in small increments with Krypton gas present.

## Slide 7: The Gun Design - Why Inverted Insulator
The inverted geometry does not use a large metal structure supporting the cathode electrode, i.e., less metal to generate field emission.
The inverted geometry makes the gun simpler and compact, and inexpensive to build.

## Slide 8: The Gun Design - Why Inverted Insulator
(a) Conventional cylindrical design (b) Inverted insulator design

## Slide 9: The Gun Design - Why Load-Lock
The load-lock design has been proved to be a better design:
No contamination of the gun components
Convenient and time saving
Gun Chamber
Preparation
Chamber
Valve

## Slide 10: The Gun Design - Photocathode
There are two popular photocathodes: GaAs and CsK2Sb. A CsK2Sb photocathode will be used.
CsK2Sb photocathodes have a better lifetime under the similar conditions.
GaAs photocathodes are prone to rapid QE loss due to poor vacuum, low field emission, etc.

## Slide 11: The Gun Design - Initial Design
18”
18”
6”
HV cable
Inverted insulator
Vacuum chamber
Cathode
Schematic of the Gun
Anode

## Slide 12: High Voltage Gun - Initial Tests

## Slide 13: High Voltage Gun - Initial Tests
Apparatus:
A gas-insulated Cockcroft-Walton 580kV, 5mA HVPS by Glassman. The HVPS sits inside a steel tank filled with SF6 at 10 psi.
Two industrial X-ray cables (gauges R28 and R30) with molded rubber plug by Essex X-ray.
A gun vacuum chamber with a dummy cathode.
A leak valve for Krypton.
A NEG pump, an ion pump, a Turbo pump, a RGA and a Pirani cold cathode gauge

## Slide 14: High Voltage Gun - Initial Tests
Two types of the insulators:
Type I: 19 cm 97% alumina.
Type II: 11.4 cm 94.4% alumina with lower bulk resistivity.
Cathode:
6-inch 304 stainless steel spherical ball.
Polished to mirror-like finish with silicon grit paper
Degreased and degassed

## Slide 15: High Voltage Gun - Insulators and Cathode
R30 alumina
R28 alumina with
lower bulk resistivity
R30 alumina
with shed
R30 alumina
with ZrO coating

## Slide 16: High Voltage Gun - Test Results

## Slide 17: The Beamline
No acceleration of beam after the gun
Is going to be a diagnostic beamline:
  Measure the emittance, bunch length, energy spread.
  Study the beams with low & high bunch charges.
  Test the fast kicking.
  Study the CsK2Sb photocathode:
    Charge lifetime
    Photocathode preparations vs. QE and beam quality
    Laser sizes vs. beam quality

## Slide 18: The Beamline – Basic design
Gun chamber
CsK2Sb depo
chamber
Solenoid
corrector
Valve
Brock cavity
Slit
Viewer
Dif pump
Dipole
BPM
Harp
Dump
Dump
Laser

## Slide 19: The Beamline Components
CsK2Sb deposit chamber – photocathode preparation chamber
Gun chamber – generation of electrons
Brock cavity – bunch length monitor and fast kicker
Slit – beam emittance measurement
Solenoid – beam focusing
Corrector – beam steering
Dipole – energy measurement
Harp – beam profile measurement
BPM – beam position monitor
Viewer – beam profile checking
Differential pump – maintaining vacuum
Dump – beam termination point

## Slide 20: Beamline components - Brock Cavity
Developed by Electrodynamic
Used as a bunch length monitor and a beam kicker
Brock cavity

## Slide 21: Evaluation of the beam - emittance
Beam bunch - many particles
Particles – moving with given
velocities
V = Vxi + Vyj + Vzk,
Vx ≠ 0, Vy ≠ 0

## Slide 22: Evaluation of the Beam - Emittance
The emittance is defined as
the area in phase space
including the particles.
ɛ =
The equation of the ellipse:
γx2 + 2αxx’ +βx’2 = ɛ,
the Courant and Snyder invariant

## Slide 23: Evaluation of the Beam – Emittance
Use a slit and a harp to
measure emittance:
The slit is a rectangular
opening on a metal plate
which samples the beam and
makes a beamlet which is
emittance dominated.
The harp measures beam
profile and current.

## Slide 24: Evaluation of the Beam – Emittance
The coordinates of the particles leaving the Slit at 𝑥0, 𝑖 and striking the harp wire at 𝑥1, 𝑖, 𝑗 are
(𝑥0, 𝑖, 𝑥’0, 𝑖, 𝑗),
where 𝑥’0,𝑖𝑗 = ((𝑥1,𝑖𝑗 - ) - (𝑥0,𝑖 - )) / L
and
=
= () /

## Slide 25: Evaluation of the Beam – Emittance
The beam emittance can be expressed by
ɛ2 = ˂𝑥𝑖2˃ ˂𝑥𝑖’2˃ - ˂𝑥𝑖𝑥𝑖’˃2 ------------ (I)
where
<𝑥02> = (𝑥0, 𝑖 - ]2 𝐼𝑖𝑗) /
<𝑥0’2> = (𝑥1, 𝑖𝑗 - ) - (𝑥0,1 - )]2𝐼𝑖𝑗)/(L2)
<𝑥0𝑥0’> =
( (𝑥0𝑖-)(𝑥1, 𝑖𝑗-)-(𝑥0, 𝑖-)2] 𝐼𝑖𝑗)/(L

## Slide 26: Evaluation of the Beam – Emittance
Measurement of the beam emittance using the Slit:
Scan the Slit across the beam, stopping in at least 10 regularly spaced locations across the beam diameter. This yields 𝑥0, 𝑖.
For each Slit location, scan the harp across the beamlet, histogramming its current distribution. This yields the 𝑥1, 𝑖𝑗 and 𝐼𝑖𝑗.
Calculate ˂𝑥02˃, ˂𝑥’02˃ and ˂𝑥0𝑥’02˃
Use formula (I) to calculate ɛ.

## Slide 27: Evaluation of the Beam – Emittance
(2)Use a solenoid and a harp to measure emittance:
The total transformation matrix is:
M = =
where f = 1/ (Lk) is the focal length of the solenoid, k is the strength of the solenoid.
d
L

## Slide 28: Evaluation of the Beam – Emittance
Some known formulas:
Ɛ2 = σ11σ22 - σ122 ------------------------- (II)
σ11=<𝑥𝑖2>, σ22 = <𝑥’𝑖2>, σ12 = <𝑥𝑖𝑥’𝑖>
σ1 = Mσ0MT ------------------------------- (III)
where σ = is beam matrix,
and M is a transformation matrix.

## Slide 29: Evaluation of the Beam – Emittance
Apply formula (III) to the beam from the solenoid to the harp:
σ1,11= (1+dLk)2 σ0,11 + 2(1+dLk) dσ0,12 + d2 σ0,22
=(d2Lσ0,11)k2+2(dLσ0,11+d2Lsσ0,12)k+(σ0,11+2dσ0,12+d2σ0,22)
Fitting σ1,11 with a parabola (ak2 + bk + c)
will determine the whole beam matrix σ0 by
σ0,11 = a/(d2L2),
σ0,12 = (b − 2dLσ0,11)/(2d2L),
σ0,22 = (c − σ0,11 − 2dσ0,12)/d2 .
Then we can calculate Ɛ using formula (II).

## Slide 30: Evaluation of the Beam – Bunch Length
The Brock cavity is a non-invasive bunch length monitor:
The cavity resonates many harmonic TM0N0 modes simultaneously.
Beam current I(t) = A0 + A1 cos (ω0t + θ1) + A2 cos (2ω0t + θ2) + A3 cos (3ω0t + θ3) … Each (An ) term is amplitude, each frequency term is harmonic, an integer multiple of the bunch frequency, (A0 ) is the average current.
When an electron bunch passes through a harmonic cavity, the harmonic modes are excited with amplitudes and relative phases proportional to the beam’s Fourier series.
The superposition of the harmonic modes on-axis, has a voltage vs. time waveform that ideally mirrors the current vs. time of the beam that induced it.

## Slide 31: Evaluation of the Beam – Bunch Length
Bunch length measurement:
Drive laser pulse.
The signal detected
by the Brock cavity.
(c) The FFT of the signal
with overlaid Gaussian fit.
Bunch length = FWHM

## Slide 32: Evaluation of the Beam – Fast Kicker
The brock cavity is a beam kicker when powered.
A pulse (f = f0/10) kicks one out of ten beam bunches

## Slide 33: Evaluation of the Beam – Momentum Spread
Frequently it is desirable to determine experimentally the particle momentum and momentum spread.
Basically a quad, a bending magnet and a harp are needed to measure them.

## Slide 34: Evaluation of the Beam – Momentum Spread
Individual particles in an intense beam are under the influence of strong repelling electrostatic forces creating the possibility of severe stability problems. This is the space charge effect.
Beam quality will be studied with low bunch change and high bunch charge.

## Slide 35: Evaluation of the Beam – Momentum Spread
Beam size 𝑥(𝑧) = ,
where Ɛ𝑥 is beam emittance,
and 𝛽𝑥(𝑧) is beta function.
If there is a finite momentum spread within the beam particles the beam size would be
𝑥(𝑧) = ---------- (IV)
where 𝐷(𝑧) is dispersion function,
and δ𝑝 = ∆𝑝/𝑝 is momentum spread.

## Slide 36: Evaluation of the Beam – Momentum Spread
Adjust the solenoid strength to get minimal
𝑥1(𝐿) so 𝑥2(𝐿) would be dispersion dominated at the same
solenoid setting when turning on the bending magnet.
From formula (IV) δ𝑝 = ∆𝑝/𝑝 = /𝐷(𝐿).
𝑝 = (3E-4) )/θ.
Bending magnet
Solenoid
𝐷(𝐿)
𝑥
𝑥2(𝐿)
𝐿
𝐿
θ

## Slide 37: CsK2Sb photocathode
The CsK2Sb photocathode has not been fully understood yet.
The charge lifetime, QE and beam emittance will be studied with different thickness of Sb, different laser spot sizes, low and high currents.

## Slide 38: CsK2Sb photocathode– Charge Lifetime
Charge lifetime, defined as the amount of charge that can be extracted before the QE falls to 1/e of its initial value, was measured by monitoring the QE evolution as a function of accumulated charge.
Measurements are made at a particular constant beam current.
Beam current, laser power and laser spot sizes are some of the factors affecting the charge lifetime.

## Slide 39: Simulations
Tosca and Poisson will be used to design the gun. These tools will help design better cathode geometry which affects the maximum achievable bias voltage.
Elegant and GPT will be used for the beamline design and beam transport. The simulated beam emittance and bunch length will be compared with the measured results.

## Slide 40: Simulations - Poisson
Electrostatic modeling results at350kV for (a) R30 pure alumina insulator, (b) R30 pure alumina insulator with shed electrode and (c) R28 lower bulk resistivity insulator.

## Slide 41: Simulations - Poisson
Radial electric field along the line between the plug and the insulator:

## Slide 42: Simulations - Poisson
Parallel electric field along the line between the plug and the insulator:

## Slide 43: Simulations - Poisson
Potentials along
the line between
the plug and the
Insulator is non-linear.

## Slide 44: Simulations - Poisson
Shield
The potential
Linearity is
Improved
with a shed

## Slide 45: Simulations - Poisson
The potential linearity is much better with a
different shed design.
Bad symmetry

## Slide 46: Simulations - Poisson
Good linearity
Good symmetry
E field < 10 MV/m

## Slide 47: Simulations - Poisson
The potential linearity: the lower trace is from the first shed design;
the top trace is from the second shed design, which is better.

## Slide 48: Simulations - GPT
Simulation results with I = 100uA, Q = 0.2pC

## Slide 49: Simulations - GPT
Simulation results with I = 100mA, Q = 1nC