## Slide 1
Bubble Chamber
Detector System

## Slide 2: ANL Bubble Chamber
Lead designer and fabricator Brad DiGiovine (ANL)
April 2009 Chamber Received Full Operation Authorization for (C4F10)
February 2010 First Bubble Chamber Received Upgrade Authorization for Superheated H2O
Two runs at HIγS
Modifications: N2O with Hg as buffer fluid
Modifications were reviewed at ANL
Extensive testing and operations ANL using N2O and mercury (neutron source)
Brought to JLAB for testing with photons
  JLAB Design Authority: Dave Meekins
  First user built pressure system (non-vacuum) post 10 CFR 851

## Slide 3
Theory of Operation
1 Cell is cooled then filled with room temperature gas
2 Gas is cooled and condenses into liquid
3 Once cell is completely filled with liquid, pressure is reduced creating a superheated liquid
3 Nuclear reactions induce bubble nucleation
2 High speed camera detects bubble and repressurizes
3 System depressurizes and ready for another cycle
Critical point
309
1051
1
2
3
Liquid
Vapor

## Slide 4: Bubble Growth and Quenching
100 Hz Digital Camera
Dt = 10 ms
3.0 cm
N2O Chamber with PuC neutron source
Notes: That is writing on the glass.

## Slide 5: Systems and Components
Bubble Chamber
  40-60 ml
Pressure Vessel
  Max operating pressure 1000 psi
Viewport, Camera and Lighting
Hydraulic Control
Cooling circuit/refrigerator
  -20C to room temp.
DAQ/Control and Instrumentation

## Slide 6: Basic Components
Heavy Wall Stainless Steel Pressure Vessel
Thin Wall Glass Active Liquid Volume
Thin Pressure Transfer Bellows
Cooling Coils
Pressure Supply
Solenoid Valves
High Speed Camera
P1
P2
High Speed Camera
Pressure Vessel
Cooling Coils
Active Liquid Volume
Pressure Equalizing Bellows
Pressure Supply
Solenoid Valves

## Slide 7: Bubble Chamber
Thin Glass Vessel Holds Active Liquid, N2O
N2O 40-60 ml Floats on Mercury 150 ml, Which Fills Remaining Inner Volume
Superheated Liquid Only in Contact With Smooth Surfaces
Thin Sensitive Edge Welded Bellows Equalize Pressure
Stainless Tube Facilitates External Connection of Pressure Transducers and Filling Valves
Wetted Materials: Stainless Steel, Kovar, Glass
Copper cooling coils (not shown) supplied by commercial refrigerator.

## Slide 8: Pressure Vessel
Houses Bubble Chamber
One Piece Construction
  No Welding
  Minimal Internal Volume
Machined From a Solid 304 S.S. Forging
Flanges Machined From 316 S.S.
  Utilize a Plug Design to Reduce Inner Volume
Maximum allowable design pressure 1700 psi.
  Component has 5.75” bore
  Treated as B31.3 (using 304.7.2)
  Div 2 analysis using elastic plastic
    Plastic collapse/local failure
  Fatigue screening indicates cycles required to squish bubbles not deep enough to consider for 500 psi ∆P maximum. Div 2 Part 5.5.2.4

## Slide 9: Viewport, Camera, and Lighting
Two Custom Designed and Fabricated by Industry Leader in High P&T Viewports
Design Paramaters:
– 260ºC Operating temperature
– 88 ATM (1300 psi) Operating pressure
One Houses High Speed 100FPS Camera
One Houses High Intensity LED Back Lighting

## Slide 10: Hydraulic Control System
Constructed of standard fittings and components with minimum working ratings of 1500 psi.
Piping is SST tube and flex lines with lowest design P=2500 psi
Accumulators are charged with N2 circuit design P=1000 psi
Pump supplies 1 gpm
Multiple hydraulic regulators and reliefs
Vented reservoir
Screened for cyclic loading

## Slide 11: Hydraulic Control System
Capped

## Slide 12: Overpressure Protection
Two separate circuits
Sources of overpressure
  Fire (not considered)
  Pump run away
N2 circuit
  ASME relief directly on bottle
  Orifice restricted flow to system piping
  1000 psi ASME relief with adequate flow capacity
Hydraulic circuit
  Multiple hydraulic reliefs non ASME below 1000 psi.
  ASME relief on chamber 1100 psi (12.7 gpm)
  N2O is liquid at room temp at this pressure
Detailed in TGT-CALC-502-003

## Slide 13: HazMat
Two materials of note:
  Mercury (Hg)
  Nitrous Oxide (N2O)
Admin and Eng controls in place
  Procedures for filling/operating/venting
  Hg not “handled” at JLAB under normal conditions
  Volume of N2O in chamber is 17 liters at STP
  Filters are in place to prevent Hg liquid and vapor from escaping confinement.

## Slide 14: Mercury
Prolonged or intense acute exposure can be very hazardous.
Active fluid is removed through
  Phase separator
  Droplet Filter
  Vapor Filter
This prevents Hg from escape.
There is secondary containment under the chamber should a mistake occur
  Two step error: Removing VCR Cap and opening valve
Monitor personnel for exposure using SKC Elemental Mercury Passive Sampler.
IH to monitor filling/venting
Filling and venting procedures shall only be performed by Brad DiGiovine (ANL).
  Spill kit on hand to properly respond in the event of containment loss

## Slide 15: N2O
Occupational exposure limit of 50 ppm
  2000 hr/year
If all in chamber is released to injector -> 25 ppm
Monitor personnel for exposure using Assay Technology 575 N2O sampler.
IH to monitor filling/venting
Filling and venting procedures shall only be performed by Brad DiGiovine (ANL).
Supply bottle is removed immediately after fill procedure is complete. (Total of 17 STPL in injector)

## Slide 16: Electrical Safety
System requires 208VAC power
Custom components were developed by ANL
  Detailed schematics can be found in pressure systems folder
  The system was inspected by both ANL and JLAB
All components can be disconnected from the power source with plug and secured.

## Slide 17: Cooling
Heavy Wall Copper Cooling Coils Installed
Bath Operating Temperature -20C to 20C

## Slide 18: Control and Instrumentation Chassis
Temperature Monitoring and Heater Control
Pressure and Temperature Transducer Retransmission to Computer
Solenoid Valve Manual Operation and Computer Interface
Hydraulic System Logic and Interlocks
Two Remote Override Control Interfaces
Electrical Safety Inspection Completed on All Chasses

## Slide 19: Control Chassis & Remote Overrides
Control Chassis Designed with Safety Interlocks
  Heating
  Solenoid Valves
Two Remote Override Interfaces Allow for Complete Control of System
  Solenoid Valves
  Hydraulic Pump
  Heaters

## Slide 20
Detector Control Rack

## Slide 21: Control Chassis & Remote Overrides
Control Chassis Designed with Safety Interlocks
  Heating
  Solenoid Valves
Two Remote Override Interfaces Allow for Complete Control of System
  Solenoid Valves
  Hydraulic Pump
  Heaters

## Slide 22: Failure Modes and Effects
Failure of glass bubble chamber
Failure of pump cut off switch
Accumulator failure
Regulator failure
Refrigerator failure
Power failure

## Slide 23: Failure of glass bubble chamber
  Failure can be detected by camera remotely or by direct observation.
  N2O (or C2F6) and Hg may be released into hydraulic fluid.
  The integrity of the system is not lost.
  The fluid mix can be drained back to reservoir and shipped back to ANL for recovery.

## Slide 24: Failure of Pump Cut off
Pump is controlled by pressure switch
  Failure causes pump to run continuously at 1 gpm
Multiple relief devices and control regulators will relieve pressure/excess flow back to reservoir
ASME relief set at 1100 psi will blow down entire system should other relief paths fail.
ΔP across glass will stay low due to equalizing bellows

## Slide 25: Accumulator Failure
This mode is considered very unlikely due to working pressure rating 3x higher than operating
Accumulators are diaphragm style.
There will be DAQ and control issues but system integrity will be maintained.
ΔP across glass will stay low due to equalizing bellows

## Slide 26: Regulator Failure
N2 regulator failure only affects accumulator filling actions. This is very rare.
Orifice limits flow and ASME relief with excess capacity is set at 1000 psi.
N2 bottle is disconnected after charging the accumulators

## Slide 27: Refrigerator Failure
DAQ system will alarm unless full power failure.
Interlocks will open CV-2 (high pressure valve) and close CV-4 (low pressure valve).
Active fluid will stay in liquid state while system warms to room temperature
No damage will occur to bubble chamber.
System expert shall address issues and repair as needed.

## Slide 28: Power Failure
During power failure DAQ, control, pump and refrigerator stop
CV-2 opens (default state) and CV-4 closes (default state)
Chamber is returned to high pressure “standby” mode which is stable.
The system will slowly warm.
The system will stay down until reset
Only system expert may perform reset.

## Slide 29: Required Training
To operate the detector users shall be trained by Brad DiGiovine
Only system expert shall perform fluid handling procedures or adjustments to controls
SAF 801 Rad worker I
SAF 103 ODH
SAF 130 Oil Spill Training (not required for all personnel and only if needed)
SAF 132 Tunnel worker safety
SAF 801kd RWP for tunnel access
SAF 100 General safety

## Slide 30: Operating Procedures
Following procedures are developed
  Fill procedure
  Venting
  Power loss
  Fire (leave area call 911)
  Basic user operation
    Start/stop runs
    Only available after system is placed in standby mode by expert
  DAQ failure procedure
Emergency de-energize procedure is available for all personnel but, will require that the system be shipped back to ANL for repair.

## Slide 31
User Interface

## Slide 32: Running Adjustments
Test is to understand and optimize the detector performance.
Only system expert shall make/approve changes to the fluid system
Possible adjustments include
  Operating pressure/temperature
  Quench pressure
  Fluid levels
  Dead time