## Slide 1: Gun2 with Tee electrode High Voltage push to 200 kV
C. Hernandez-Garcia
September 1 2020

## Slide 2: Present configuration.
No conditioning resistor
HVPS filled to 60 psig SF6
Electrode operating at 130kV, no field emission
Goal: Evaluate for field emission to 200kV

## Slide 3: First: Write ATLis for inserting conditioning resistor
If R30 resistor tank (Pig) ready (~4 hours)
  Ask Neil to crane the resistor tank & HV cables down to the injector tunnel
  Bubba or Carlos to remove R350-R28 cable from HVPS to gun
  Plug R350-R30 from HVPS to resistor tank
  Plug R30-R28 from resistor tank to gun
If R30 resistor tank NOT ready (~10 hours)
  Ask Neil to crane DILO cart down to the injector tunnel
  Transfer SF6 from HVPS tank to storage bottles (5 hours)
  Vent HVPS and replace short with conditioning resistor (1 hour)
  Evacuate HVPS tank (1 hour)
  Backfill HVPS to 60 psig of SF6 (2 hours)

## Slide 4: High voltage evaluation: ~ 8hr shift
Open Gun2 HVPS controls screen.
Set HV Op limit to 205kV and HVPS current limit to 0.2 mA.
Monitor DecaRad signals, Gun2 vacuum, anode current and HVPS currentas a function of HVPS voltage using LivePlot
Set the HV ramp up rate to 25 kV/min.
Set voltage to 130kV and click HV Go. Soak for a few minutes to onbtain a baseline for all the monitored signals.
Set the HV ramp up rate to 5 kV/min
Increase the voltage by 5 kV and wait for about 1 minute before increasing the voltage again by 5kV.
Continue increasing voltage up to 200kV, keeping track of field emission, anode current, HVPS current and gun vacuum levels.
Record voltage for field emission onset.
High voltage evaluation complete. Input voltage set point to 0 V and click HV Go.
Turn HVPS OFF.
Procedure to be performed under vacuum conditions