## Slide 1: 200 keV Wien filter V3
Third iteration modeled in CST by Gabriel with improved 40 kV feedthroughs has demonstrated (2019) 20kV per plate at UITF on a test stand but test still undergoing. Intended as a testbed for 20kV switch chassis test in preparation for implementation at CEBAF.
The version of the Wien filter described here is an improved version of the one tested in UITF beamline with the following changes:
40 kV CeramTec HV feedthroughs (3/4" diameter).
Spring-guide to couple to 40 kV feedthrough
Spring to match design, but we had to use some that are slightly magnetic. Danny ordered new non-magnetic springs and gave them to Phil.
HV tests have been performed without the magnet coil and without the iron yokes
October 9 2019

## Slide 2: HV switch chassis (K. Cole)
Keith is ready to start building the prototype to be tested at UITF
  He wants to see if the cable guide used by the retired oil tank PS can be utilized to draw the 30kV new cable. He would like to take opportunity now that injector is open day shift.
Has mostly everything on hand to start
Direct connection between Matsusada HVPS and switch chassis
Connection to Wien will have interlock cable
  We need to decide on the HV connection to Wien
    Mate 30kV cable directly to CeramTec 40 kV feedthrough cable, or
    Mate 30 kV cable to bayonet ‘shv’ style connecter, then that to junction box to CeramTec 40kV feedthrough cable

## Slide 3: Wien HV switch chassis wiring diagram

## Slide 4: Wien HV switch chassis schematic

## Slide 5: HV controls (G. Croke)
Digital I/O and ADC channels we'll plan to use:
Switch box control and status readback:
  Output: iocitfgun XY244 digital I/O card, output 6, PV = IGLK100DIOLWTB.B6
  Input: iocitfgun XY244 digital I/O card, input 21,PV = IGLK100DIOHRD
ADC readbacks of power supply HV:
  iocitfgun XY566 board, channel 1, PV = IGLK100AI2
  iocitfgun XY566 board, channel 2, PV = IGLK100AI3
One unresolved question is whether we want to have an 'injector steering script' for Wien plates and magnets at UITF like in CEBAF?

## Slide 6: HV controls systems block diagram (G. Croke)