## Slide 1: Near-XHV Pressure Characterization for the Jefferson Lab Polarized Electron Source Marcy Stutzman, P. Adderley, Md. A. Mamun, A.A. Elmustafa, M. Poelker
Notes: Title Page Design 3

## Slide 2: Thomas Jefferson National Accelerator Facility
12 GeV electron accelerator for Nuclear Physics
Up to 90% polarization from DC photoemission source
Electron currents to 200μA beam (CW) to four experimental halls
  Presently recommissioning machine after upgrade to 12 GeV: first beam October 2014
Hall A
Hall B
Hall C
Hall D
Source

## Slide 3: DC Photoemission Source
Strained superlattice GaAs/GaAsP photocathode
Residual gasses ionized, limit operational lifetime
SAES WP NEG modules
Gamma Vacuum SEM style ion pump
Base pressure approaching XHV
XHV ≡ P < 1x10-10 Pa = 1x10-12 mbar = 7.5x10-13 Torr
-130kV
light in
electrons out
0.36
GaAs
GaAsP
GaAs
Strained
Superlattice
Photocathode
14 layers

## Slide 4: Photocathode Lifetime
Ion bombardment – with characteristic QE “trench” from laser spot to electrostatic center of photocathode – damages NEA of GaAs
High energy ions are focused to electrostatic center: create QE “hole”
  Don’t run beam from electrostatic center.
QE can be restored, but takes about 8 hours to heat and reactivate
Photocathode “QE scan”
  Active area = 5 mm
  Laser size = 0.35 mm

## Slide 5: Outline
Outgassing rates:
Measured vs. expected pressure:
Diffusion vs. Recombination:
Temperature effects:
Three chambers: Bare, a-Si, TiN coating
Do the measured outgassing rates lead to corresponding ultimate pressures?
Where is Fick’s diffusion law applicable for predicting outgassing rate after heat treatment?
Can the temperature dependence of the outgassing rate be exploited to improve ultimate pressure?

## Slide 6: Outgassing rates vs. heat processing and coatings
Identical 304L chambers
Bare steel
TiN coating (9 µm)
a-Si (SilcoGuard-1000®~900 nm)
Outgassing rates
Spinning Rotor Gauge
Accumulation technique
Multiple bakes
Q(Troom) after each bake
Published 2014
  Md. A. Mamun, A. A. Elmustafa, M. L. Stutzman, P. A. Adderley and M. Poelker, “Effect of heat treatments and coatings on the outgassing rate of stainless steel chambers”, J. Vac. Sci. Technol. A 32, 021604 (2014);

## Slide 7: Outgassing rate vs. room temperature after each bakeout
Factor of 20 reduction in outgassing after 400°C heat treatment
Linearity on Arrhenius plot:
Temperature dependent diffusion activation energy
First tests of SilcoGuard-1000™
Reproduction of SS results @ 400°C bake
Reproduce and question TiN outgassing
Bake temperatures, in order
Stainless Steel

## Slide 8: Outgassing vs. Room temperature: a-Si and TiN coated chambers
Linearity on Arrhenius plot:
Temperature dependent diffusion activation energy
Deviations: error in outgassing setup
a-Si
TiN

## Slide 9: Outgassing rates: bare and coated chambers
Chart: Outgassing Rate (Torr L s-1 cm-2)
| First Bake 90°C, 30h | Second Bake 150°C, 30h | Third Bake 250°C, 30h | Fourth Bake 400°C, 100h | Fifth Bake 150°C, 30h | Sixth Bake 250°C, 30h | Seventh Bake 150°C, 30h | Eighth Bake 360°C, 150h
TiN | 6.44e-13 | 3.7e-15 | 2.4e-16
SilcoII | 1e-12 | 1.3e-13 | 1.5e-13
SS | 4e-12 | 1.7e-13 | 1e-13 | 2.8e-13 | 2.2e-13 | 9.9e-14
Silco | 2.4e-12 | 2.1e-12 | 1.6e-12 | 3.6e-12 | 2.8e-12 | 1.9e-12
All data at Troom = 20°C
Notes: Steel:
Expected reduction
with 400°C heat
treatment
Silco:
Little variation with
heat treatment
Silco on heated SS:
Retains prior
outgassing rate
TiN coating:
Very low outgassing
rates, consistent with
literature, elemental
Ti pumping?

## Slide 10
Chart: 
| Stainless Steel | Silco on Heated SS | TiN
Expected Pressure delta Q delta A delta S Delta P/P Delta P 3bg extractor | 8.5997e-13 | 4.5897e-13 | 7.3435e-16
S = 1680 L/s from NEG and ion pumps

## Slide 11: Gauge current contributions
  Igas: pressure dependent gas phase ions arriving at collector
  Ix-ray: x-ray induced electron desorption from collector
  Iheating: due to pressure rise in chamber with filament on
  IESD: electron-stimulated desorption of gas ions from grid

## Slide 12: XHV gauges
Extractor gauge
  x-ray limit reduced through geometry
Axtran gauge
  Bessel box energy discrimination
  electron multiplier to assist in low current measurements
Watanabe BBB (Bent Belt Beam) gauge
  230° deflector BeCu housing
  JVSTA 28, 486 (2010)
extractor
BBB
Axtran

## Slide 13: Pressure measurement backgrounds
measured
minimized
measured
Chart: Repeller voltage (V)
| 85.381 | 94.900999999999996 | 105.176 | 114.982 | 124.747 | 135.21600000000001 | 145.67400000000001 | 154.90899999999999 | 165.51 | 175.01 | 184.816 | 194.85599999999999 | 205.01900000000001 | 214.85499999999999 | 225.04900000000001 | 234.93600000000001 | 245.17 | 255.44499999999999 | 265.69 | 275.29199999999997 | 285.73 | 295.57600000000002 | 305.58600000000001 | 315.31 | 325.34100000000001 | 335.09500000000003 | 345.57400000000001 | 354.78800000000001 | 364.69600000000003 | 375.42 | 384.93 | 395.12299999999999 | 405.20499999999998 | 415.68400000000003 | 425.35700000000003
10 v steps, 5 min | 20.313 | 20.445 | 21.715 | 22.557 | 23.773 | 24.596 | 28.104 | 29.26 | 31.013 | 35.232 | 36.703 | 42.438 | 42.778 | 42.53 | 42.565 | 40.863 | 40.414 | 38.585 | 35.604 | 35.173 | 31.557 | 26.468 | 23.498 | 22.187 | 19.501 | 20.027 | 18.644 | 19.714 | 19.81 | 19.364 | 19.902 | 19.082 | 19.082 | 18.923 | 21.317
background | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12
x-ray limit
heating
ESD

## Slide 14: Coated chambers with NEG & ion pumps Measured vs. calculated pressures
Chart: Pressure (x10-12 Torr)
| bare | silco | TiN
Extractor | -0.20968 | 4.505 | 1.2903
3BG | 0.93916 | 1.6567 | 1.0143
Calculated | 0.85997 | 0.45897 | 0.00073435
Notes: Preliminary conclusions: is the sensitivity correct? Were the outgassing measurements correct? (no for TiN) was the pump speed constant through the three. Work on error bars until the extractor error is more believable. Poor signal to noise on the extractor gauge makes it difficult to use?

## Slide 15: Diffusion Limited Outgassing: bare steel
Diffusion equation:
Fick’s law:
flux is related to the outgassing,
D is the diffusion coefficient (temperature dependent)
dC/dx is the concentration gradient
Bare steel: diffusion limited below Fo = 1

## Slide 16: Diffusion Limited Outgassing: a-Si coating
a-Si:
Recombination limited for
Coated before heat treatment
Insufficient data for pre-heat-treatment

## Slide 17: Diffusion Limited Outgassing: TiN coating

## Slide 18: Does the TiN coating pump?
Ti:N ratio ~ 1:1, but inhomogeneous
Ti:N ranges from 0.58:1 through 1.17:1

## Slide 19: Can we chill to reduce outgassing, pressure?
Chart: Outgassing vs. Temperature
| 294.54999999999995 | 294.45 | 295.75
SS_250C(1) | 5.352e-12 | 5.3115e-12 | 5.78e-12
SS_400C(2) | 1.1874e-13 | 1.7889e-13 | 2.686e-13 | 5.0038e-13
SS_150C(3) | 6.1454e-14 | 1.0098e-13 | 1.8009e-13 | 2.7221e-13
SS_250C(4) | 1.6181e-13 | 2.7958e-13 | 4.5347e-13 | 6.2098e-13

## Slide 20: Ice chilling tests
+
+

## Slide 21: Chill chamber in ice water (a-Si)
Chart: Elapsed time (hours)

## Slide 22: Outgassing rate extrapolated from ultimate pressure, pump speed
Chart: T, (K)
| 288.14999999999998 | 293.14999999999998 | 298.14999999999998 | 303.14999999999998
Silco_dried(1) | 1.6443e-12 | 2.355e-12 | 3.1376e-12 | 4.426e-12
Silco_150C(2) | 1.4559e-12 | 2.1382e-12 | 2.8768e-12 | 4.1318e-12
Silco_250C(3) | 1.1845e-12 | 1.5559e-12 | 2.2602e-12 | 3.2401e-12
Silco_400C(4) | 2.6975e-12 | 3.6165e-12 | 4.9612e-12 | 6.0085e-12
Chart: T, (K)
| 292.60000000000002 | 278.93
from extractor | 2.2611e-12 | 1.2484e-12
from bbb | 1.6194e-12 | 7.8134e-13
Silco_dried(1) | 1.6443e-12 | 2.355e-12 | 3.1376e-12 | 4.426e-12
Silco_150C(2) | 1.4559e-12 | 2.1382e-12 | 2.8768e-12 | 4.1318e-12
Silco_250C(3) | 1.1845e-12 | 1.5559e-12 | 2.2602e-12 | 3.2401e-12
Silco_400C(4) | 2.6975e-12 | 3.6165e-12 | 4.9612e-12 | 6.0085e-12
Q = PS/A
Ultimate pressure variation less than anticipated from outgassing temperature dependence

## Slide 23: Heated and coated results
Measured outgassing varies by factor of 1000
  Calculated base pressures vary correspondingly
Measured pressures must be corrected for various gauge background effects
Measured and calculated pressures
  bare steel chamber: within experimental error
  Silco: measured higher by factor of ~3
  TiN: measured higher by > factor of 1000, supporting theory that TiN coating pumps

## Slide 24: Conclusions
Room temperature outgassing rates for bare steel chamber deviates from Fick’s diffusion law for Fo>1
  a-Si coating unchanged with heat treatment
  TiN coating outgassing rate drops faster than Fick?
Chilling chambers leads to decreased pressure, decrease smaller than expected
  Decrease in NEG pump speeds at reduced temperature?

## Slide 25: Acknowledgements
Matt Poelker, Joe Grames, Don (Bubba) Bullard,
Marcy Stutzman, John Hansknecht, Shukui Zhang,
Carlos Hernandez Garcia, Philip Adderley, Riad Suleiman
U.S. DOE Contract No. DE-AC05-06OR23177 and with funding from the
DOE R&D for Next Generation Nuclear Physics Accelerator
Facilities Funding Opportunity Number: DE-FOA-0000339
Jefferson Lab
Center for
Injectors and Sources