## Slide 1
Vacuum Characterization and Improvement for the Jefferson Lab Polarized Electron Source
Marcy L. Stutzman, P.A.Adderley, M.Poelker (Jefferson Lab),
Md.A.Mamun (Old Dominion University)
Residual gasses ionized, limit operational lifetime
SAES WP NEG modules
Gamma Vacuum SEM style ion pump
Base pressure approaching XHV
XHV ≡ P < 1x10-10 Pa = 1x10-12 mbar = 7.5x10-13 Torr
Improved vacuum for increased photocathode lifetime
Identical 304L chambers
Bare steel
TiN coating (9 µm)
a-Si (SilcoGuard-1000®~900 nm)
Outgassing rates
Spinning Rotor Gauge
Accumulation technique
Multiple bakes
Q(Troom) after each bake
Chart: Outgassing Rate (Torr L s-1 cm-2)
| First Bake 90°C, 30h | Second Bake 150°C, 30h | Third Bake 250°C, 30h | Fourth Bake 400°C, 100h | Fifth Bake 150°C, 30h | Sixth Bake 250°C, 30h | Seventh Bake 150°C, 30h | Eighth Bake 360°C, 150h
TiN | 6.44e-13 | 3.7e-15 | 2.4e-16
SilcoII | 1e-12 | 1.3e-13 | 1.5e-13
SS | 4e-12 | 1.7e-13 | 1e-13 | 2.8e-13 | 2.2e-13 | 9.9e-14
Silco | 2.4e-12 | 2.1e-12 | 1.6e-12 | 3.6e-12 | 2.8e-12 | 1.9e-12
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12
Axtran gauge
Bessel box energy discrimination
electron multiplier for amplification
~$8k
Acknowledgement: U.S. DOE Contract No. DE-AC05-06OR23177 and with funding from the DOE R&D for Next Generation Nuclear Physics Accelerator Facilities Funding Opportunity Number: DE-FOA-0000339
Leybold Bakable cryopump
LN2 chill circuit
  cools cryosorber panels and compressor
  allows bake of cryopump body
Operation
  10K pumping surface
  30K baffles
Presented at AVS 55 by Dieter Mueller, Leybold
Chart: Time (days)
Ionized residual gas accelerated into photocathode, limits operational lifetime
400°C heat treatment compares favorably to tested coatings
Low TiN outgassing rate due to Ti pumping, does not improve base pressure
XHV gauges require background measurement and corrections
XHV cryopump improves base pressure
  Adds surface area and complexity
  Ongoing investigation
Motivation
Outgassing Reduction
Conclusions
XHV Gauge
Characterization
XHV Pumping
Typical pumps: non-evaporable getter + Ion pump
Do the ion pumps have low pressure limitation? Could the NEG pumps or infrastructure limit base?
Can adding a commercial “XHV” cryopump improve electron gun base pressure?
Base Pressures:
Predicted vs. Achieved
Expected pressure calculations
V. Shutthanandan, …, M.L. Stutzman et al.,Phys.Rev.ST Accel.Beams 15 (2012) 063501.
a-Si
bare
TiN
a-Si
Accumulation setup
Chart: Pressure vs. time, bare steel chamber, varied room temperatures
| 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53
1March2012 20C 1st ROR | 7.62e-06 | 7.78e-06 | 7.98e-06 | 8.19e-06 | 8.45e-06 | 8.61e-06 | 8.83e-06 | 9.07e-06 | 9.27e-06 | 9.45e-06 | 9.69e-06 | 9.92e-06 | 1.01e-05 | 1.03e-05 | 1.05e-05 | 1.07e-05 | 1.09e-05 | 1.11e-05 | 1.13e-05
4March2012 38C | 5.35e-06 | 5.88e-06 | 6.44e-06 | 6.99e-06 | 7.55e-06 | 8.13e-06 | 8.6e-06 | 9.16e-06 | 9.76e-06 | 1.02e-05 | 1.08e-05 | 1.13e-05 | 1.18e-05 | 1.24e-05 | 1.3e-05 | 1.35e-05 | 1.4e-05 | 1.45e-05 | 1.5e-05
5March2012 25C | 4.79e-06 | 5.15e-06 | 5.63e-06 | 5.98e-06 | 6.29e-06 | 6.64e-06 | 6.94e-06 | 7.28e-06 | 7.7e-06 | 8.07e-06 | 8.35e-06 | 8.72e-06 | 9.07e-06 | 9.43e-06 | 1e-05 | 1.03e-05 | 1.07e-05 | 1.11e-05 | 1.14e-05 | 1.18e-05
20C early data | 2.76e-06 | 2.95e-06 | 3.18e-06 | 3.36e-06 | 3.57e-06 | 3.81e-06 | 3.98e-06 | 4.18e-06 | 4.41e-06 | 4.63e-06 | 4.83e-06 | 5.07e-06 | 5.25e-06 | 5.53e-06 | 5.7e-06 | 5.93e-06 | 6.13e-06 | 6.28e-06 | 6.55e-06 | 6.73e-06 | 6.97e-06 | 7.41e-06
Published 2014
  Md. A. Mamun, A. A. Elmustafa, M. L. Stutzman, P. A. Adderley and M. Poelker, “Effect of heat treatments and coatings on the outgassing rate of stainless steel chambers”, J. Vac. Sci. Technol. A 32, 021604 (2014)
Commercial XHV Ionization gauges require mitigation and measurement of backgrounds
x-ray limit
gauge heating effects
electron stimulated desorption
Outgassing Progression with successive heat treatments
Watanabe 3BG (Bent Belt Beam) gauge
230° deflector BeCu housing
~15.5k ($13k sensor + Leybold extractor controller)
Extractor gauge
x-ray limit reduced through geometry
~$4k
Chart: Repeller voltage (V)
| 85.381 | 94.900999999999996 | 105.176 | 114.982 | 124.747 | 135.21600000000001 | 145.67400000000001 | 154.90899999999999 | 165.51 | 175.01 | 184.816 | 194.85599999999999 | 205.01900000000001 | 214.85499999999999 | 225.04900000000001 | 234.93600000000001 | 245.17 | 255.44499999999999 | 265.69 | 275.29199999999997 | 285.73 | 295.57600000000002 | 305.58600000000001 | 315.31 | 325.34100000000001 | 335.09500000000003 | 345.57400000000001 | 354.78800000000001 | 364.69600000000003 | 375.42 | 384.93 | 395.12299999999999 | 405.20499999999998 | 415.68400000000003 | 425.35700000000003
10 v steps, 5 min | 20.313 | 20.445 | 21.715 | 22.557 | 23.773 | 24.596 | 28.104 | 29.26 | 31.013 | 35.232 | 36.703 | 42.438 | 42.778 | 42.53 | 42.565 | 40.863 | 40.414 | 38.585 | 35.604 | 35.173 | 31.557 | 26.468 | 23.498 | 22.187 | 19.501 | 20.027 | 18.644 | 19.714 | 19.81 | 19.364 | 19.902 | 19.082 | 19.082 | 18.923 | 21.317
background | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408
Gauges used in combination
Gauge performance
Base chamber pressure
  Cryopump
  Outgassing test chambers
Add 1340 L/s NEG /ion pumps
Characterize all gauge limitations
Determine base pressure
Bare and Silco agree with calculated base pressure
TiN pressure ~1000x too high
Elemental Ti excellent pump for hydrogen
XPS analysis shows inhomogeneous TiN coating
Net pump speed of 0.003 L/s in accumulation measurement would explain discrepancy
Chart: Pressure (x10-12 Torr)
| bare | silco | TiN
3BG | 0.93916 | 1.6567 | 1.0143
Calculated | 0.85997 | 0.45897 | 0.00073435
Chart: Pressure (x10-12 Torr)
| bare | silco | TiN
3BG | 0.93916 | 1.6567 | 1.0143
Calculated | 0.85997 | 0.45897 | 0.00073435
Outgassing reduction with 400°C bake
First test of a-Si “Silco” coating
TiN coating: remarkably low outgassing
Does it yield similarly remarkable base pressure?
 | NEG, Ion pump only | NEG, Ion and Cryopump
Getter pump (l/s) | 2150 | 2150
Ion pump (l/s) | 45 | 45
Cryopump (l/s) | - | 800-1300 (conductance dependent)
Area (cm2) | 8000 (chamber only) | 8000 (chamber) + 3400 (bellows) + 13700 (gate valve)
Outgassing (TorrLs-1cm-2) | 1x10-13 | ≥ 1x10-13
Expected pressure | 4 x 10-13 Torr | 3 - 7 x 10-13 Torr
Measured pressures | 4x10-12 Torr | 1.5x10-12 Torr
Moderate improvement when cryopump added
Significant infrastructure required: gate valve, LN2 plumbing circuit, vibration isolation
Further work to understand discrepancy between predicted and measured pressures
-130kV
light in
electrons out
Strained
Superlattice
Photocathode
Structure
0.36
GaAs
GaAsP
GaAs
14 layers
For high polarization (near 90%),
Strained superlattice GaAs/GaAsP
How to improve vacuum?
Less gas in:
Reduce outgassing rate
Better gas removal:
XHV pumping
Measure improvements:
XHV gauge characterization
Backgrounds | X-ray | Heating | ESD
Extractor | ~1x10-12 Torr | ~1x10-12 Torr | degas well
3BG | 0 | ~1x10-12 Torr | degas well
AxTran | 0 | ~1x10-12 Torr | energy discrimination
extractor
3BG