## Slide 1
>200 keV Wien filter spin rotator
Gabriel G. Palacios Serrano
gabrielp@jlab.org

## Slide 2
Introduction
Thomas Jefferson Nation Accelerator Facility (2019)
Experimental halls
RF acceleration
Bending magnets
Refrigeration
Photogun
+
Wien
Aligned Spin!

## Slide 3
Wien filter
B. Steiner, W. Ackermann. (2007)
The design used at CEBAF is originally from SLAC, and has been replicated at MAMI, MIT-Bates, and S-DALINAC.

## Slide 4
Model description
CEBAF’s 100 keV Wien filter
High voltage plates
High voltage feedthrough
Magnetic mirrors
Vacuum chamber
Mounting insulators
Magnet coil
Steel shield
42 cm
E
v
B

## Slide 5
Two objectives
Do not deviate the beam trajectory
Provide Spin angle control
Beam trajectory
100 keV -> 200 keV

## Slide 6
Beam trajectory
Magnetic field
Beam trajectory
Don’t deviate the beam
Equilibrium condition:
Preserve this relation:

## Slide 7
Electrostatics simulations
Solidworks models
CST microwave studio

## Slide 8
Beam dynamics
Done with Pulsar physics GPT
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]

## Slide 9
We have built four Wien filters at JLab:
  Two 200 keV at CEBAF
    In operation since February 2021
    32 cm electrode plate
    1.5 cm gap
    20 kV max achieved voltage per plate
    2.7 MV/m max gap electric field
    Undergoing optics matching studies.
  A shorter-gap 300 keV at UITF
    In operation since July 2022
    32 cm electrode plate
    1.3 cm gap
    26 kV max achieved voltage per plate
    3.9 MV/m max gap electric field
  One with integrated ion pump at test lab
    Patent!
    Still on the works
Experiments
Jefferson Lab (2019)

## Slide 10: 300 keV Wien spin rotator
KE [keV] | beta | E [MV/m] | Vapp [+-] | B [mT]
300 | 0.7765 | 4.3873 | 21.9368 | 18.8963
Is it possible with the existing set up?
Operation parameters for 300 keV beam at 10 mm gap and 100°spin rotation

## Slide 11
Model
Isometric view showing magnetic mirrors (blue)
At 10 [mm] gap, we are at a shorter gap than the mirror aperture. But stay tuned for beam dynamics.
10 [mm] gap
Frontal view with transparent mirrors to show gap

## Slide 12
Electric field magnitude
+21.9368 [kV]
-21.9368 [kV]
< 3.0 [MV/m]
< 3.5 [MV/m]

## Slide 13
Beam dynamics
Based on the beam dynamics digest by the Darmstadt team
Done with Pulsar physics GPT
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Test Gaussian distribution beam @300 keV
The 10 mm gap is just fine!

## Slide 14
Finally, for 300 keV Wien
Caveats:
  Magnetic field requires more current, larger cross section wire?
    Care must be taken on the feedthrough pipe
  Magnetic mirror aperture optimization
  Electrode Rogowski profile optimization
  Electrode manufacturing
  Must adjust HV spring guide, spring coil, and electrode mounting spring (or mounting insulators)
Also: for 100° spin rotation
  How about other gaps?
Gap [mm] | KE [keV] | beta | E-field [MV/m] | Bias at plate [kV] | B- field [mT]
10 | 300 | 0.7765 | 4.3873 | 21.9368 | 18.8963
13 | 300 | 0.7765 | 4.3873 | 28.5178 | 18.8963
15 | 300 | 0.7765 | 4.3873 | 32.9052 | 18.8963
HV spring guide
Electrode mounting spring

## Slide 15: FIN
gabrielp@jlab.org

## Slide 16
References
K. Tsuno & D. Ioanoviciu. Advances in Imaging and Electron Physics 176. Academic Press (2013)
M. Salomaa & H. A. Enge. Nuclear Instruments and Methods 145 (1977)
B. Steiner, W. Ackermann. Proceedings of PAC07, Albuquerque, New Mexico, USA. (2007)
J. Grames et al. Proceedings of 2011 Particle Accelerator Conference, New York, NY, USA. (2011).
V. Tioukine, K. Aulenbacher. Nuclear Instruments and Methods in Physics Research A 568 (2006)
MOLLER Collaboration. arXiv:1411.4088 [nucl-ex], (2014). https://moller.jlab.org/moller_root/
J. D. Jackson. Wiley, (1975)
J. Grames et al. Physical Review Special Topics - Accelerators And Beams,Volume 7, 042802 (2004)
Solidworks https://www.solidworks.com
CST Studio Suite, “CST Microwave Studio,” 2019. http://www.cst.com
H. Miller, “Surface flashover of insulators,” IEEE Transactions on Electrical Insulation, vol. 24, no. 5, 765 (1989).
G. Palacios et al. Review of Scientific Instruments 89, 104703 (2018)
Pulsar Physics General Particle Tracer http://www.pulsar.nl/gpt

## Slide 17: Extra slides
Just in case

## Slide 18
Provide Spin angle control
Relativity requires Thomas-BMT to describe the change of spin with respect to time:
Magnetic moment anomaly
So the spin rotation angle:
For our case:
Wait, what?
z
y
t
z
0=0°
y
Final angle t
Initial angle 0
Notes: Bargmann–Michel–Telegdi

## Slide 19: Extras
Using the calculated value of E field = 1.65 MV/m
kicks beam out of z-axis
Using 97 % E (1.61 MV/m) restores the trajectory

## Slide 20: Extras
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]

## Slide 21: Extras
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]

## Slide 22
Beam dynamics
Remember me?
Kick due to fringe fields!
Done with Pulsar physics GPT
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
“No transmission”
Chart: avg z [m]
Chart: avg z [m]
Chart: avg z [m]
Fy
Fz
vs transmission (97% Ey)*
*Produced with Gaussian distribution test beam

## Slide 23
Extras
Modelling
  The Wien filter is composed of
    HV electrode plates with mounting tabs and Rogowski profile
    HV spring guide contacts with spring coils
    HV electrical feedthroughs with alumina insulators
    MACOR mounting insulators
    Vacuum chamber with electrode mounting springs
    Magnet coils
    Magnet mirrors, pole plates and return yokes
Isometric view of 3D model of the existing CEBAF 130 keV Wien filter based on SLAC’s original design.

## Slide 24
For 100° spin rotation angle and we know the energy:
  = 1.3914
  The normalized velocity:
  And we obtain the magnetic field magnitude (at homogeneous region):
Extras
Will use this to get Ey
Kinetic energy
Relativistic factor

## Slide 25
We use the magnetic field in the equilibrium condition:
  So the potential per plate for a 0.015 [m] gap:
Extras
To keep the beam trajectory
Which defines the power supply set point

## Slide 26
Extras
Buy/manufacture
Test, repair
Assembly/S&A
Bake, vacuum check, troubleshoot
Final product!

## Slide 27
Extras
Monitor on EPICS
Data processing with
Python
Visualization and analysis
Done!