Here’s are input conditions, from PR’s, drawings, as-founds, e-mails:

- Schneedle

  - JLAB dwg 200-3000-0219

  - Dimameter 0.742” to 0.744”

- Moly Puck

  - JLAB dwg JL0047248

  - Keith uses drawing as-is

- Indium Foil

  - PR = https://misportal.jlab.org/reqs/prs/293572#ShowItemsTab

  - Thickness from PR = 0.002”

  - Keith used thickness = 0.001”

- Wafer (micron)

  - AXT (from PR) = 625 +/- 25

  - SVT (spec) = 625

  - UCSB = 625

  - RIT = 350

  - Keith uses 650 um = 0.026”

- Tantalum cup

  - Drawing JLAB dwg 32709 indicates 0.005” thick

  - Keith uses as-found measurement says 0.004” thick

Keith created a 3d model of the stackup of a photocathode on a puck with a schneedle with highlighted values above.

The distance from left face of Ta cup (green) to Puck crown is 0.001”

- Ideally, we’d like this to be close to zero, so pretty close

  - If puck 600 um => gap = 0.003”

  - If puck 625 um => gap = 0.002”

  - If puck 650 um => gap = 0.001”

The distance from right face of Ta cup (green) to Puck (orange) is 0.0066” = 168 um, which is greater than any of the stack-up tolerances

- Our wafers sit proud of the orange piece, although too high and this can lead to cracking the wafer.

- Do we recall the design goal?:

- My intuition says make that gap as short as possible given tolerances

Finally, we should have a design for a “625” and “350” micron thick.

Schneedle

Marcy’s mark-up of Moly puck dwg

Tantalum cup print

CIS-TN-18-01

March 1, 2018

FABRICATION OF MOLYBDENUM PUCK BODY

CIS

Jefferson Lab

This note summarizes the steps for fabricating the molybdenum “PUCK BODY” as described in drawing JL0047248.

- Machine and Assembly Steps

  - Machine puck, key and dowel pin

  - Trial fit-up of puck, key and dowel pin

  - Initial cleaning of components with micro and acetone rinse (do not acid etch before key is pinned)

  - Install key and pin into puck body

- Acid Etch Steps

  - Remove indium if needed (isopropyl + indium ball to get indium off)

  - Moly Etch Aqua Regalis (1:3 HCl:Nitric), 10 seconds, rinse, check surface, repeat for 10 more seconds if needed, rinse, check surface

  - Ultrasonic cleaning in DI water after (USC twice to ensure removal of acid from around ear)

  - Rinse, air dry in hood, bag for heat treatment

- Vacuum Bake Steps

  - Use a vacuum oven to degas puck.

  - Use a ramp rate of 10-15 degrees Celsius per minute.

  - Ramp to 900 degrees Celsius.

  - Soak for 120 minutes, recording pressure every 30 minutes.

  - Ramp and cool to room temperature

  - Bag puck, if possible in GN2.