## Slide 1: UITF Operator Training – SAF162
Matt Poelker, March 5, 2020

## Slide 2: SAF 162 Training Outline: PowerPoint slides + Tour
Part 1: UITF Safety Awareness Training
  Geographical info, documentation, ODH, etc.,
Part 2: UITF Operator information
  What to expect at the controls
Part 3: Tour of UITF

## Slide 3: Part 1: UITF Safety Awareness Training
UITF overview: capabilities and uses
About UITF: geographic location, regions Cave1 Cave2 rooftop, fire extinguishers, pull stations, phones, dosimeter requirements, etc.,
Rooftops Cave1 and Cave2: electronics racks, names and locations, access restrictions
UITF utilities: electrical power, LCW, GN2 and compressed air
Control Room
ODH
UOD, ASE and FSAD
“Commissioning”, then “Operating” the UITF accelerator
Beam Authorization
Credited Controls and Defense in Depth: current limits at UITF
PSS description, including Sweeping the enclosure
Laser Light
Laser Bypass mode
Making UITF safe after operation – shutting down and rapid access beacon

## Slide 4: UITF, what can it be used for?
Application | Beam Energy | Beam Current | Experiment Duration | Notes | Presenter
Commission QCM for CEBAF | 6 MeV, but prefer up to 10 MeV | up to 100 uA | three or four 1-week long tests | tests complete before long shutdown of 2020, when QCM to be installed at CEBAF | R. Kazimi
Commission HDIce for CEBAF | ~ 8 MeV | up to 100 nA for tuning, 0.25 to 5 nA for production | four or five run periods, two-weeks long each | target provides transverse polarization required for 3 A-rated Hall B experiments | A. Sandorfi
Manufacturing polarized targets for CEBAF via DNP | 10 to 18 MeV | 1 to 10 uA | hours, days | likely some R&D to determine optimum polarizing conditions | C. Keith
Bubble Chamber astrophysics | 4 - 10 MeV | 0.01 to 100 uA | 3 weeks, ~ 3 runs/year | UITF better location than CEBAF injector, when CEBAF shutdowns are short | R. Suleiman
MeV parity violation experiment | 10 MeV | milliamps preferred, will reduce experiment duration | months to years | requires polarized electron beam, transmission geometry offers advantages | R. Carlini
Testing Nb3Sn-coated cavities | determining the beam energy of test cavity is point of test | up to 100 uA | as many tests as possible | Nb3Sn cavities require only 4K Helium | G. Eremeev
Wastewater treatment | 2- 10 MeV | 100 uA | imagine week-long test durations over three years | together with local partners | G. Ciovati
Polarized positron source | 5 - 10 MeV | up to 100 uA | staged tests, likely many required, 1-week long duration | requires polarized electron beam | J. Grames
EIC: fast kicker tests | 5 - 10 MeV | up to 100 uA | two 1-week long tests | together with sbir-partner | H. Wang
EIC: testing high bunch charge | 5 - 10 MeV | up to 100 uA | two 1-week long tests | requires polarized electron beam | J. Grames and J. Guo
from a workshop December 12, 2018

## Slide 5: UITF is a 10 MeV spin-polarized electron accelerator
Today, UITF is used as a test bed to evaluate new hardware destined for CEBAF. Examples include Wien filters, 200 kV photogun, new booster to replace CEBAF ¼ CM, HDIce target, PSS beam current monitor, new I-to-V amplifier system, non-invasive polarimeter, etc.,

## Slide 6: Geographic Features of UITF
Located inside Test Lab building 58 high bay
CMTF
VTA
TL1137
Physics workspace
CM fabrication
CTF refrigerator

## Slide 7: Cave1, Cave2, Control Room, Laser Room, Mechanical Room, Labyrinth
Bubba’s office
Room 1125
Room 1126A
Room 1127
Room 1126B
There is fire suppression inside these rooms, and in the UITF enclosure, i.e., sprinklers

## Slide 8: Cave1, Cave2, Control Room, Laser Room, Mechanical Room, Labyrinth
Bubba’s office
Room 1125
Room 1126A
Room 1127
Room 1126B
there’s good cell phone reception in these areas

## Slide 9: Cave1 rooftop, Cave2 rooftop, moveable shielding
Cave2 rooftop
Tiles are movable!
labyrinth
Cave1 rooftop
Helium vent
There is NO fire suppression in building 58 Test Lab High Bay, there are NO sprinklers in the high bay. There IS fire suppression inside the UITF enclosure, control room and laser room.

## Slide 10: See how roof tiles are moved? Test the spin-polarized target “HDIce”
Roof tiles are “credited control”: we must lockout radiation sources when tiles are moved. There’s an OSP for this

## Slide 11: UITF rooftop, electronics and storage….
Cave2 rooftop
Cave1 rooftop
In between rows of racks, there are penetrations in the roof tiles, for the electrical cable to devices inside the enclosure

## Slide 12: UITF utilities: electrical power, LCW, GN2 and compressed air
Cave2 rooftop
Cave1 rooftop

## Slide 13: Access Restrictions, dosimeter requirements
Inside UITF enclosure:
no dosimetry required inside UITF
Signs indicate no drilling into mildly radioactive concrete walls
UITF Rooftops
Cave1 rooftop can be accessed even during operation but must have Lab-issued dosimeter
No access permitted to Cave2 rooftop during accelerator operation
Magenta beacons flashing = Accelerator in operation

## Slide 14: Entering the Enclosure when OPEN, no dosimetry required
Access to the Enclosure permitted only when PSS in OPEN/Safe state
Magenta beacon OFF

## Slide 15: Magenta Beacons
There are four magenta beacons: at each entrance, and two on the Cave1 rooftop between electronics racks
SAF 162: UITF Operator Training

## Slide 16: UITF Rooftops when Magenta Beacons are ON
Cave2 rooftop
Cave1 rooftop
Cave1 rooftop access to electronics racks: YES with dosimeter
but NO access to Cave2 rooftop

## Slide 17: UITF Control Room
Cave2 rooftop
Cave1 rooftop
Operators at left, Users at right
You will need key Z3AA1X to unlock the door

## Slide 18: Oxygen Deficiency Hazards (ODH) at UITF
ODH is a serious concern at UITF
There are two cryogenic liquids inside UITF enclosure: LHe and LN2
There are two noteworthy room temperature gases: GN2 and SF6
For elevation < 9’, the enclosure was designed to ensure ODH0 conditions unless u-tubes are being stabbed, or when HDIce brings large dewars of LHe into the enclosure
We don’t “take credit” for the two exhaust fans inside UITF
The booster vent stack is plumbed outside the UITF enclosure
In case of large LHe spill, passive vents ensure safe ODH0 conditions inside UITF enclosure
Chimney stacks ensure safe ODH0 conditions at Cave1 rooftop electronics racks
Flow limiting orifice installed in the room temperature GN2 gas supply lines
Memory foam inserted in ceiling penetrations and trenches
Safety System Group has installed an ODH monitoring system: (3) heads for helium, (1) head for nitrogen
There is not enough SF6 inside the high voltage power supply tank to pose an ODH problem

## Slide 19: Oxygen Deficiency Hazards (ODH) at UITF
Liquid Helium (LHe): ~ 100 L in the booster cryostat, ~ 500 L in HDIce dewar
Liquid Nitrogen (LN2): 50 L HDIce
Gaseous Nitrogen (GN2): boil off from dewar outside Test Lab
Sulfur Hexaflouride SF6, gun HV power supply tank to suppress corona and HV discharge

## Slide 20: Oxygen Deficiency Hazards (ODH) at UITF
and ODH-1 when u-tubes are being stabbed
Conditions are described on white board at entryway

## Slide 21: Oxygen Deficiency Hazards (ODH) at UITF
We don’t “take credit” for ventilation
Blowers need electricity and this might not be available when things go wrong
Volume of SF6 is small compared to volume of room
SF6 is heavy compared to air, it would occupy space < 1” thick on floor
Orifice used to restrict GN2 flow to UITF

## Slide 22: Oxygen Deficiency Hazards (ODH) at UITF
Cave1: two penetrations provide passive venting, note tall chimneys to keep workers on Cave1 rooftop safe
Cave1: the other penetrations are blocked to provide safety for people working on Cave1 rooftop, working inside the electronics racks

## Slide 23: Oxygen Deficiency Hazards (ODH) at UITF
Large passive vent inside Cave2 (fans in photo not required for ODH safety)

## Slide 24: Oxygen Deficiency Hazards (ODH) at UITF
We vent the circle-seal relief valve out of the UITF enclosure, and we direct the flow of gas from parallel plate relieve valve upward

## Slide 25: Oxygen Deficiency Hazards (ODH) at UITF
All of the trenches are packed with foam to restrict the flow of heavy gasses to occupied areas, like the control room
(trenches also packed with sand and lead brick, to provide radiation shielding)

## Slide 26: Oxygen Deficiency Hazards (ODH) at UITF
ODH monitoring system, built and maintained by Safety System Group
Periodic checks of the heads
Three heads for helium, one head for nitrogen
ODH beacons and sirens

## Slide 27: UITF Accelerator…also a “Gun Test Stand”
keV
MeV
Gun Test Stand here…..
“accelerator” over here…..
The “mode” matters, Gun Test Stand versus Accelerator. How will you know which mode we are in? More on this topic later when we discuss “beam authorization”

## Slide 28: Official Documentation
https://wiki.jlab.org/ciswiki/index.php/UITF_Official_Documents
https://www.jlab.org/eshq/ProgramDocs
These documents were approved by Lab Leadership and TJSO

## Slide 29: Radiation Shielding Assessment and Commissioning Plan
https://wiki.jlab.org/ciswiki/index.php/UITF_Official_Documents
https://wiki.jlab.org/ciswiki/index.php/UITF_Safety_Documents
Shielding is adequate for 100 uA CW beam

## Slide 30: Procedures: preferred way of doing business at UITF
https://wiki.jlab.org/ciswiki/index.php/UITC_Operations
https://wiki.jlab.org/ciswiki/index.php/Document-Controlled_Procedures_and_Assessments_for_UITF

## Slide 31: Procedures: preferred way of doing business at UITF
https://wiki.jlab.org/ciswiki/index.php/UITF_Procedures_not_managed_via_document_control
if you Operate UITF, you are going to contribute procedures

## Slide 32: UITF Accelerator…also a “Gun Test Stand”
keV
MeV
Gun Test Stand here…..
“accelerator” over here…..
The “mode” matters, Gun Test Stand versus Accelerator. How will you know which mode we are in?

## Slide 33: Beam Authorization for Accelerator use

## Slide 34: Beam Authorization Rescinded: Lock-out the valve upstream of booster
MeV possible
MeV beam not possible

## Slide 35: Credited Controls
From UOD…

## Slide 36: Credited Control: Radiation Shielding
The grating is credited control too

## Slide 37: Credited Control – lead brick and other moveable shielding

## Slide 38: Defense in Depth: CARMS
There’s a CARM in the control room, in trench under floor

## Slide 39: Personnel Safety System (PSS)
PSS chassis inside the UITF Control Room. No computer console, no epics readback
Modes: OPEN, SWEEP, RUN
Lights on the PSS chassis display status and system alarms
Laser Bypass mode

## Slide 40: Sweeping the Cave
Time-out interval from control room to door closure

## Slide 41: Laser Light

## Slide 42: Laser Light

## Slide 43: Laser Bypass mode
With Laser Bypass key inserted into the Safety Interface Chassis, the laser can be turned ON in CW mode (i.e., it bypasses the PSS, opens the PSS laser shutters)
In Laser Bypass mode, the gun and booster CANNOT be energized
Why do this?
Can align the laser in CW mode with the enclosure OPEN (e.g., align the pockels cell, the tune mode generator)
Needed for QE scans with gun biased at low voltage

## Slide 44: Laser Bypass mode also provides safe environment for laser alignment inside UITF
When in Laser Bypass mode, and when the lid to the laser box is OPEN inside the UITF enclosure, “light barriers” at both entrances will trip OFF laser (close a shutter)
Safety cones with labels inform would-be UITF occupants that laser alignment is taking place

## Slide 45: Rapid Access System at UITF
Gamma and neutron probe, inside the enclosure, near the booster
Test button, make sure the lights function properly
OK for “High” and “Alert” indicators to illuminate when making beam
But we don’t Open the enclosure unless lights turn OFF when beam is terminated

## Slide 46: Rapid Access System at UITF
CARM at exit of booster
Neutron probe above booster

## Slide 47: Making UITF Safe after Beam Operations
Turn OFF sources of radiation: the photogun and the booster cryomodule RF
Wait for the Rapid Access indicators to show “safe”
  If Alarm indicators persist after X minutes, call RadCon for sweep
Open the entry way gate

## Slide 48: Take a break

## Slide 49: Part 2: UITF Operator Information
JMenu for UITF
Gun modes
Gun HV
Drive laser
Valves and vacuum
Viewers
Magnets used at UITF including field maps and dipole calibrations (I versus B)
Beam position monitors
Current monitoring: cups, dumps, apertures, BCMs
RF systems: field control chassis, drive laser, choppers, buncher, two cell, seven cell
Controls at UITF are mostly identical to controls at CEBAF, this is intentional

## Slide 50: Part 2: UITF Operator Information, cont.
Booster: SRF and Cryo features related to CTF
MPS description, including clearing faults, masking FSD nodes
Save and restore
Archiver
UED
Channel Access
User accounts, logging on
Signs!
Emergency response

## Slide 51: Part 2: UITF Operator Information
Find an open LINUX terminal and logon using your Accelerator username and password
….then from the Redhat icon, click JMenu
….now stamp in

## Slide 52: Part 2: UITF Operator Information “Drive Laser”
Familiar laser controls screen
Access to other items from this screen: STAC5 stepper control, FSD overview, Expert Beam Mode, etc., These other screens are available from other pages too
At UITF, there is no master control of gun mode, but software was copied from CEBAF. So you must click twice on each gun mode for the change to be made

## Slide 53: Part 2: UITF Operator Information “Gun High Voltage”
Familiar Gun high voltage screen
This will only work if you have swept and set PSS to Run mode
PSSA and PSSB should be green
SF6 pressure should be green
You might need to set the dipole to the correct current (from an all-save, for 200 keV beam)
Turn ON the gun HV
Enter setpoint voltage: 200 kV
Hit the HV Go button

## Slide 54: Part 2: UITF Operator Information “Magnets”

## Slide 55: Part 2: UITF Operator Information “Vacuum”

## Slide 56: Part 2: UITF Operator Information “Viewers”

## Slide 57: Part 2: UITF Operator Information “QE scans”
Lock out the gun HV
Open the valve to the prep chamber and insert the manipulator to ground the cathode electrode
Insert the laser power meter, Open the gun valve
Take laser by-pass key and insert it in the PSS interface chassis rack ITF26, turn key
Open the screen Expert Beam Mode screen for laser control page: Viewer Mask OFF, and CW enable
Go to CW mode and set laser attenuator to ~ 120 to provide ~ 1uA current measured at anode
Hit QE scan button

## Slide 58: Part 2: UITF Operator Information “BPMs”
“stripline” BPMs are used at UITF
(antenna BPMs used at CEBAF injector)

## Slide 59: Part 2: UITF Operator Information “Current monitoring: cups, dumps, apertures”
Expert screen

## Slide 60: Part 2: UITF Operator Information

## Slide 61: Part 2: UITF Operator Information