## Slide 1: R30-4 gun S&A during electrode assembly installation and laser retro-reflection evaluation
C. Hernandez-Garcia, G. Palacios-Serrano, M. Stutzman, B. Mares, C. Gould, S. Hardisty, R. Ammons
April 2024

## Slide 2
R30-4 back end zoomed in showing orientation of puck retainer assembly to be used for installing puck with ear (tab) up
Up
Puck keys horizontal
Puck tab vertical at 12 o’clock orientation
Puck retainer spring at 6 o’clock orientation
R30-4 gun puck retainer assembly orientation
R30-4 gun electrode assembled on April 25, 2024 with 16/30 front end

## Slide 3
D = 1380 mm
Y reflected laser spot displacement (mm)
Reflected laser spot measuring plane (card with pinhole)
Puck plane
(mirror or GaAs)
Reflected light angle is 2 X electrode angle
The electrode z-axis is tilted down 0.124 deg with respect to chamber z-axis
S&A reports that the electrode z-axis is 0.124 deg pointing down wrt to the gun HVC z-axis
532 nm laser beam path
S&A setup to evaluate laser retro-reflection with R30-4 electrode installed to gun high voltage chamber

## Slide 4
Laser colinear with R30-4 gun high voltage chamber z-axis
Mirror puck ear up
GaAs wafer on Moly puck # 35 ear up
Mirror puck 𝜃=
Y = 6 mm
D = 1380 mm
𝜃 = 0.249∘
Electrode z-axis angle = 𝜃/2 = 0.124∘
Agrees with S&A measurement
GaAs on Moly puck # 35. 𝜃=
Y = 15 mm
D = 1380 mm
𝜃 = 0.624∘
GaAs puck # 35 z-axis angle = 𝜃/2 = 0.312∘
𝜃 = 0.124 ∘
High voltage chamber
Electrode & mirror puck
High voltage chamber
Electrode & puck # 35
𝜃 = 0.312 ∘
Is the Moly puck z-axis tilted?
Or is the GaAs wafer plane tilted wrt to the moly puck z-axis?

## Slide 5
Ear to the left
Ear up
Ear to the right
Laser colinear with R30-4 gun electrode z-axis
Mirror Puck

## Slide 6
Laser colinear with R30-4 gun electrode z-axis
GaAs on Puck # 35
Ear to the left
Ear up
Ear to the right
The rotation (same r = 9 mm) of the reflected laser spot indicates the GaAs plane is not parallel to that of the puck

## Slide 7
𝜃 = 0.124 ∘
𝛽 = 0.312 ∘
𝛼 = 0.312- 0.124 ∘
𝛼 = 0.188 ∘
𝛼 = 0.188 ∘
Y= 15 mm
X = Y∙tan (0.188 ∘) = 0.048 mm or ~ 50 microns
Indium foil is 50 microns thick
A 50 micron gap offset in the GaAs wafer wrt to the puck plane explains the observed laser retro-reflection spot displacement and rotation