## Slide 1: Foil thickness measurements
5June2015

## Slide 2: FE-SEM images

## Slide 3: Measure using ImageJ
measurement | Angle | Length
1 | 94.988 | 0.546
2 | 95.232 | 0.543
3 | 95.648 | 0.543
4 | 96.108 | 0.54
5 | 95.69 | 0.539
6 | 95.711 | 0.537
7 | 95.877 | 0.541
8 | 96.34 | 0.538
9 | 95.583 | 0.529
10 | 95.369 | 0.529
11 | 95.46 | 0.539
12 | 95.251 | 0.541
13 | 95.44 | 0.543
14 | 95.232 | 0.543
15 | 95.382 | 0.549
16 | 95.176 | 0.548
17 | 93.705 | 0.551
18 | 93.474 | 0.555
19 | 94.899 | 0.556
20 | 94.574 | 0.546
21 | 94.145 | 0.548
22 | 94.145 | 0.548
23 | 94.383 | 0.544
24 | 94.493 | 0.556
25 | 94.493 | 0.556
26 | 94.042 | 0.561
27 | 95.102 | 0.556
28 | 94.847 | 0.562
29 | 93.45 | 0.559
30 | 93.414 | 0.565
31 | 92.612 | 0.564
32 | 93.191 | 0.569
33 | 92.37 | 0.574
34 | 91.968 | 0.576
35 | 92.532 | 0.582
36 | 92.135 | 0.584
37 | 92.218 | 0.562
38 | 91.206 | 0.564
39 | 91.608 | 0.564
40 | 94.295 | 0.552
42 | 91.206 | 0.529
43 | 96.34 | 0.584
 | AVERAGE | 0.552738
 | STDEV | 0.014422
 | MIN | 0.529
 | MAX | 0.584

## Slide 4: Measured vs. Lebow thicknesses
Chart: Lebow thickness (nm)
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 750 | 750 | 625 | 625 | 625 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 355 | 355 | 355 | 355 | 355 | 225 | 225 | 225 | 225 | 50 | 50 | 50
all decent data | 971.556 | 928.094 | 916.089 | 968.0135 | 941.007 | 966.716 | 749.135 | 752.624 | 808.417 | 809.928 | 807.487 | 861.042 | 876.638 | 844.427 | 824.635 | 836.0366 | 782 | 772 | 539 | 576 | 552 | 508 | 430.7273 | 465 | 463 | 534 | 494 | 477 | 477 | 489 | 361 | 388 | 394 | 391 | 400.61 | 211.068 | 215.514 | 217.646 | 223.289 | 51.254 | 53.553 | 51.824
 | 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200
Foil | Lebow (nm)
5385B | 1000
3057C | 870
5134B | 750
7028B | 625
5275C | 500
5613A | 350
7029B | 225
6809B | 50

## Slide 5: Weighted means of data
Chart: Lebow thickness (nm)
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 750 | 750 | 625 | 625 | 625 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 355 | 355 | 355 | 355 | 355 | 225 | 225 | 225 | 225 | 50 | 50 | 50
weighted averages | 954.5461 | 827.5667 | 776.3014 | 555.1243 | 482.0916 | 390.1635 | 215.6478 | 52.1871
 | 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200

## Slide 6: Systematic errors
Foil tilt
  0.2% of measurement at tilts up to 2.5 degrees
Pixel resolution of image
  assume within 4 pixels of actual edge
Film uniformity
  quoted 1%/cm by Lebow
Judgement on edges
  repeat measurements

## Slide 7: Foil tilt analysis
First foils measured were not carefully controlled
  Roll
  Pitch
Later measurements
  Roll set to ~zero
  Pitch varied intentionally to study effect
3057C, ~870 nm
0 deg pitch | -1.7 deg pitch | +2.5 deg pitch
749 nm | 808 nm | 861 nm

## Slide 8
Chart: Foil 3057: nominally 870 nm Tilt analysis
| 0 deg | 0 deg | neg 1.7 | neg 1.7 | neg 1.7 | 2.5 deg | 2.5 deg | 0 deg | 0 deg | 0 deg | spot 1 | spot 1 | spot1 | spot1 | spot1 | spot 1 | spot 2 | spot 2 | spot 3 | spot 4 | center | center | center | center | center | center | center | edge | edge | edge
 | 749.135 | 752.624 | 808.417 | 809.928 | 807.487 | 861.042 | 876.638

## Slide 9
Chart: Foil 3057: nominally 870 nm Tilt analysis
| 0 deg | 0 deg | neg 1.7 | neg 1.7 | neg 1.7 | 2.5 deg | 2.5 deg | 0 deg | 0 deg | 0 deg | spot 1 | spot 1 | spot1 | spot1 | spot1 | spot 1 | spot 2 | spot 2 | spot 3 | spot 4 | center | center | center | center | center | center | center | edge | edge | edge
 | 749.135 | 752.624 | 808.417 | 809.928 | 807.487 | 861.042 | 876.638
Zero suppressed for detailed look
Tilts of 1.7-2.5 degrees should reduce measured thickness ~0.2%
  about 1.5 nm at 870 nm
  tilted foils measure thicker?
Differences larger than the expected tilt error
  focus/magnification issues with tilt
  Image measurement issue only?

## Slide 10: Position dependence: 3 spots along same foil
Chart: Foil 3057: nominally 870 nm
| 0 deg | 0 deg | 0 deg | spot 2 | spot 3 | spot 4 | edge | edge | edge
 | 844.427 | 824.635 | 836.0366
All consistent with an average of 835 nm within error bars – no strong position dependence
Lebow uniformity quote 1%/cm implies 1x10-4%/ µm
Foil thickness measurements here not sensitive enough to test uniformity

## Slide 11
Chart: Asym (%)
| 5.219E-2 | 0.21565000000000001 | 0.38800000000000001 | 0.48209000000000002 | 0.55511999999999995 | 0.77629999999999999 | 0.82757000000000003 | 0.95455000000000001
Timing cut asym | 43.4 | 40.92 | 39.16 | 38.61 | 37.21 | 35.61 | 34.61 | 33.65
fit | 43.3127 | 41.2005 | 39.1855 | 38.1665 | 37.4114 | 35.2965 | 34.8399 | 33.7584
Asymmetry analyzed with timing cuts only, measured foil thicknesses

## Slide 12: Continuing work
Re-analyze outlier data points
Work to add any magnification/focus errors from FE-SEM into analysis
Consider re-measuring some of the foils if needed (for example, only two images on 750 nm foil, 3 on 50 nm foil)
Incorporate more refined asymmetry into asym. vs. thickness plot