AIPINJ – Phase 1

Magnet Lead Wiring Diagrams

Version 6

- Version 6 – Added Joe Meyer’s slides for wiring the FX solenoids

- Version 5 – Posted this version to ATLIS originally

- Version 4 – Has all wiring information except for FX

- Version 3 – Dr, FG, Wiens added

- Version 2 – QW drawing corrected/updated

- Version 1 - Original

This document describes the wiring diagrams for:

- QW = (4) quads

- HD = (4) steering magnets

- DR = (1) dipole

- FX = (3) CW solenoids

- FG = (2) SW solenoids

- WF = (2) Wien filters (Horizontal and Vertical)

- BH = (many) nested Haimson horizontal/vertical steerers

“QW” Quads on Viewer Crosses

There are (4) quad magnets.

Each quad magnet is composed of 4 coils mounted on a vacuum cross.

Each coil has 2 leads, for a total of 8 leads per magnet.

Magnet leads should be dressed out to a terminal block for jumpering.

Q1 = upper left, Q2=lower left, Q3=lower right, Q4=upper right

“HD” Steering magnets on A1 & A2 Cubes

There are (4) steering magnets, mounted on the A1 & A2 cubes.

Each magnet is composed of 2 opposing coils.

Each coil has 2 leads, for a total of 4 leads per magnet.

The (4) leads for each magnet should be dressed to a terminal block.

Wiring diagram shows VERTICAL magnet using UPPER and LOWER coils

For HORIZONTAL magnet replace UPPER=>LEFT and LOWER=>RIGHT

“DR” 15 deg dipole (looks like old DS dipole)

There are (1) of these dipoles, mounted on the Y-chamber.

The DR dipole has two halves (upper and lower):

A terminal block should be affixed to upper coil half (tapped holes provided) and the leads may terminate there.

The jumpering at the termainal block is like this:

  - Dipole red/black leads are for +/- of PS

  - Two “Jumper” leads are jumpered together

“FX” Green Solenoid Magnet

There are (3) of these green solenoids.

“FG” Solenoids

There are (2) FG solenoids.

Each solenoid has two leads.

The leads should be dressed out to a terminal block.

The connections should be made as shown in picture with +/- going to the +/- polarities of the PS.

“WF” Wien Filters

There are (2) Wien filters, we call “Vertical” and “Horizontal”.

Each Wien filter has two coils, each with two leads.

The four leads should be dressed to a terminal block.  Tapped holes are provided.

The leads are marked as in the picture for connection like this:

- The +/- leads to the PS are shown

- The jumper leads between coil halves are labeled “A”

“MBH” Haimson Nested Steering Coils

The technical note attached describes the wiring diagram.