## Slide 1
Steigerwald results reported in SPIN 2000
JLab measurements in June 2012
after 12 years good agreement
Results depend on (energy=same, targets=same, TOF cut=same)
Do we agree at a precision level?
similar at few percent, but what about <1% ??? => motivates experiment

## Slide 2
Significance of target thickness
Our vendor Lebow claims 10% absolute and 2% relative to foil
dS(d)/S(d) = (-b * ad) / ( 1+ ad), where b = dd/d

## Slide 3
dS(d)/S(d) = (-b * ad) / ( 1+ ad), where b = dd/d
0 – 10 mm
0 – 1 mm
b
2%
4%
6%
8%
10%

## Slide 4
Thermionic SEM images :
encouraging, consistent with vendor claim
in future will co-deposit onto Si wafer
employ field emission SEM (1 nm = 0.001 mm resolution)

## Slide 5: Initial electron beam condition before and after the target

## Slide 6
105 e- pass 1um Au target & illuminate 1” thick dump : Cu, Al, C, Be
Particles entering target (z = -1)
Particles exiting target (z = 1, R=12.7)
 | Cu 29 | Al 13 | C 6 | Be 4
e- | 105+65 | 105 +66 | 105+63 | 105+62
g | 226 | 228 | 227 | 231
 | Cu 29 | Al 13 | C 6 | Be 4
e- | 105 | 105 | 105 | 105
g | 0 | 0 | 0 | 0
Pencil Beam
Pz = 5.5
Px = Py = 0
sx = sy = 0.5
spx = spy = 0
Units are millimeters and MeV/c unless stated
Electrons are RED and Photons are GREEN unless stated

## Slide 7: Electron and photon spatial distributions in front of dump (going to or coming from)

## Slide 8
Pz>0
Pz<0
Copper (Z=29)
1 mm upstream of Dump

## Slide 9
Pz>0
Pz<0
Aluminum (Z=13)
1 mm upstream of Dump

## Slide 10
Pz>0
Pz<0
Carbon (Z=6)
1 mm upstream of Dump

## Slide 11
Pz>0
Pz<0
Beryllium (Z=4)
1 mm upstream of Dump

## Slide 12: Electron and photon momentum distributions in front of dump returning toward target

## Slide 13
Number of Electrons
Total Momentum [MeV/c]
Cut: Pz < 0
Cu
Al
C
Be
Electrons with Pz < 0
1 mm upstream of Dump

## Slide 14
Cu Z=29
Photons with Pz < 0
1 mm upstream of Dump
Al Z=13
Be Z=4
C Z=6
Number of photons
Number of photons
Number of photons
Number of photons
Total Momentum [MeV/c]
Total Momentum [MeV/c]
Total Momentum [MeV/c]
Total Momentum [MeV/c]

## Slide 15: Electron and photon spatial distributions exiting downstream of dump

## Slide 16
Cu Z=29 (37 e-)
Pz > 0 (no events with Pz < 0)
1 mm downstream of Dump
Al Z=13 (69 e-)
Be Z=4 (17 e-)
C Z=6 (21 e-)
 | Cu | Al | C | Be
e- | 10 | 189 | 337 | 466
g | 1428 | 2649 | 1966 | 1521
radial edge of Dump

## Slide 17: Response versus Z

## Slide 18
total
total
upstream
upstream
downstream
downstream
sidewall
sidewall
Electrons
Photons
Electrons and Photons traversing surfaces about 1mm outside of a 1” dump flange
upstream means returning from dump
downstream means exiting dump
sidewall means exiting OD surface

## Slide 19: Electron and photon back-illuminating target plane

## Slide 20
Assuming “illuminous dump” is uniform in backward 2pi then 324cm2 detector plane 1.8m from dump represents 1.5% angular acceptance. Testing for Aluminum…
1.5% * 1644 = 25
1.5% * 10084 = 151
target
So, what reaches the target plane ?

## Slide 21
Summary
C or Be are similar, both better than Al and by
  about 3-4 for e-
  about 2-3 for photons
Cu is better absorber of gamma by factor of about 2
  Could be a good backing material, but should not be proud in chamber
Simulation rates fine for characterizing dump; poor for back-illuminate target
  105 target events (per second) is 16nA or 106 overnight still low stat
  Target – Dump distance is about 1.8m so need to bias simulation
More realistic model will be helpful
  Use a reasonable e- beam distribution
  Verify multiple-scattering distribution
  Include dump dipole
  About 90% of e- reach the dump, so more careful detail on surfaces
  Figure out how to integrate deposited energy using G4Beamline

## Slide 22
Mott Progress Report on Schedule, Target Ladder and Beam Dump Plate
Joe Grames
August 14

## Slide 23
Present work schedule
  Sept 28 HCO begins
  Nov 1 Mott tunnel work complete
  Nov 4 2K recovery begins
  Nov 8 Mott commissioning plan ready
  Nov 9 HCO ends
  Nov 9 Injector setup begins
  Nov 11 2K recovery ends
  Nov 11 Pulsed beam to FC1
  Nov 16 Pulsed beam to FC2
  Nov 16 Mott ready for beam
  Nov 21 Pulsed beam to INJ spectrometer
  Nov 24 Pulsed beam to 0R dump (end of injector)
  Nov 28 Run through Thanksgiving
  Dec 18 1-pass pulsed beam at 2.2 GeV
  Dec 20 Winter shutdown
  Feb 5 ACC-II begins
  May 7 ACC-II ends
Schedule

## Slide 24
Target Ladder
Budget approved for new target ladder parts and assembly.
Gears and fixture plate for precision position control designed and ordered
STAC-5 stepper motor controller software written and testing = good
Target ladder (0.16”) should be ordered today
Viewer machining and adapter plate requirements known, yet incomplete

## Slide 25
Dump Plate : motivation to reduce backscatter
“…angular distribution of backscattered electrons and the backscattering coefficient were measured for Cu, Ag, and Au targets of various thicknesses at the incident energy of 6.08 MeV, and for Be, C, Al, Cu, Ag, Au, and U targets of effectively semi-infinite thickness in the energy range 3.24-14.1 MeV.”
From Leo

## Slide 26
Dump Plate : motivation to reduce backscatter
From same paper, angular distributions for semi-infinite thickness targets.

## Slide 27
Dump Plate : energy loss
200keV – 8MeV energy loss is nearly all collisional at 1.6 MeV-cm2/g
KE | Uniform Loss (1.6 MeV-cm2/g) | Katz and Penfold Rev. Mod Phys 24 (1952) 28 | Practical Range (rBe = 1.848 g/cm3) | Turn-Around Range
3 MeV | 1.88 g/cm2 | 1.48 g/cm2 | 0.8cm | 0.40cm
5 MeV | 3.12 g/cm2 | 2.54 g/cm2 | 1.4cm | 0.70cm
8 MeV | 5.00 g/cm2 | 4.14 g/cm2 | 2.3cm | 1.15cm
10 MeV | 6.25 g/cm2 | 5.14 g/cm2 | 2.8cm | 1.4cm

## Slide 28
Dump Plate : worst case thermal distribution
I asked Dave Meekins if he would givs us a first pass look at thermal distribution given uniform energy loss in volume (beam radius * E/(dE/dx).

## Slide 29
Existing | Upgrade | Labor | Procure | Status
Collimator position unknown | Survey collimator | S&A | None | Complete
Competing detector acceptance | Test w/ nosepiece removed Modify lead + new nose | TBD | TBD | Riad/Marty to advise
Vacuum window thickness SPIN2000: 0.05mm = 2 mil | Marty suggests 4 or 6 mil better | TBD | TBD | TBD
Detector (Air Side)
vs.