## Slide 1: Extreme High Vacuum: why and how Marcy L. Stutzman Jefferson Lab Center for Injectors and Sources
Notes: Title Page Design 3

## Slide 2: What is vacuum?
Vacuums are nothings. We only mention them to let them know we know they're there.
Middle school student’s answer on a science test

## Slide 3: Modern Vacuum History
‘Ultimate’ Vacuum vs. Time
Improvements with pumps
Improvements with gauges
P.A. Redhead Vacuum 53 (1999) 137
10-14
Ultimate Vacuum (Torr)
10-6
10-12
10-10
10-8
1900
1920
1940
1960
1980
2000

## Slide 4: Vacuum level definitions
Vacuum | Pressure (Torr) | Molecules/m^3 (bottom of range) | Mean free path
Atomsphere | 760 | 2.7x1025 | 68 nm
Low vacuum | 300 – 1 Torr | 1022 | 100 µm
Medium vacuum | 1 Torr – 1 mTorr | 1019 | 100 mm
High vacuum | 1 mTorr – 1x10-7 | 1015 | 1 km
Ultra High vacuum | 1x10-7 – 1x10-12 | 1010 | 105 km
Extreme high vacuum | <1x10-12 | <1010 | >105 km

## Slide 5
Atmosphere
Medium Vacuum
Vacuum Level | Where found | If people were molecules, the your closest neighbor is
Atmosphere | This room | About 1 foot away
Medium Vacuum | Bell jar, scattering chambers | Near the Newport News airport, almost 3 miles
High Vacuum | Low earth orbit, beamline | At the moon, 240,000 miles
Ultra High Vacuum (UHV) | Lunar atmosphere, Electron gun | Near Pluto, 4 billion miles
Extreme High Vacuum (XHV) | Interstellar space, Cryomodules | At the nearest star, Alpha Centauri, 25,000 billion miles
High Vacuum
UHV
XHV

## Slide 7: Beam required at Jefferson Lab
24 Hours
High polarization for 3 (4) experimental halls simultaneously
24/7 beam operation
Beam Current: 200 µA
Charge Delivered: 16 Coulombs polarized electrons / day
Outline
How do we emit polarized electrons?
What is required to keep the electron source working for the duration of experiments?
How do we achieve and maintain the required vacuum

## Slide 8: Photoemission Source
-130 kV DC (vs. RF) electrode bias
x-ray standard “inverted” insulator
Pumps with NEG modules and ion pump
Base pressure approaching XHV ≡ P < 1x10-12 Torr
-130kV
light in
electrons out
photocathode

## Slide 9: Metals, Insulators and Semiconductors
gap
filled
unfilled
unfilled
Metal:
Unfilled valance band
Insulator:
Large band gap
Semiconductor:
Small band gap
gap
filled
unfilled
gap
filled
unfilled
EF
EF
EF
Fermi
Level, EF

## Slide 10: Photoelectric effect
No emission below cutoff wavelength
Increase in photocurrent with increase in intensity of light above cutoff wavelength
Metallic emission: typically UV wavelengths
Particle nature of light
The Nobel Prize in Physics 1921 was awarded to Albert Einstein "for his services to Theoretical Physics, and especially for his discovery of the law of the photoelectric effect".

## Slide 11: Photoemission from GaAs
Bare GaAs surface
Large Work Function
Positive Electron Affinity (PEA)
PEA=4.07 eV
Egap=1.4 eV
Conduction
Band
Valence
Band
GaAs
Vacuum
Dipole layers of Cs + NF3 or Oxygen
Reduces Work Function
Negative Electron Affinity (NEA)
+
-
-
χNEA ~ 0.2 eV
GaAs
Vacuum
Light
NEA Surface Activation
GaAs Spin Polarized Electron Source, D.T. Pierce, R.J. Celotta, G.-C. Wang, W.N. Unertl, A. Galejs, C.E. Kuyatt, and S.R. Mielczarek, Review of Scientific Instruments 51(4), 478 (1980).

## Slide 12: NEA Activation of GaAs
“Activate” GaAs photocathode by applying about one mono-layer of Cesium and NF3 to very clean surface
QE
0
NF3
Cs
Time
GaAs
Cs
F
F
F
F
Cs
Cs
Cs
Vacuum must be “deep-UHV” ≡ < 5x10-11 Torr
System baked to remove water vapor
Leak tight, all metal seals, no organics
Heat GaAs prior to activation to remove any surface oxides
Maintain surface activation layer
water vapor, oxygen destroy surface chemistry

## Slide 13: Polarized electrons using circularly polarized light
Crystalline (bulk) GaAs
Circularly polarized light
780 nm ~ band gap
Helicity reversal: 1kHz
Maximum theoretical polarization: 50%
GaAs Spin Polarized Electron Source, D.T. Pierce, R.J. Celotta, G.-C. Wang, W.N. Unertl, A. Galejs, C.E. Kuyatt, and S.R. Mielczarek, Review of Scientific Instruments 51(4), 478 (1980).

## Slide 14: High P: Breaking GaAs Degeneracy
Split degeneracy of P3/2 : Introduce strain on GaAs crystal by growing it on substrate (GaAsP) with different lattice constant
High polarization by circularly polarized laser excitation from
P3/2 to S1/2 : Egap < E < Egap+δ
Higher QE: Alternating layers of GaAs and GaAsP
Strained Superlattice GaAs
+3/2
Egap
δ=20–50
meV
mj = -1/2
+1/2
S1/2
P3/2
mj = -3/2
-1/2
+1/2
σ+
σ-
Bulk GaAs
Max 100% theoretical polarization

## Slide 15: Strained Superlattice QE and Polarization
+3/2
Egap
δ=20–50
meV
mj = -1/2
+1/2
S1/2
P3/2
mj = -3/2
-1/2
+1/2
σ+
σ-
Highest polarization at wavelength just over the band gap
Higher QE with shorter wavelengths
J.L. McCarter, … M.L. Stutzman et al., NIMA 618 (2010) 30.

## Slide 16: GaAs Photocathode Evolution
QE ~ 20%, 120 μA/mW
Unpolarized @ 532 nm
QE ~ 5%, 30 μA/mW
Pol ~ 35% @ 780 nm
Bulk GaAs
QE ~ 0.2%, 1 μA/mW
Pol ~ 75% @ 850 nm
Strained GaAs:
GaAs on GaAsP
100 nm
Superlattice GaAs:
Layers of GaAs on GaAsP
No strain relaxation
QE ~ 1%, 6 μA/mW
Pol ~ 85% @ 780 nm (1.59 eV)
100 nm
14 Layers
2 μm
625 μm
Timeline of Polarized Source at CEBAF
0.36
GaAs
GaAsP
GaAs
0.36
GaAs
GaAsP
GaAs
350 μm
350 μm

## Slide 17
  Any gas in chamber can be ionized by electron beam, accelerated back toward the photocathode and
  limit photocathode operational lifetime

## Slide 18: Photocathode Lifetime
Ion bombardment – with characteristic QE “trench” from laser spot to electrostatic center of photocathode – damages NEA of GaAs
High energy ions are focused to electrostatic center: create QE “hole”
  Don’t run beam from electrostatic center.
QE can be restored, but takes about 8 hours to heat and reactivate
Photocathode “QE scan”
  Active area = 5 mm
  Laser size = 0.35 mm

## Slide 19: Operational Lifetime Limitations
Static vacuum approaching XHV (1x10-12 Torr)
  Baked, leak free, all metal vacuum systems
    Ceramic insulators, glass windows bonded to metal
    No machining oil, no fingerprints
    Bake for 30 hours 250°C to largely eliminate water vapor
  Hydrogen outgassing reduction
  Pumping for nearly XHV
Dynamic vacuum issues
  Eliminate field emission from electrodes
  Eliminate beam halo, scraping

## Slide 20
How to reach XHV ≡ P < 1x10-12 Torr?
Less gas in
More gas out
Measure pressure

## Slide 21: More gas out: improve pumping
Only some pumps suitable for UHV
Ion pumps
  High voltage to ionize gas
  Magnetic field to direct ionized gas into plates to trap gasses

## Slide 22: Another proven UHV pump
Getter Pumps
  Chemically active surface
    Titanium sublimed from hot filament
    Non-Evaporative Getters
  Molecules stick when they hit
    Does not work for inert gasses such as Argon, Helium

## Slide 23: A novel system: Leybold XHV Cryopump
LN2 chill circuit
  cools cryosorber panels and compressor
  allows bake of cryopump body
Operation
  10K pumping surface
  30K baffles
Presented at AVS 55 by Dieter Mueller

## Slide 24
True gas ionization
  Positive current
Cathode
Anode
1
+
3
e-
2
e-
X-ray effect
  e- on anode -> photons emitted
  Photons on collector -> electrons emitted
  Extra positive current
Inverse X-ray effect
  e- on anode -> photons
  Photons on walls -> electrons
  Electrons to collector
  Extra negative current
4
+
Electron stimulated desorption
  e- strike gas molecules on anode
  Gas ionized and reaches collector
    can be distinguished with energy analysis
  Neutral atom desorbed
    Ionized within grid
    Indistinguishable from real gas
    Must eliminate souce
Measure Pressure:
Hot cathode gauge operation and errors

## Slide 25: XHV gauges
Extractor gauge
  available commercially for decades
  x-ray limit reduced through geometry
  x-ray limit quote: 7.5x10-13 Torr
Axtran gauge
  Bessel box energy discrimination
  electron multiplier to assist in low current measurements
  Quoted limit: 3.75x10-13 Torr
Watanabe BBB (Bent Belt Beam) gauge
  Uses Leybold IE540 controller
  230° deflector BeCu housing
  JVSTA 28, 486 (2010)
  Quoted limit: 4x10-14 Torr
extractor
BBB
Axtran
Hot filament

## Slide 26: x-ray limit measurement
Extractor Gauge
Chart: Repeller voltage (V)
| 85.381 | 94.900999999999996 | 105.176 | 114.982 | 124.747 | 135.21600000000001 | 145.67400000000001 | 154.90899999999999 | 165.51 | 175.01 | 184.816 | 194.85599999999999 | 205.01900000000001 | 214.85499999999999 | 225.04900000000001 | 234.93600000000001 | 245.17 | 255.44499999999999 | 265.69 | 275.29199999999997 | 285.73 | 295.57600000000002 | 305.58600000000001 | 315.31 | 325.34100000000001 | 335.09500000000003 | 345.57400000000001 | 354.78800000000001 | 364.69600000000003 | 375.42 | 384.93 | 395.12299999999999 | 405.20499999999998 | 415.68400000000003 | 425.35700000000003
10 v steps, 5 min | 20.313 | 20.445 | 21.715 | 22.557 | 23.773 | 24.596 | 28.104 | 29.26 | 31.013 | 35.232 | 36.703 | 42.438 | 42.778 | 42.53 | 42.565 | 40.863 | 40.414 | 38.585 | 35.604 | 35.173 | 31.557 | 26.468 | 23.498 | 22.187 | 19.501 | 20.027 | 18.644 | 19.714 | 19.81 | 19.364 | 19.902 | 19.082 | 19.082 | 18.923 | 21.317
background | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408 | 21.408
Bent Belt Beam Gauge
Chart: Deflector voltage (V)
| 84.983000000000004 | 89.492999999999995 | 95.216999999999999 | 99.768000000000001 | 104.86 | 110.655 | 114.982 | 120.57299999999999 | 125.634 | 130.73599999999999 | 135.328 | 140.34800000000001 | 145.64400000000001 | 150.30699999999999 | 155.184 | 160.26599999999999 | 164.77600000000001 | 169.72399999999999 | 174.59100000000001 | 180.26499999999999 | 184.816 | 189.82599999999999 | 194.815 | 199.53899999999999 | 205.15100000000001 | 209.6 | 214.83500000000001 | 220.13
3BG31July2013c | -0.896 | -0.777 | 0.593 | 0.068 | -0.372 | -0.942 | -0.56 | -0.942 | 0.959 | 3.165 | 8.951 | 22.336 | 23.51 | 26.686 | 26.598 | 27.171 | 26.82 | 24.874 | 23.827 | 23.587 | 20.655 | 18.987 | 15.92 | 12.836 | 9.892 | 8.155 | 4.943 | 3.189
1Aug2013c | -2.217 | -2.015 | -1.433 | -1.983 | -1.614 | -1.245 | -1.04 | -1.718 | -1.666 | 1.702 | 9.296 | 19.034 | 24.498 | 26.288 | 0 | 27.273 | 26.8 | 25.239 | 24.748 | 22.495 | 20.626 | 17.823 | 16.92 | 12.962 | 11.379 | 7.691 | 6.384 | 3.917 | 1.485 | 1.028 | 0.867 | 0.167 | -0.198 | -0.005 | -0.028 | -0.257 | -0.916 | -0.655 | 0.697
5-Aug | -2.141 | -0.445 | -0.616 | -0.186 | 0.531 | -1.762 | -0.492 | -0.957 | -0.018 | 2.264 | 7.857 | 19.282 | 25.038 | 28.272 | 27.187 | 27.076 | 26.488 | 25.398 | 23.501 | 23.45 | 21.717 | 20.151 | 17.646 | 14.883 | 11.377 | 7.791 | 4.666 | 3.808 | 1.375 | 1.625 | 0.108 | -0.731 | -1.928 | 0.206 | 0.195 | -0.041 | -0.112 | 1.503 | 0.405
extractor back on | -1.06 | -0.108 | 0.468 | -0.249 | -2.083 | -0.381 | -0.542 | 0.582 | -0.315 | 2.079 | 8.736 | 19.289 | 25.349 | 27.684 | 27.806 | 28.181 | 26.13 | 26.001 | 24.21 | 24.452 | 21.961 | 20.271 | 16.861 | 13.62 | 10.78 | 7 | 5.526 | 2.461 | 1.233 | 0.3 | 0.331 | 0.606 | -0.239 | -0.292 | -1.197 | -1.136 | -1.142 | -2.154 | -1.53
x-ray background
x-ray background

## Slide 27: Gauge heating effects
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12

## Slide 28: Days to stabilize: minimization of ESD effects?
Chart: Days
Extractor:Bare Steel | 2.2e-11 | 6e-12 | 3.83e-12 | 3.52e-12 | 2.58e-12 | 2.29e-12
BBB:Bare Steel | 1.58e-11 | 4.3e-12 | 4.61e-12 | 3.86e-12 | 3.57e-12 | 2.8e-12 | 2.46e-12 | 2.5e-12 | 3.8e-12 | 4.84e-12 | 2.96e-12 | 2.5e-12 | 2.96e-12
Bare Steel
Silco on Heated
TiN

## Slide 29: Effectiveness of adding a cryopump
Chart: Time (hours)
dummy secondary y axis | 3 | 6 | 9 | 12 | 15 | 18 | 21 | 24 | 27 | 30

## Slide 30: Less gas in: reduce outgassing
Primary source in stainless system: hydrogen outgassing
To reduce outgassing
  reduce thick flange area
  heat treatment
    400°C bake
    10+ days
thin dome replaces thick flanges

## Slide 31: Outgassing rates vs. heat processing and coatings
Identical 304L chambers
Bare steel
TiN coating (9 µm)
a-Si (SilcoGuard-1000®~900 nm)
Outgassing rates
Spinning Rotor Gauge
Accumulation technique
Multiple bakes
Q(Troom) after each bake

## Slide 32: Outgassing rates: bare and coated chambers
Chart: Outgassing Rate (Torr L s-1 cm-2)
| First Bake 90°C, 30h | Second Bake 150°C, 30h | Third Bake 250°C, 30h | Fourth Bake 400°C, 100h | Fifth Bake 150°C, 30h | Sixth Bake 250°C, 30h | Seventh Bake 150°C, 30h | Eighth Bake 360°C, 150h
TiN | 6.44e-13 | 3.7e-15 | 2.4e-16
SilcoII | 1e-12 | 1.3e-13 | 1.5e-13
SS | 4e-12 | 1.7e-13 | 1e-13 | 2.8e-13 | 2.2e-13 | 9.9e-14
Silco | 2.4e-12 | 2.1e-12 | 1.6e-12 | 3.6e-12 | 2.8e-12 | 1.9e-12
All data at Troom = 20°C
Notes: Steel:
Expected reduction
with 400°C heat
treatment
Silco:
Little variation with
heat treatment
Silco on heated SS:
Retains prior
outgassing rate
TiN coating:
Very low outgassing
rates, consistent with
literature, elemental
Ti pumping?

## Slide 33: Coated chambers with NEG & ion pumps Measured vs. calculated pressures
Chart: Pressure (x10-12 Torr)
| bare | silco | TiN
Extractor | -0.20968 | 4.505 | 1.2903
3BG | 0.93916 | 1.6567 | 1.0143
Calculated | 0.85997 | 0.45897 | 0.00073435
Outgassing rate for TiN coating incorrect:
There must be some pumping from elemental
Ti in the system, making accumulation method
invalid (supported by surface analysis)
Notes: Preliminary conclusions: is the sensitivity correct? Were the outgassing measurements correct? (no for TiN) was the pump speed constant through the three. Work on error bars until the extractor error is more believable. Poor signal to noise on the extractor gauge makes it difficult to use?

## Slide 34
Remember: Better Vacuum, longer lifetime
System Leak
~5x10-11 Torr
~2.5x10-11 Torr

## Slide 35: Summary
High polarization strained superlattice GaAs-GaAsP photocathodes susceptible to ion damage, gun vacuum level
Outgassing reduction, pumping combinations and measurements to verify improvements are being optimized
Vacuum improvements required for future proposed higher current polarized electron sources

## Slide 36: Thank you for your attention
(L-R) Matt Poelker, Joe Grames, Bubba Bullard
Marcy Stutzman, John Hansknecht, Shukui Zhang
Carlos Hernandez Garcia, Phil Adderley, Riad Suleiman
Students: Md. A. Mamun, Edgar Abarca Morales, Wei Liu