## Slide 1: Beam Commissioning Injector Upgrade Phase 1
Reza, Alicia, Yan, and Dennis

## Slide 2: Goal: High current Beam to FC1 for FY2021 Physics Running
Including:
  Initial Check of the newly installed line elements with beam.
    Do all elements behave as expected? Check for miswiring of corrector, lenses, Wiens, etc.
  Vertical and Horizontal Wien setups
  Prebuncher setup
  Acceptable transmission from gun to FC1.
    Low halo, lossless transmission through 15 deg bend and Wiens, above 90% transmission through A1A2 apertures.
    Vacuum quality
  Choppers setup
  Optics measurements
    FOPTs, Emittance measurements
  Bunch length measurements at choppers.
  PQB measurements
  What else?

## Slide 3: Gun to Chopper commissioning
Initial beam threading (Prebuncher, Choppers, and Wiens OFF)
  1.1 Pcup inserted
    1.1.1 Correcting for the gun kick
    1.1.2 Setup the 15 deg bend (This could be tricky. How much stray field is there?)
    1.1.3 Check magnets with beam (correctors, solenoids, quads) for miswiring/shorts.
    1.1.4 Change the Wien magnets by < 10 Gauss.cm to check its wiring. (HV check??)
    1.1.5 Scan aperture from Gun to Pcup to ensure enough clearance for the beam.
    1.1.6 Center on lenses and the Wien pipes.
  1.2 Pcup out, FC1 inserted
    1.2.1 Thread beam through A1A2 and Chopper (off) to FC1
    1.2.2 FOPT systematic corrector checks
    1.2.3 Check the beam shape on the viewers.
    1.2.4 Double check beam centering and clearance from gun to FC1
    1.2.5 Check transmission
    1.2.6 Emittance measurement

## Slide 4: Gun to Chopper commissioning
RF elements setup
  2.1 Setup the choppers
  2.2 Setup the prebuncher
  2.3 Setup all laser phases in the choppers

## Slide 5: Gun to Chopper commissioning
Major Elements Setup
  3.1 Vertical Wien Setup (with Horizontal Wien off)
    3.1.1 Check the beam quality (emittance) and transmission before and after
    3.1.2 Check the beam phase in the choppers before and after
  3.2 Horizontal Wien setup (with Vertical Wien off)
    3.2.1 Check the beam quality (emittance) and transmission before and after
    3.2.2 Check the beam phase in the choppers before and after
  3.3 Both Wiens on
    3.3.1 Check the beam quality (emittance) and transmission before and after
    3.3.2 Check the beam phase in the choppers before and after

## Slide 6: Gun to Chopper commissioning
Beam Measurements
  4.1 Bunch length measurement in the choppers with prebuncher on and off
  4.2 Measure focal length of the solenoids
  4.3 Measure Emittance with Wiens on and off

## Slide 7: Gun to Chopper commissioning
CW to FC1
  5.1 Usual CW checks including transmission, vacuum monitoring, etc.
  5.1 Run CW beam with both Wiens off. (From 10uA to 100uA to FC1)
  5.2 CW beam with Wiens on.

## Slide 8: Preparation for Phase 2 in CEBAF Injector
Gun HV at 200 kV to 1D00 dump
  Solenoids
  Wiens
  Chopper Setup
  Prebuncher
Anything else?