## Slide 1: Spot-Size Reduction
4/5/2017

## Slide 2: UVa Results – Reducing Spot-size Helps
wpc= 1.35mmX,1.46mmY, wqpd=1.43mmX,1.58mmY
  4θ terms ~ 1.3-1.5umX, 0.76-1.1umY
wpc~0.324mmX,0.341mmY,wqpd= 0.886mmX,0.891mmY ~ 7x
  4θ terms (NOT realigned) ~ 2.2-3.1umX, 2.5-4.3umY
  4θ terms(aligned) ~ 0.29-0.27umX,0.24um-0.09umY ~ 5.2x
wpc= 0.324mmX,0.341mmY,wqpd= 2.06mmX,1.907mmY ~ 3x
  4θ terms ~ 0.47-0.54umX,0.33-0.36umY ~ 2.7x
~800nm
Hpol
IHWP(780nm) + Vpol
KD*P
qpd
1m lens
-0.75m lens
RHWP(808nm)
analyzer
4θ terms ~ wpc * wqpd

## Slide 3: 3 scenarios (Injector Table)
Now: wpc=1.9mm, wc=1.1mm
Upstream 1m lens (z=-1.29m): wpc=0.53mm, wc=0.85mm
Up+Downstream(z=0.81) 1m lenses: wpc=0.53, wc=1.1mm
A 1m lens 0.2m after the Pockels cell, would keep the wpc the same=1.9mm and enlarge the wc from 1.1mm to 2.05mm (so 4sigma~4mm).
Scenario #2 has a ~20% spot size reduction on the cathode
Scenario #3 goes through a focus after the vacuum window

## Slide 4: Layout (Conceptual)
Different Halls Beams
combiner
Upstream f=1m lens
Downstream f=1m lens
Steering lens
f=2m
window
cathode
~1.29m
0.81m
25.6cm
8.3cm
2.288m
May have Upstream Lens before or after combiner (either affecting all Halls or affecting only Hall A)
Upstream lens permanently installed
May keep downstream lens installed or not
(20% spot size difference on cathode)

## Slide 5: Predictions (not including vacuum window or 2θ terms)
NOW: model UVa KD*P, Jlab spotsizes
  ~100nm offsets, 10-40nm 4θ terms
  Measure bpm0I01 first after cathode 25-45nm 4θ
1m lens upstream: model UVa KD*P, Jlab predicted spotsizes
  20-30nm offsets, <10nm 4θ terms
+Photocathode Rotation
  (<5nm offsets & 4θ terms prediction)
  Really beneficial for the vacuum window birefringence gradient which is not modeled here…but is important

## Slide 6: Resources Needed
2 f=1m lenses (we have 2 at UVa, but for permanent Jlab installation, so maybe order lenses for the week after next)
2 4-axis lens mounts (pitch,yaw, X,Y)
Control over the helicity board
3-5mW of Hall A laser for alignment (CW or pulsed, either is fine)
HallA Electron beam >20uA (70uA is good) going up to at least FC1
(Conditions of injector beamline should be as if accelerator were going to run 70uA of 1GeV beam, or 150uA of 2GeV beam)
Walk-through of injector laser for Amali (M)
Someone in control room who can change beam current, turn on/off autogaining on bpms (M morn., T even., W morn., W even., Th morn., Fri morn.)
Electron beam (M morn, M aft., T even., W, Th 3x, F)
Access to injector laser room (M even., T, W aft.,Th 3x, F morn.)
Someone who can get the laser to give us 3-5mW of Hall A laser beam(M even., T, W aft.,Th 3x, F morn.)
Someone who can help us rotate the photocathode Angle (Th 3x, F morn)

## Slide 7: Plan
Day1 - benchmarking, bpm/bcm calibration, RHWP scans (#1), QPD laser setup
Day2 - QPD setup, camera measurements, lens alignment, PC realignment, maybe RHWP scans(#2)
Day3 – bpm/bcm calibration, RHWP scans(#2), downstream lens insertion, camera measurements, bpm/bcm calibration, RHWP scans(#3)
Day4 - Photocathode rotation, bpm/bcm calibration, RHWP scan (repeat 3X)
Day5 – Final photocathode angle selection, bpm/bcm calibration, RHWP scan (final)

## Slide 8: Day1
Day1 benchmarking
Morning – bcm/bpm calibration
Need someone in control room who can change beam current, turn on/off autogaining on bpms
HallA Electron beam >20uA (70uA is good) going up to at least FC1
BCM/BPM calibration scan - 5uA steps of current up to max current, auto gaining on injector bpms off
Autogaining of injector bpms back on
Afternoon – RHWP scans
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2 hours
Tweak Pockels cell translation – 1 hour
Repeat 4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2 hours
Evening –laser QPD setup
Need Access to injector laser room
Need someone who can get the laser to give us 3-5mW of Hall A laser beam
Setup pick off to QPD

## Slide 9: Day 2
Day 2 upstream lens insertion and PC re-alignment
Morning – QPD setup, camera measurements, lens alignment
Need Access to injector laser room
Need someone who can get the laser to give us 3-5mW of Hall A laser beam
3-5mW Hall A laser (CW or pulsed, either is fine)
Finish Setup of QPD pickoff / calibration
CHECK PC alignment with no analyzer (steering) and Aq in S2 (do PITA scan to make sure in S2)
Get spiricon measure of spot size at cathode
Repeat measure of spot size at pockels cell (will bring our own spiricon for this)
Insert 1m lens upstream of Pockels Cell at predetermined z-position z=…
measure of spot size at pockels cell (will bring our own spiricon for this)
Measure divergence of laser at Pockels cell
Measure spiricon spot size at cathode
Measure spot size at vacuum window(if possible)
Measure spot size at QPD
Afternoon – PC realignment
Calibrate QPD
Check PC alignment starting point– S1, S2, no anal, RHWP scan
Align Pockels cell
Evening – PC realignment (maybe RHWP scans)
PC alignment

## Slide 10: Day 3
Day 3 PC -RHWP scans + downstream lens insertion
Morning – RHWP scans
Need someone in control room who can change beam current, turn on/off autogaining on bpms
HallA Electron beam >20uA (70uA is good) going up to at least FC1
BCM/BPM calibration scan - 5uA steps of current up to max current, auto gaining on injector bpms off
Autogaining of injector bpms back on
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours
Afternoon – downstream lens insertion + camera measurements
Need Access to injector laser room
Need someone who can get the laser to give us 3-5mW of Hall A laser beam
3-5mW Hall A laser (CW or pulsed, either is fine)
Insert 1m lens downstream of Pockels Cell at predetermined z-position z=…
Measure spiricon spot size at cathode
Measure spot size at vacuum window (if possible)
Evening – bpm/bcm calibration + RHWP scans
Need someone in control room who can change beam current, turn on
HallA Electron beam >20uA (70uA is good) going up to at least FC1
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours
decide to keep downstream lens or remove + RHWP scans
Need Access to injector laser room
Remove 1m lens downstream of PC (if decided)

## Slide 11: Day 4
Day 4 Photocathode rotation
Morning
Need Access to injector room
Need someone who can help us rotate the photocathode Angle #2
HallA Electron beam >20uA (70uA is good) going up to at least FC1
Need someone in control room who can change beam current, turn on/off autogaining on bpms
HallA Electron beam >20uA (70uA is good) going up to at least FC1
BCM/BPM calibration scan - 5uA steps of current up to max current, auto gaining on injector bpms off
Autogaining of injector bpms back on
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours
Afternoon
Need Access to injector room
Need someone who can help us rotate the photocathode Angle #3
HallA Electron beam >20uA (70uA is good) going up to at least FC1
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours
Evening
Need Access to injector room
Need someone who can help us rotate the photocathode Angle #4
HallA Electron beam >20uA (70uA is good) going up to at least FC1
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours

## Slide 12: Day 5
Day 5 Photocathode rotation final
Morning/Afternoon/Evening
Need Access to injector room
Need someone who can help us rotate the photocathode FINAL ANGLE
HallA Electron beam >20uA (70uA is good) going up to at least FC1
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours
PC translation to optimize
4 RHWP scans (IHWP in/out PITA 0/ PITA non-zero) – 2-4 hours

## Slide 13
wpc= 1.35mmX,1.46mmY,wqpd=1.43mmX,1.58mmY
  4θ terms
  1.3um(IHWPout)/1.5um(IHWPin)X, 0.76um(IHWPout)/1.1um(IHWPin)Y
UVA data

## Slide 14
wpcx<=0.324mmX,0.341mmY wqpdx~0.886mmX,~0.891mmY
PC NOT realigned, angle=(1mrad yaw,6mrad pitch)
  4θ terms
  2.2um(IHWPout)/3.1um(IHWPin)X, 2.5um(IHWPout)/4.3um(IHWPin)Y
UVA data

## Slide 15
wpcx<=0.324mmX,0.341mmY wqpdx~0.886mmX,~0.891mmY
PC realigned, angle = (-0.3mrad yaw, 5.5mrad pitch)
  4θ terms
  0.29um(IHWPout)/0.27um(IHWPin)X, 0.24um(IHWPout)/0.09um(IHWPin)Y
UVA data

## Slide 16
  4θ terms
  0.47um(IHWPout)/0.54um(IHWPin)X, 0.33um(IHWPout)/0.36um(IHWPin)Y
wpcx<=0.324mmX,0.341mmY wqpdx~2.06mmX,~1.907mmY
PC realigned, angle = (-0.3mrad yaw, 5.5mrad pitch)
UVA data