## Slide 1: Test of Target Materials Irradiated by 180 MeV Electron Beam at MAMI
Andriy Ushakov1,2, C. Le Galliard1, Eric Voutier1, Joe Grames2, Silviu Covrig Dusa2, Shukui Zhang2 1Université Paris-Saclay, CNRS/IN2P3, IJCLab, Orsay; 2Jefferson Lab
December 1, 2022
This project has received funding from the European Union’s Horizon 2020 research and innovation program under grant agreement No 824093; the U.S. Department of Energy, Office of Science, Office of Nuclear Physics under contract DE-AC05-06OR23177; UT-Battelle, LLC, under contract DE-AC05-00OR22725 with the US Department of Energy (DOE).

## Slide 2: JLab e+ Source: e‒ Beam and e+ Target
Electron beam energy: 120 MeV
Beam current: 1 mA cw
Beam size on target: 1.5 mm (rms)
Target material: pure tungsten
Target thickness: 4 mm
Average power deposited in target: 17 kW
Beam passes through tungsten rim at radius of 17.5 cm
Tangential speed at R = 17.5 cm: 4 m/s
Tungsten cooled by copper disk with 6 water channels
Cross-section of water channels: 10 mm 20 mm
Velocity of water flow: 2 m/s (fully developed turbulent flow)
3D View of 60° Target Sector Cut in the Middle

## Slide 3: JLab e+ Source: Average Temperature and Thermal Stress
Tmax = 258 °C
σvM max = 282 MPa
Average Temperature of Tungsten
Equivalent von Mises Stress

## Slide 4
JLab e+ Source: Thermal Cycling of a Spinning Target
Thermal Cycle
Heating Phase
Temperature and Stress at the End of Heating Phase
ΔTmax ≈ 120°C
Tmax ≈ 340°C
σvM max ≈ 150 MPa
+ 280 MPa due to Tave = 258°C
Duration of Heating ≈
2.5 ms
Period of Cycles =
275 ms

## Slide 5: JLab e+ Source: Radiation Damage of Tungsten
Max Damage per e‒ = 5.7· 10 ‒22 dpa/e‒
Max Damage per 5000 h = 0.215 dpa
Radiation Damage per e‒ in Stationary Target
Damage of Rotated Target vs z

## Slide 6: What We Would Like to Test at MAMI
Rad. Damage of Tungsten ≈ 0.215 dpa
Tungsten at cycling temperatures expected at JLab e+ source:
  Average T ≈ 300 °C
  ΔTpulse ≈ 120 °C
MAMI Beam and Target
Electron beam energy: 180 MeV
Beam current: 22 μA
Beam size on target: 250 μm (RMS)
Target material: pure Tungsten
Target thickness: 100 μm

## Slide 7: Geometry of W Target and Al Holder
Beam
W
Al2O3
3D View of ¼ of Target
Thickness of W target: 100 μm
Length of W target: 80 mm
Width of W target: 3 mm
Thickness of Al2O3 ceramic: 1 mm
Al
Temperature, Radiation and Convection

## Slide 8: Power and Temperature for Duty Cycles of 0.17
Average P = 25.9 W/mm3
Average Temperature
Ppulse = 152 W/mm3
ΔT = 125 °C
Cycling Temperature for Pulse length = 34 ms, Pulse Period = 200 ms
Tave = 302 °C

## Slide 9: Radiation Damage by 22 μA @ 180 MeV e- beam with 250 μm RMS spot size
Rad. Damage of W per e- = 1.96E-20 dpa/e-
22 μA → 1.375E+14 e-/s
For cw e- beam
To reach the damage of 0.215 dpa in W:
• 8E+4 seconds of irradiation is needed (22.2 h).
• = 619 °C in W foil (0.1 mm x 10 mm x 80 mm)
Tungsten
Material | Displacement Threshold Energy [eV] | Rad. Damage [dpa/e-]
Ta | 90 | 1.94E-20
W | 90 | 1.96E-20
Re | 60 | 3.37E-20
Cu | 30 | 
W40Cu |  | 5.27E-20

## Slide 10
# | Thickness [μm] | Length x Width | Pulse Length [ms] | Pulse Period [ms] | Tave [°C] | ΔTpulse [°C] | Rad. Damage [dpa] | Number of cycles | Irradiation Time [h] | Total Charge [C]
1 | 100μm | 80mm x 10mm | cw |  | 619 |  | 0.215 | - | 22.16 | 1.75
2 | 100μm | 80mm x 3mm | 34 | 200 | 302 | 125 | 0.001 | 104 | 0.556 | 0.01
Ʃ |  |  |  |  |  |  |  |  | 22.72 | 1.76
# | Thickness [μm] | Length x Width | Pulse Length [ms] | Pulse Period [ms] | Tave [°C] | ΔTpulse [°C] | Rad. Damage [dpa] | Number of cycles | Irradiation Time [h] | Total Charge [C]
1 | 100μm | 80mm x 10mm | cw |  | 619 |  | 0.215 | - | 22.16 | 1.75
2 | 100μm | 80mm x 3mm | 34 | 200 | 302 | 125 | 0.001 | 104 | 0.556 | 
Ʃ |  |  |  |  |  |  |  |  | 22.72 | 1.76
Target #1: test of radiation damage
Target #2: test 10k thermal cycles with expected temperatures at JLab e+ target
Plan of Irradiation Test of W Targets

## Slide 11: Test at MAMI (180 MeV e-)
To test the radiation damage of target materials at moderate elevated average temperatures (0.215 dpa at 300 °C for W)
(If possible) Short tests of cycling temperature expected at JLab e+ injector
Tasks
Define target geometry (thickness and other dimensions), target holder, and required beam time structure.
Simulate radiation damage and calculate target temperatures of different materials (W, Ta, WRe- and WCu- alloys,).
Prepare/buy materials and equipment for the experiments and measure material properties before irradiation
Post-analysis of irradiated materials:
  Surface analysis (cracks) using microscopes
  Changes in a bulk material (crystal structure), thermal and mechanical properties
  Measure radiation damage using X-rays (PETRA, DESY, Hamburg)
Test at JLab (Laser)
To test high peak heat load (PEDD) in a wide temperature range (up to 1000 °C)
To test fatigues limits at different T

## Slide 12
Backup Slides

## Slide 13: Pulsed MAMI Beam: Temperatures of W foil (0.1mm x 3mm x 80mm) Pulse Length = 34 ms and Pulse Period = 200 ms
Red profile: T at the end of heating (t = 34 ms)
Blue profile: T at the end of thermal cycle (t = 200 ms)

## Slide 14: CW MAMI Beam: Temperature of W foil (0.1mm x 10mm x 80mm)

## Slide 15: Tungsten: Temperature Dependence of Yield Strength and Fatigue Limit vs Number of Cycles
Yield Strength vs Temperature (taken from Plansee website)
Plansee: Yield Strength (T = 420 °C) ≈ 923 MPa
ANSYS Granta: Yield Strength (T = 22 °C) = 1506 MPa
T = 22 °C
Fatigue Limit vs Number of Cycles (taken from ANSYS Granta)
Fatigue Limit (Ncycles = 6.55e+7*) = 774 MPa
* For Ø35 cm target spinning with 4 m/s after 5000h