Mott Commissioning / Spring 2014 Highlights
(January 1, 2014 – June 1, 2014)

January
3	Discriminator thresholds changed
3	New 31MHz timing signal to DAQ
6	CAMAC to VME upgrade completed w/ future punchlist
13	Hall C laser 31MHz w/ 1ns seed pulse
13	New Cu dump end flange received
13	Discriminator thresholds changed
13	Found Mott camera mis-aligned
13	Found 12” long mu-metal on 3D line mis-steered beam
13	499MHz & 31MHz beam to Mott from A/C lasers
13	Commissioning summary plots runs 7237-7243
27	Mott setup script tested
27	Bad Down-dE is found to be a bad cable
28	Mott tested with Hall C laser 62MHz w/ 1ns seed pulse run 7249

February
10	Hall C laser returned 31MHz and KE=5MeV setup
10	Tested CW BLM, BCM, target motion FSD all OK
10	Completed PMT HV calibration w/ Hall A laser 499MHz
10	Instrumental asymmetry vs. position @ 31MHz 1um Au runs 7266-7272
11	Verified no rate w/ beam off, i.e. no 0L02 FE
11	Start of instrumental systematics studies
11	New dE PMT HV setpoints
11	Sleuthing L/R instrumental asymemtries
11	Identified dE HV too low leading to missed dE-E coincidences
11	End with 3 dump runs (0, +5, -5 Amps) @ 499 MHz
12	Added useful run signals to S2 (ungated counters)
12	Solved a dead time issue – increased OR’d Mott trigger from 1ns to 20ns
13	Cu dump cleaned, cooling circuit finished
13	Archiving 96 channels in EPICS + MottRunNumber added
13	Commissioning 31MHz 1um Au logbooks.jlab.org/entry/3269732
- Asymmetry vs. current 0-5uA runs 7312-7318
- Asymmetry vs. position 1uA runs 7319-7328
- Asymmetry vs. dump dipole/empty 1uA runs 7329-7334
- Asymmetry vs. spot size/shape 1uA runs 7335-7337
- Measurement summary logbooks.jlab.org/entry/3269965
16	Identified 900kB data file limit issue
16	Asymmetry vs. kinetic energy (3,4,5,6) 31MHz 1um Au runs 7344-7355
16	Null vertical polarization runs 7356-7358
16	Explore Cu (1.0, 4.1, 8.0 um) 31MHz runs 7359-7361
17	Cu (3 foils) runs 7365-7384 logbooks.jlab.org/entry/3270433
17	Vacuum performance during Cu runs logbooks.jlab.org/entry/3270435
17	Changed dE delays from boxes to cables and then backed out
17	Explore Au (7 foils) runs 7387-7393
18	Au (7 foils) logbooks.jlab.org/entry/3270650
- Foil#16 - 50min (Au:5um) : Run 7394 – 7403
- Foil#15 - 60min (Au:1um) : Run 7404 – 7409
- Foil #4 - 60min (Au: 0.75um) : 7410 – 7415
- Foil #5 - 75min (Au: 0.50um) : 7416 – 7421
- Foil #14 - 120min (Au: 0.35um) : 7422 – 7427
- Foil #12 - 2h45m (Au: 0.05um) : 7428 – 7434
- Foil #13 - 1h0m (Au: 0.05um) : 7435 – 7436
18	Explore Ag (5 foils) runs 7438-7442 logbooks.jlab.org/entry/3270728
18	0L02 trip low liquid level
18	Target ladder brass worm gear loosened, needs replacement
19	Vacuum vs. operation + planning logbooks.jlab.org/entry/3270788
19	Tightened ladder gear, verified home never lost
19	Asymmetry error calc error fixed logbooks.jlab.org/entry/3270814
19	Compare 31/499MHz with/without chopping apertures using Mott TDC
20	Ag (4 foils) runs 7452-7483 logbooks.jlab.org/entry/3270977
- Ag: 10um, Runs 7452-7457
- Ag: 4.5um, Runs 7458-7463
- Ag: 1.6um, Runs 7464-7474
- Ag: 0.45um, Runs 7475-7483
20	Au (foil#13) runs 7484-7485
20	Mott shutdown for SAD
25	Vent Mott: adjust mirror, window port, add GP-100 NEG, replace dump
25	Upgrade details: logbooks.jlab.org/entry/3271331
26	Mott baked 18h, NEG activated, dump LCW logbooks.jlab.org/entry/3271368
27	Vacuum good, improvement logbooks.jlab.org/entry/3271417
27	Dump LCW FSD interlocked OK
27	Camera aligned to target ladder OK
27	Dump operational limits changed to 1kW w/ LCW and 34W w/o LCW
27	Dump temp monitor: webdaq.acc.jlab.org logbooks.jlab.org/entry/3271442
28	Dump monitoring described logbooks.jlab.org/entry/3271535

March
8	Dump tests Au 1um Hall C 31MHz and Hall A 31/499MHz w/ (0, +5, -5 A)
30	Dump temp and vacuum to 30uA CW logbooks.jlab.org/entry/3278662
31	OTR 6uA Cu 1um visible logbooks.jlab.org/entry/3279045

April
4	High current testing logbooks.jlab.org/entry/3280208
- Dump to 75uA
- Au 0.5um to 40uA, no OTR
- Au 0.05um to 40uA, OTR>1uA but foil failed (suspect=formvar)
8	DAQ MVME6100 (iocmdaq1) replaced by SBC (sbcmdaq0)
18	First DAQ upgrade beam test logbooks.jlab.org/entry/3282109
19	Cryounit exit emittance (3-7 MeV/c) logbooks.jlab.org/entry/3282199
23	Able to analyze data files >2GB
24	Mott DAQ upgrade summary logbooks.jlab.org/entry/3282791
26	Second DAQ upgrade beam test logbooks.jlab.org/entry/3283024
28	Cryounit entrance emittance (550keV) logbooks.jlab.org/entry/3283454
28	Mott DAQ Testing: 3283480
29	MOTTLOG created (all previous entries in POLOG)
30	DAQ upgrade done devweb.acc.jlab.org/CSUEApps/atlis/task/14092

May
9	BeCu Dump study @ 31 MHz
- 7650-7662 Different foils with beam current adjusted to created 1.5KHz trigger rate, dump dipoles both at 0A
- 7663-7667 Thru target with various dump dipole settings – ({5,0,-5},{5,0,-5}) Amps
- 3286038
9	Mott Dipole Polarity might be wrong – 3286039
9	7668-7671 Mott charge asymmetry vs DAQ mode study –	3286049
9	7673-7684 Measurement of MFL0I07 B-field – 3286084
22	7685-7695 Tested/characterized cryo-unit at 4K – 3286878