## Slide 1: HDIce Leak scenarios
Jan 17, 2020

## Slide 2: Assumption: leak at the HDIce Target
Concern: pressure increase at the SRF booster
Simulation of beamline as cylinder, all pumping surfaces simple 20 L/s facets in approximate pump locations
For steady state, all ion pumps “on”
For leak state, downstream ion pumps turned off to simulate supply trips

## Slide 3: Steady state simulation
Ion pumps at nominally correct locations, approximated as pumping surface in beamline
  Could simulate pump drops, but pumps will trip off in vent event so not critical
Outgassing simulated at 5e-12 mbarL/scm2
  Approximately what we measured for stainless steel with no heat treatment
No dynamic effects from beam scraping or heating
Highest pressure in waist high quad section ~3e-9 mbar
Best pressure in DP can region ~1e-11 Torr

## Slide 4: Simulate leak 100 mbarL/s at end of pipe
Burst to 100 for 0.3 seconds at 0.1 seconds
No orifice
Pumps downstream of DP can all off for leak event
No simulated valve closing yet
DP can pressure reduction: factor of 10-100
  Line of sight lets some gas through

## Slide 5: Simulate leak 100 mbarL/s at end of pipe, 5 mm orifice
Burst to 100 for 0.3 seconds at 0.1 seconds
5 mm orifice immediately upstream of raster
Pumps downstream of DP can all off for leak event
No simulated valve closing yet
DP can pressure reduction: factor of 10-100
  Line of sight lets some gas through

## Slide 6: 100 mbarL/s burst for 0.3 seconds with and without 5 mm orifice upstream of raster
Chart: no orifice (x) and 5 mm orifice (-)

## Slide 7: Other simulations
Lower “leak” rates for scaling, all no orifice

## Slide 8: Simulation 1: Starting at 0.1 seconds, burst to 1 mbar, duration 0.3 seconds
Pressure vs. time with burst at t=0

## Slide 9: Simulation 2: Starting at 0.1 seconds, burst to 10 mbar, duration 0.3 seconds
Pressure vs. time with burst at t=0