## Slide 1
Mott DAQ Timing
R. Suleiman January 29, 2013

## Slide 2: Helicity Signals Timing
Pair -Sync
Delayed Helicity
T_Settle
Pattern-Sync
T_Settle starts 1.0 µs earlier than all other helicity signals
ΔT(nTSettle-Delayed Hel) = 950 ns
ΔT(nTSettle-PatSync) = 1.1 us
ΔT(nTSettle-PairSync) = 940 ns

## Slide 3: Scalers / FADC_Int
FADC Integration Window
nT_Settle
S1 LNE
nT_Settle
180 ns
TID Trigger
nT_Settle
428 ns
S1 GATE
nT_Settle
70 µs

## Slide 4: Mott Signals
TID Trigger
Mott DetTr
TDC
Mott DetTr
0.322 µs
TDC – COMM START
L1A
216 ns
FADC TRG IN
L1A
33 ns
Jitter 16 ns

## Slide 5
Busy Delays
TOT Busy (us) | Mode | Use | Channels
47.6 - 52.4 | FADC Int | Electron production | At 30 and 960Hz
214 | Mott Sample | Mott production | 
All the measurements were done using the configuration (or control readout list “crl”) files used during the data taking.

## Slide 6
Busy Delays
The evaluation of the busy signal for each single piece was done starting from the configuration files used during the data taking. To evaluate each contribution all the other boards were removed from “trigger” session of the crl file. ONLY the correspondent board was read.
Busy (us) | Module | Mode | 
106 | FADC | Mott Sample | 16 channels and 500 samples
83 | FADC | Mott Sample | 16 channels and 148 samples
25 - 26 | FADC | FADC_Int | 16 channels
14.8 – 17.4 | TDC |  | 
31 - 32 | Scalers (S1+S2) |  | Both boards together
11.8 - 13.2 | Scaler S1 |  | Only Scaler S1
8.8 - 9.9 | All the environment |  | NO readout of any board

## Slide 7
Delays & Signals
Signal | Delay | Width | Jitter
S1 LNE = nTSettle delayed | 180ns | 70us | 
TID Trigger = nTSettle delayed | 428ns | 70us | 
TDC Mott Det .Trigger | 322ns |  | 
L1A - Det. Trigger | 216ns | 33ns | 16ns
L1A FADC Tr input | 1.8us |  | 
TDC PEPPo Detectors | 760ns |  | 
TDC CH19 (TS1) | 812ns |  | 
TDC CH21 (TS1B) | 912ns |  | 
121kHz clock | 8.16us |  | 
TS1 octal unit PEPPo NIM crate | 252ns | 158ns | 
PMT5 (FCD) | 165ns |  | 
220ns ribbon cable + 510 JLab discrl + 10ns modules delay
Signals from Mott LVT2, Coincidence signal
Signal from PEPPo Lin Fan In/Out, cable delay only (156ns)
Signal | Amplitude | Width
TTL detector signals in S2 | +5V | 20ns
NIM detector signals in S1 | -0.7V | 20ns
From PEPPo LVT