## Slide 1: Beam Energy Measurements for Mott Run II : Dry run for Bubble ? Joe Grames
March 9, 2016

## Slide 2
Oct 2015 studied Mott analyzing power vs. beam energy.
Varied beam kinetic energy 4.5-5.3 MeV in 0.2 MeV steps.
Record cavity gradient, Bubble dipole, steering coils, beam positions.
q = 25.0°

## Slide 3
BL = M0 + M1 P + M2 P2 + M3 P3 + M4 P4 + M5 P5
M0 = +4.811
M1 = -1416.2
M2 = +1.2399
M3 = -0.1646
M4 = +0.009795
M5 = -0.00021257
R028 | MDL0L02 | P | dP
MV/m | G-cm | MeV/c | MeV/c
3.35 | 7109.57 | 5.035 | 0.005
3.74 | 7384.34 | 5.229 | 0.005
4.12 | 7646.01 | 5.415 | 0.005
4.5 | 7927.59 | 5.614 | 0.006
4.89 | 8185 | 5.797 | 0.006
J. Benesh, “A detailed examination of the MDL field map and the TOSCA model of this “5 MeV” dipole”, JLab-TN-15-017.
TN provides model for ideal operation with dP/P = 0.1%

## Slide 4
Magnetic fields other than dipole play important role:
  Stray By field (red points) from Earth and Ion Pumps
  Distributed mu-metal helps shield beam from stray field
  Steering coils provide distributed point-correction
Constructed simple model to track fields
  Plots show trajectories for 4.5-6.5 MeV/c in 0.5 MeV/c increments
  Without steering coils beam is “lost” to pipe wall x=1.75cm
  With steering coils orbit is realistic and quasi-independent of momentum
With Stray. Without Mu-Metal, Steering Coils
With Stray, Mu-Metal, Steering Coils

## Slide 5
Record SRF gradient, steering coils, Bubble dipole and beam positions.
Convert recorded beam positions (.XPOS) to absolute survey positions (.XCOR).
  Assumed calibration of beam position monitor to quadrupole s = 0.50 mm
  Assumed survey of quadrupole to absolute coordinates s=0.25 mm

## Slide 6
Model trajectories using beam positions and propagate uncertainties
  Use 0L BPM’s to constrain orbit and predict beam (X,X’) at dipole MDL0L02
  Use (X,X’) at dipole and 5D BPM’s to determine how much q <> 25.0°
  Correct Jay’s model calculation proportionally : PTOSCA(25.0°)[25.0°/(25.0°+q)]
Model of Undeflected 0L beam line
Model of Deflected 5D beam line
<q> = 1.311 ± 0.267 mrad = 0.0751° ± 0.015°
Model predicts dipole deflected beam in excess of 25.0° by <q>:

## Slide 7
Error budget for Mott Run II
Summary for Mott Run II
Recommendations for Bubble
  Shielding helpful, but probably not global solution => still need model
  Improve beam position monitoring around (0L) or further from (5D) dipole
  Greatest “bang for effort” systematic study of model for non-ideal orbits