18Jan 2015

Discriminator threshold adjustments are made on the 6th card of the NIM crate, crate 1, located lowest in the Mott DAQ rack.

The channels are color coded at the bottom, with
Blue = Left
Orange = Right
Green = Up
White = Down

The discriminator pot readout is read between the port on the discriminator board and ground somewhere on the rack.

A reading of -1 V is equivalent to a discriminator threshold of -100 mV, so there is a factor of 10x multiplier applied to the output voltage on the front panel readout.

The scaling between the daq counts that are seen on the graph of hLEFT_E, etc, scale between discriminator mV readings and counts with the following equation:
counts = [-35 * mV] + 988

The graph of the mapping is attached next line item