## Slide 1: Mott DAQ Speed
February 16, 2015

## Slide 2: VME CRATE (VME8200 9U 21Slot - VME64x backplane with no P0 connector)
VME SBC – sbcmdaqo | TID |  | CAEN V538A |  | DB 1 | Mott FADC |  |  | DB 2 | INT FADC |  |  |  | S1 (SIS3801) | S2 (SIS3801) |  |  | TDC CAEN V775 |  |  | FT
Opsmdaq0
L1A
LEMO
4
TRG IN
Common
TRG IN
1 In
17 BFM
18 Mott
DetTr
Gate
1
1
nT_Settle
LN1
L1A
ECL
4
INH
BUSY

## Slide 3
Name | Readout | Trigger
Scalers | Scaler S1 (helicity gated), S2 (un-gated) | Delayed nT_Settle
Mott_Sample | Mott FADC (Mode=1), S1, S2, TDC | Mott Detector
Mott_SemiInt | Mott FADC (Mode=7), S1, S2, TDC | Mott Detector
PEPPo_Int | INT FADC, S1, S2 | nT_Settle
SemiIntFast | Mott FADC (Mode=7), BlockLevel=1 | Mott Detector
SemiIntBlock | Mott FADC (Mode=7), BlockLevel=50 | Mott Detector
Data Taking Configurations

## Slide 4: Use Faster DAQ Mode
For DAQ to be faster:
  No Readout of CAEN V775 TDC or SIS3801 Scalers; only FADC readout
  Use block readout
FADC has a mode with smaller data size and time readout: Pulse Integral and High Resolution Time Mode
What about scalers?
We took data with this mode … results are shown next …

## Slide 5
Comparison of:
Raw ADC Data Samples Mode
vs.
Pulse Integral and High Resolution Time Mode

## Slide 6: Raw ADC data samples – Run 7673

## Slide 7: Raw ADC Data Samples
PL = 100 ns – 8 µs
PTW = 100 ns – 2 µs
Mott Settings:
Programmable Latency (PL) = 60 samples
Programmable Trigger Window (PTW) = 50 samples
Threshold = 0
Each Sample = 4 ns (250 MHz), 0 – 4096 (212)
Mott Readout:
50 samples

## Slide 8: Samples

## Slide 9: Samples

## Slide 10: Analysis: Pedestals

## Slide 11: Analysis: Energy
σE ~ 2.7%

## Slide 12: Analysis: Time-of-flight
σt ~ 0.2 ns

## Slide 13: Pulse Integral and High Resolution Time – Run 7672

## Slide 14: Pulse Integral & High Resolution Time
PL
PTW
Mott Settings:
PL = 60 samples, PTW = 50 samples, NW = NSB + NSA
Programmable Trigger Energy Threshold (TET)
Number of pulses (NP) in PTW = 1 (up to 3 pulses)
Number of Samples Before threshold crossing (NSB) = 5
Number of Samples After threshold crossing (NSA) = 28
TET
NSB
NSA

## Slide 15: TET

## Slide 16: Pulse Integral & High Resolution Time
PL
PTW
Mott Readout:
Pedestal = Average of first 4 samples of window
Peak Value (VP)
Pulse Integral = Sum of raw samples MIN(NSB+NSA, PTW)
Pulse Coarse Time = Sample number N1 such that V(N1) ≤(Vp-Ped)/2)<V(N2)
Pulse Fine Time = 64 ( Vmid – V(N1) )/( V(N2)-V(N1) ). From 0 to 63 in steps of 62.5 ps
TET
NSB
NSA
VP
Pedestal
TC
See
Note 2
See
Note 1
Vmid

## Slide 17: Pulse Pedestal

## Slide 18: Pulse Peak Value (VP)

## Slide 19: Pulse Integral

## Slide 20: Pulse Coarse Time
See
Note 3

## Slide 21: Pulse Fine Time

## Slide 22: Analysis: Pedestals

## Slide 23: Analysis: Energy
σE ~ 3.2%
The increase partially due to average pedestals being subtracted. New data will have this fixed
Energy from run 7642
Pedestal from run 7672

## Slide 24: Analysis: Self Timing Peak
σt ~ 0.33 ns
T->Draw("(TimeCoarse11*4.0 + TimeFine11*0.0625) - (TimeCoarse1*4.0 + TimeFine1*0.0625) >> h1", "TimeCoarse1>1.0 &&TimeCoarse11 >1.0")

## Slide 25: Analysis: E-dE Timing Peak
σt ~ 0.40 ns
T->Draw("(TimeCoarse5*4.0 + TimeFine5*0.0625) - (TimeCoarse1*4.0 + TimeFine1*0.0625) >> h1", "TimeCoarse1>1.0 &&TimeCoarse5 >1.0")

## Slide 26: Analysis: Time-of-flight
Use CH1-CH4
T->Draw("(TimeCoarse1*4.0 + TimeFine1*0.0625) - (TimeCoarse9*4.0 + TimeFine9*0.0625) >> h1", "TimeCoarse1>1.0 &&TimeCoarse9 >1.0")

## Slide 27: Analysis: Time-of-flight
Use CH11
T->Draw("(TimeCoarse11*4.0 + TimeFine11*0.0625) - (TimeCoarse9*4.0 + TimeFine9*0.0625) >> h1", "TimeCoarse1>1.0 &&TimeCoarse9 >1.0 ")

## Slide 28: Analysis: Time-of-flight
Use CH11
T->Draw("(TimeCoarse11*4.0 + TimeFine11*0.0625) - (TimeCoarse9*4.0 + TimeFine9*0.0625) >> h1", "TimeCoarse1>1.0 &&TimeCoarse9 >1.0 &&VPeak9>0.0")

## Slide 29: Notes
Note 2
Note 3
MIN(NSB+NSA, PTW): Make sure NSB+NSA is less than PTW.
Otherwise from CODA readout, cannot tell how many samples
were summed; this is needed for pedestal subtraction.
How FADC deal with a periodic square wave. Possible Solutions:
  Add same signal to another channel delayed by ½ period
  Use signal and signal_bar
PL
PTW
bad
good
Note 1
If first sample above threshold then:
course time =1, fine time = 0, ped = 0, and VPeak = 0

## Slide 30: Conclusions – so far …
Pulse Integral and High Resolution Time Mode has good energy and time resolutions. Must have:
  Correct FADC parameters: Threshold, NSB, NSA
  How to deal with a periodic signal?
    Add a second laser ToF signal delayed by ½ period
    Use laser ToF signal and ToF_bar
Again, what about scalers?
A new CODA Configuration was created with this mode (no CAEN v775 TDC and no SIS3801 Scalers) with block readout. Next …

## Slide 31: Conclusions – so far …
Measure deadtime with 2.2 uA on 1 um gold foil:
  Run 8225: Mott_SemiInt configuration (FADC+TDC+Scalers), Deadtime = 28% at 5.1 kHz.
  Run 8227: SemiIntFast configuration (only readout of FADC, Blocklevel = 1) , Deadtime = 17% at 5.1 kHz.
  Run 8228: SemiIntBlock configuration (only readout of FADC, Blocklevel = 50) , Deadtime = 1% at 5.1 kHz.
Learning how to unblock and analyze