## Slide 1: MDL0L02 Dipole Field
July 6, 2016

## Slide 2: Field Map
Meas. Date: 8/29/2014
Coil used: Hall Probe Stepper
Current (A) Strength (Gauss-cm)
-------------- -------------------
-9.992 -23944.2
-8.996 -21569.6
-7.991 -19169.0
-6.990 -16769.5
-5.990 -14360.7
-4.993 -11954.6
-3.994 -9542.8
-2.989 -7116.1
-1.989 -4698.2
-0.990 -2283.6
0.003 126.0
1.009 2548.4
2.009 4960.8
3.009 7374.6
4.010 9785.8
5.010 12192.0
6.010 14589.8
7.011 16980.4
8.013 19360.4
9.015 21720.5
10.014 24038.1

## Slide 3: Field Map Effective Length
Meas. Date: 8/29/2014
Coil used: Hall Probe Stepper
Current (A) Eff. L (cm)
-------------- -------------------
10.00 12.88
9.00 12.89
8.00 12.89
7.00 12.90
6.00 12.90
5.00 12.91
4.00 12.92
3.00 12.93
2.00 12.95
1.00 13.01
0.00 16.39
-1.00 12.73
-2.00 12.81
-3.00 12.83
-4.00 12.85
-5.00 12.85
-6.00 12.86
-7.00 12.86
-8.00 12.86
-9.00 12.86
-10.00 12.87

## Slide 4: 0 BdL

## Slide 5
Distance from dipole center (cm) | <By> (G)
(Upstream) -15 | 0.4
-14 | 0.4
-13 | 0.3
-12 | 0.3
-11 | 0.2
-10 | 0.1
-9 | 0.0
-8 | -0.6
-7 | -2.2
-6 | -3.7
-5 | -3.9
-4 | -4.1
-3 | -3.9
-2 | -4.2
-1 | -3.9
0 | -3.8
1 | -3.9
2 | -3.9
3 | -4.0
4 | -3.8
5 | -3.9
6 | -2.1
7 | -0.6
8 | 0.0
9 | 0.1
10 | 0.3
11 | 0.6
12 | 0.6
13 | 0.6
14 | 0.6
(Downstream) 15 | 0.6
Dipole Survey (March 20, 2016)
On Hysteresis
I = -0.0489 A
BdL = 0.000 G-cm
Hall Probe = -3.98 G*
Zero is center of dipole and survey was done along 0L Region using a hand-held Hall probe
*Accuracy: ±(0.03% of reading +0.006% of full scale) max
Hall Probe = -3.98 G ± 0.18 G

## Slide 6: 0 Current

## Slide 7
Degaussed

## Slide 8: Degaussed

## Slide 9: Test at Magnet (March 15, 2016)
With cables from trim card swapped at magnet (after degaussing), Hall Probe = -0.26 G
With current cables disconnected, Hall Probe = -0.06 G
With cables back to normal at magnet (after degaussing), Hall Probe = +0.50 G
With current set to zero and magnet degaussed, the Hall Probe measures +0.5 G. At this setting, we measured about +1.3 mA of current to magnet using KEITHLEY DMM7510 in series with MDL0L02.

## Slide 10
Trim Card Set Current | DMM Measured Current
0.0 | +0.0011
1.0 | 1.0008
2.0 | 2.0002
3.0 | 3.0010
4.0 | 4.0009
5.0 | 5.0014
6.0 | 6.0014
7.0 | 7.0018
8.0 | 8.0034
9.0 | 9.0042
9.0 | 9.0044
8.0 | 8.0044
7.0 | 7.0034
6.0 | 6.0030
5.0 | 5.0030
4.0 | 4.0025
3.0 | 3.0024
2.0 | 2.0013
1.0 | 1.0009
0.0 | +0.0016
-1.0 | -0.9984
-2.0 | -1.9985
-3.0 | -2.9992
-4.0 | -3.9988
-5.0 | -4.9986
-6.0 | -5.9982
-7.0 | -6.9980
-8.0 | -7.9987
-9.0 | -8.9986
0.0 | +0.0013
Use KEITHLEY DMM7510 in series with MDL0L02 to measure current powering magnet
For Bubble Chamber, we are especially interested in momenta around 5.5 MeV/c (fluorine measurement) and 8.5 MeV/c (oxygen measurement)
Corresponding magnet currents are 3.2 A and 5.0 A. At these currents, Trim card is good to about 1 to 2 mA.

## Slide 11: Beam Momentum in 5D (opera – Jay)
BdL = -4.811 + 1416.2*p - 1.2399*p2 + 0.1646*p3 - 0.009795*p4 + 0.00021257*p5
Jay's Tech Note TN-15-017, page 9.

## Slide 12: Scanned Dipole current from 9.0 to -9.0 A (with ON hysteresis) and recorded Bdl and Hall Probe readings
March 20, 2016

## Slide 13
Current (A) | Core Field (G)
10.01 | 1863.40
9.01 | 1682.70
8.01 | 1499.15
7.01 | 1314.35
6.01 | 1128.70
5.01 | 942.75
4.01 | 756.15
3.01 | 569.25
2.01 | 382.25
1.01 | 195.15
0.00 | 7.40
-0.99 | -179.45
-1.99 | -366.65
-2.99 | -554.15
-3.99 | -742.50
-4.99 | -929.65
-5.99 | -1116.25
-6.99 | -1302.90
-7.99 | -1489.25
-9.00 | -1675.60
-9.99 | -1860.05
From Field Map
Current (A) Hall Probe (G)
9.0 1676.84
8.0 1494.811
7.0 1311.098
6.0 1126.72
5.0 941.599
4.0 755.813
3.0 569.625
2.0 382.975
1.0 196.247
0.0 9.73
-1.0 -176.974
-2.0 -363.63
-3.0 -550.311
-4.0 -736.641
-5.0 -922.818
-6.0 -1108.806
-7.0 -1294.494
-8.0 -1480.191
-9.0 -1665.526
From Current
Scan

## Slide 15: Field Ratio
BdL / Hall Probe

## Slide 16
(BdL+8.0) / Hall Probe
(i.e., added small offset to field map)

## Slide 17: Scanned Dipole current from 9.0 to -9.0 A (with ON hysteresis) and recorded Bdl and Hall Probe readings
May 11, 2016

## Slide 18
Current (A) | Core Field (G)
10.01 | 1863.40
9.01 | 1682.70
8.01 | 1499.15
7.01 | 1314.35
6.01 | 1128.70
5.01 | 942.75
4.01 | 756.15
3.01 | 569.25
2.01 | 382.25
1.01 | 195.15
0.00 | 7.40
-0.99 | -179.45
-1.99 | -366.65
-2.99 | -554.15
-3.99 | -742.50
-4.99 | -929.65
-5.99 | -1116.25
-6.99 | -1302.90
-7.99 | -1489.25
-9.00 | -1675.60
-9.99 | -1860.05
From Field Map
Current (A) Hall Probe (G)
9.0 1679.413
8.0 1497.175
7.0 1313.228
6.0 1128.575
5.0 943.145
4.0 757.186
3.0 570.716
2.0 383.856
1.0 196.884
0.0 10.084
-1.0 -176.862
-2.0 -363.784
-3.0 -550.767
-4.0 -737.396
-5.0 -923.895
-6.0 -1110.12
-7.0 -1296.09
-8.0 -1481.945
-9.0 -1667.458
From Current
Scan

## Slide 20: Field Ratio
BdL / Hall Probe

## Slide 21
(BdL+12.0) / Hall Probe
(i.e., added small offset to field map)

## Slide 22: Summary - I
Hints that there are problems with Field Map:
  Measured field in Injector is not zero when BdL is set to zero
  Effective Length from Field Map depends on magnet current
Recommend to use spare magnet to resolve this problem
For magnet currents below 6 A, Trim Card is good to 2 mA
Is MPT-231 Hall Probe mounted right inside MDL0L02? Should use G10 to hold probe parallel to magnet
Hall Probe MPT-231 measurements at very small fields vary by about 0.2 – 0.3 G, due to exact hysteresis history
Jay’s model is good to 0.1% (see Tech Note TN-15-017)
Either due to errors in current measurement and/or environmental fields

## Slide 23: Summary - II
For Beam Energy Measurement:
  CEBAF : BdL 0 (due to field map error), instead:
  BdL = Hall Probe * Magnetic Length (or use my field survey)
  BdL -3.6 G * 12.9 cm = -46 G-cm. Treat as another horizontal corrector.
  Spectrometer Lines (2D, 3D, 5D): magnet currents between 2 A and 6 A are most relevant. Till we check field map, we will assume a -46 G-cm error on field measurements.

## Slide 24: Summary - III
 | Error
Trim Power Supply | 2 mA
Magnet Model | 0.1%
Field Map | -46 G-cm
For Mott Energy Measurements