## Slide 1: Parity Quality Beam (PQB) Working Group Report
OPS StayTreat
Riad Suleiman
Friday July 17, 2015

## Slide 2: Upcoming Parity Violation Experiments
Experiment | Energy (GeV) | Pol (%) | I (µA) | Target | Apv (ppb) | Maximum Charge Asym (ppb) | Maximum Position Diff (nm) | Maximum Angle Diff (nrad) | Maximum Size Diff (δσ/σ)
PREx-II | 1.0 | 90 | 70 | 208Pb (0.5mm) | 500±15 | 100±10 | 1±1 | 0.3±0.1 | 10-4
C-REx | 2.2 | 90 | 150 | 48Ca (5mm) | 2000±42 | 100±10 | 1±1 | 0.3±0.1 | 10-4
MOLLER | 11.0 | 90 | 60 | 1H (150 cm) | 35.6±0.74 | 10±10 | 0.5±0.5 | 0.05±0.05 | 10-4
  PREx-II is tentatively scheduled for Hall A in Fall 2016
  C-REx is tentatively scheduled for Hall A in Spring 2017
  MOLLER is planned for Hall A in 2022

## Slide 3: Issues from QWeak
  Beam Halo: there was beam halo that could develop a large helicity-correlated charge asymmetry
    Beam halo charge asymmetry was found to depend on:
    Beam loss at Injector Apertures
    Laser phase
    Machine tuning
    What can we do to help?
  BCM Resolution is not suitable for MOLLER:
    QWeak achieved 65 ppm but MOLLER requires 10 ppm
    BCM digital receiver bench studies found that local oscillator phase and amplitude noise being responsible for 65 ppm noise floor
    Try new digital receivers
    Improve phase/amplitude noise of local oscillator

## Slide 4: Hall A 2015/2016 Beamline Plan for PQB
  Install QWeak halo monitor in Hall A beamline
  Equip two BCMs with new-style digital receivers
  Install QQQ cavity triplet in Hall A beamline
  Reinstate Hall A beam modulation system (air-core coils and associated control/drive electronics)
    Beam modulation and accelerator Fast Feedback: is there a need to pause FFB when modulation is ON? does FFB implement a digital notch filter to attenuate modulation frequency (QWeak: 125 Hz for 20 s every 320 s) while passing all other frequencies?
    Need software support
Beam Studies of these systems

## Slide 5: Synchrotron Radiation @ 11 GeV
  What are the issues? Any helicity-correlated effects?
    Synchrotron radiation causes energy spread and energy tails – clipping might be helicity-correlated (e.g., due to helicity-correlated position differences etc.)
    Emittance growth
    Effect on adiabatic damping
    Power is helicity dependent
    Is there any helicity-correlated emittance growth
    Depolarization due to spin-flip synchrotron radiation (Sokolov–Ternov effect) is very small

## Slide 6: Tasks Summary

## Slide 7: Laser Table Tasks
Task | Sub Tasks | Date | Description
2 kHz Helicity Reversal |  | MOLLER | Requires 10 µs settle time – No ringing (not required for PREX, but hoped to test at this time). No Kerr Cell.
 | RTP Pockels Cell |  | Buy test crystals to characterize, design RTP quarter-wave system.
 | KD*P re-design |  | Model E-field to maximize PC uniformity, buy a properly engineered, one with the correct cell-diameter-to-laser-beam-diameter aspect ratio
Pockels Cell Stewart Platform |  | 2016 | For remote optimization using e-beam. Assemble, build control software, qualify

## Slide 8: Injector Tasks
Task | Sub Tasks | Date | Description
Reinstate Injector Parity DAQ |  |  | 
Improve 2-Wien Flip Optics |  | PREx-II | 
Injector Matching |  | PREx-II | Maximize damping
Helicity-correlated Beam Size Monitor |  | PREx-II | Looking for ideas!
Upgrade Helicity Magnet controls |  | PREx-II | 
Locate Helicity Magnets to span (x,x') and (y,y') to minimize both position and angle |  | Fall 2015 | 
Augment helicity steering dipoles with helicity size quads |  | PREx-II | 
Share Injector apertures' current read-back with parity DAQ |  | Fall 2015 | 
MOLLER Feedback to minimize transverse polarization |  |  | Once a shift, adjust Wien angle

## Slide 9: Accelerator Tasks
Task | Sub Tasks | Date | Description
Study Depolarization at Higher Passes |  |  | 
 | Energy stability and precession to Hall |  | 
Synchrotron Radiation |  | MOLLER | 
 | Helicity-correlated emittance growth |  | 
 | Energy spread and energy tails |  | Clipping might be helicity dependent
 | Polarization dependence |  | 
 | Adiabatic damping |  | 
MOLLER (g-2) Spin Flip |  |  | Change beam energy by 100 MeV (few reversals)

## Slide 10: Hall A Tasks
Task | Sub Tasks | Date | Description
Reinstate Hall A Parity DAQ |  |  | 
Beam Halo |  |  | 
 | Install QWeak halo monitors in Hall A | Fall 2015 | 
BCM Resolution |  |  | 
 | New cavities and receivers | Fall 2015 | 
 | BCM receiver bench studies and beam studies | Fall 2015 | 
Beam Modulation |  | Fall 2015 | 
Beam Polarimetry |  |  | 
Spin Dance |  |  | 
Beam Matching and Optics |  |  | 
Phase Trombone |  |  |