## Summary
Magnet: | Solenoid #2 38mm | MFX002 |  |  | Inductance: | Freq (Hz) | 20 | 100 | DC Res.( Ω) | 3.78
Staff: | Mike Beck |  |  |  |  | Ls (mH) | 80.84 | 34.78 | Temp. (C) | 46.6 | @ 4A
Date: | Aug 17,2020 | **Hysteresis +/-4A**
Measurment #1: Single-Wound Centerline Integrals
Current (A) | Field Strength (g-cm) | Core Field (g) | Efective Length (cm)
3.9995 | -2703.9951 | -467.71 | 5.7813
3.0008 | -2031.2099 | -353.928 | 5.739
2.7994 | -1895.3366 | -330.717 | 5.731
2.0005 | -1355.755 | -238.096 | 5.6942
0.99884 | -679.2538 | -121.137 | 5.6073
-0.0030953 | -1.4813 | -3.508 | 0.42226
-0.99879 | 672.5707 | 113.723 | 5.9141
-1.9995 | 1349.5008 | 231.644 | 5.8258
-2.7992 | 1891.1695 | 325.902 | 5.8029
-3.0003 | 2026.7307 | 349.534 | 5.7984
-3.9987 | 2702.3001 | 466.778 | 5.7893
Measurement #2: Bz v I (Center)
Current (A) | Core Field (g) | Voltage (V)
-3.9986 | 466.84 | -15.576
-3.0002 | 353.18 | -11.655
-2.7991 | 329.99 | -10.871
-1.9995 | 237.27 | -7.756
-0.99876 | 120.31 | -3.869
-0.00106 | 3.14 | 0
0.99896 | -114.61 | 3.91
2.0005 | -232.62 | 7.83
2.7995 | -326.73 | 10.958
3.0007 | -350.39 | 11.746
3.9995 | -467.79 | 15.61
3.0008 | -354.05 | 11.746
2.7995 | -330.82 | 10.958
2.0006 | -238.17 | 7.83
0.99885 | -121.12 | 3.91
-0.00121 | -3.68 | 0
-0.99882 | 113.82 | -3.869
-1.9994 | 231.75 | -7.756
-2.7992 | 325.9 | -10.871
-3.0002 | 349.54 | -11.655
-3.9985 | 466.81 | -15.576
Measurement #3: Counter-Wound Centerline Integrals
Current (A) | Field Strength (g-cm) | Core Field (g) | Upstream Coil (g) | Downstream Coil (g)
3.0008 | 6.654 | -9.03 | -228.112 | 228.347
Field Map Data:
Magnet Name: | FX002 |  | Magnet Name: | FX002
Magnet Location: |  |  | Magnet Location:
Measurement Type: | Stepper Probe |  | Measurement Type: | Stepper Probe taken at coil#1 peak
Coil Configuation: | Single-Wound |  | Coil Configuation: | Counter-Wound
Current | Strength |  | Current | Strength
(Amps) | (Gauss) |  | (Amps) | (Gauss)
------ | ------- |  | ------ | -------
3.9987 | -467.63 |  | 3.9988 | -299.11
3.7971 | -445.18 |  | 3.797 | -285.26
3.5984 | -422.63 |  | 3.5981 | -271.27
3.3989 | -399.89 |  | 3.3989 | -257.1
3.1998 | -377.02 |  | 3.1997 | -242.77
2.9999 | -354.08 |  | 2.9999 | -228.29
2.7986 | -330.85 |  | 2.7985 | -213.61
2.5973 | -307.57 |  | 2.5971 | -198.82
2.398 | -284.5 |  | 2.3978 | -184.15
2.1988 | -261.42 |  | 2.1988 | -169.35
1.9995 | -238.19 |  | 1.9995 | -154.55
1.7999 | -214.94 |  | 1.7999 | -139.68
1.5984 | -191.38 |  | 1.598 | -124.59
1.3963 | -167.82 |  | 1.3962 | -109.49
1.1961 | -144.38 |  | 1.1961 | -94.44
0.99788 | -121.19 |  | 0.99782 | -79.55
0.79932 | -97.85 |  | 0.79904 | -64.57
0.60112 | -74.62 |  | 0.60108 | -49.64
0.39991 | -50.99 |  | 0.39976 | -34.46
0.19847 | -27.34 |  | 0.19845 | -19.26
-0.00193 | -3.74 |  | -0.00214 | -4.09
-0.20104 | 19.69 |  | -0.20124 | 11
-0.4 | 43.11 |  | -0.4 | 26.08
-0.59951 | 66.64 |  | -0.59978 | 41.21
-0.79985 | 90.22 |  | -0.79978 | 56.37
-0.99978 | 113.78 |  | -0.99983 | 71.53
-1.1997 | 137.35 |  | -1.1998 | 86.69
-1.4001 | 160.96 |  | -1.4003 | 101.91
-1.5998 | 184.51 |  | -1.5998 | 117.08
-1.8003 | 208.15 |  | -1.8004 | 132.28
-2.0003 | 231.74 |  | -2.0004 | 147.45
-2.1999 | 255.2 |  | -2.1999 | 162.55
-2.3997 | 278.74 |  | -2.3999 | 177.68
-2.5996 | 302.27 |  | -2.5996 | 192.84
-2.8002 | 325.87 |  | -2.8003 | 208.02
-3.0013 | 349.56 |  | -3.0012 | 223.23
-3.2017 | 373.14 |  | -3.2017 | 238.38
-3.401 | 396.53 |  | -3.401 | 253.43
-3.6007 | 419.99 |  | -3.6006 | 268.48
-3.8003 | 443.43 |  | -3.8003 | 283.56
-3.9998 | 466.84 |  | -3.9997 | 298.58
MFX 38mm (2nd article) on 3m Test stand
Single Wound
Counter Wound
Field Directions:

## CW Run 3A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002018.stp | ! data file
2020-08-12 | ! date
10:04 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, CW/Bucked, Hyst x 4 (+/-4A) I=3A
! end File Header |  |  |  |  | I | 3.0008 | A
! begin Step Header |  |  |  |  | BdL | 6.654 | g-cm
0 | ! set current |  |  |  | B_center | -9.03 | g
0 | ! actual current at start of step |  |  |  | B_peak1 | -228.112 | g
0 | ! actual current at end of step |  |  |  | B_Peak2 | 228.347 | g
! end Step Header |  |  |  |  | Eff_L | -0.73688 | cm
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.745 | 16.952 | -0.346 | 0.30011 | 0.246 | -24.998 | -0.346
13.745 | 17.444 | -0.36 | 0.30007 | 0.251 | -24.506 | -0.36
13.745 | 17.946 | -0.38 | 0.30007 | 0.249 | -24.004 | -0.38
13.745 | 18.444 | -0.41 | 0.30006 | 0.2505 | -23.506 | -0.41
13.745 | 18.945 | -0.441 | 0.30007 | 0.249 | -23.005 | -0.441
13.745 | 19.443 | -0.476 | 0.30007 | 0.249 | -22.507 | -0.476
13.745 | 19.941 | -0.516 | 0.30007 | 0.252 | -22.009 | -0.516
13.745 | 20.445 | -0.56 | 0.30008 | 0.249 | -21.505 | -0.56
13.745 | 20.943 | -0.613 | 0.30006 | 0.25 | -21.007 | -0.613
13.745 | 21.443 | -0.669 | 0.30008 | 0.2505 | -20.507 | -0.669
13.745 | 21.944 | -0.731 | 0.30007 | 0.2485 | -20.006 | -0.731
13.745 | 22.441 | -0.806 | 0.30007 | 0.252 | -19.509 | -0.806
13.745 | 22.945 | -0.895 | 0.30007 | 0.2495 | -19.005 | -0.895
13.745 | 23.444 | -0.993 | 0.30008 | 0.2485 | -18.506 | -0.993
13.745 | 23.941 | -1.11 | 0.30007 | 0.251 | -18.009 | -1.11
13.745 | 24.443 | -1.239 | 0.30007 | 0.249 | -17.507 | -1.239
13.745 | 24.941 | -1.398 | 0.30007 | 0.2505 | -17.009 | -1.398
13.745 | 25.442 | -1.58 | 0.30009 | 0.2515 | -16.508 | -1.58
13.745 | 25.945 | -1.795 | 0.30006 | 0.2485 | -16.005 | -1.795
13.745 | 26.442 | -2.05 | 0.30008 | 0.2505 | -15.508 | -2.05
13.745 | 26.943 | -2.348 | 0.30007 | 0.2495 | -15.007 | -2.348
13.745 | 27.442 | -2.713 | 0.30008 | 0.2495 | -14.508 | -2.713
13.745 | 27.941 | -3.152 | 0.30009 | 0.2515 | -14.009 | -3.152
13.745 | 28.444 | -3.687 | 0.30008 | 0.249 | -13.506 | -3.687
13.745 | 28.942 | -4.326 | 0.30008 | 0.2495 | -13.008 | -4.326
13.745 | 29.441 | -5.111 | 0.30007 | 0.2505 | -12.509 | -5.111
13.745 | 29.942 | -6.581 | 0.30008 | 0.2495 | -12.008 | -6.581
13.745 | 30.441 | -8 | 0.30008 | 0.252 | -11.509 | -8
13.745 | 30.945 | -9.63 | 0.30007 | 0.249 | -11.005 | -9.63
13.745 | 31.443 | -12.168 | 0.30007 | 0.2485 | -10.507 | -12.168
13.745 | 31.94 | -14.986 | 0.30008 | 0.1005 | -10.01 | -14.986
13.745 | 32.141 | -16.467 | 0.30009 | 0.1005 | -9.809 | -16.467
13.745 | 32.342 | -18.2 | 0.30008 | 0.1 | -9.608 | -18.2
13.745 | 32.542 | -19.784 | 0.3001 | 0.0985 | -9.408 | -19.784
13.745 | 32.739 | -22.16 | 0.30007 | 0.101 | -9.211 | -22.16
13.745 | 32.941 | -24.492 | 0.30008 | 0.1015 | -9.009 | -24.492
13.745 | 33.144 | -27.074 | 0.30007 | 0.1005 | -8.806 | -27.074
13.745 | 33.345 | -29.734 | 0.30008 | 0.0985 | -8.605 | -29.734
13.745 | 33.542 | -32.548 | 0.30007 | 0.1 | -8.408 | -32.548
13.745 | 33.742 | -36.504 | 0.30008 | 0.101 | -8.208 | -36.504
13.745 | 33.944 | -40.441 | 0.30007 | 0.1 | -8.006 | -40.441
13.745 | 34.144 | -44.682 | 0.30007 | 0.0985 | -7.806 | -44.682
13.745 | 34.341 | -49.896 | 0.30008 | 0.1 | -7.609 | -49.896
13.745 | 34.541 | -55.472 | 0.30008 | 0.1005 | -7.409 | -55.472
13.745 | 34.742 | -61.22 | 0.30007 | 0.1 | -7.208 | -61.22
13.745 | 34.942 | -67.805 | 0.30007 | 0.0985 | -7.008 | -67.805
13.745 | 35.139 | -75.13 | 0.30007 | 0.1005 | -6.811 | -75.13
13.745 | 35.34 | -83.55 | 0.30009 | 0.1015 | -6.61 | -83.55
13.745 | 35.543 | -92.01 | 0.30009 | 0.1005 | -6.407 | -92.01
13.745 | 35.744 | -101.687 | 0.30007 | 0.099 | -6.206 | -101.687
13.745 | 35.942 | -111.009 | 0.30008 | 0.1 | -6.008 | -111.009
13.745 | 36.142 | -121.4 | 0.30008 | 0.101 | -5.808 | -121.4
13.745 | 36.344 | -131.996 | 0.30008 | 0.1 | -5.606 | -131.996
13.745 | 36.544 | -143.148 | 0.30007 | 0.0985 | -5.406 | -143.148
13.745 | 36.741 | -154.197 | 0.30007 | 0.1 | -5.209 | -154.197
13.745 | 36.941 | -165.013 | 0.30006 | 0.1005 | -5.009 | -165.013
13.745 | 37.142 | -175.592 | 0.30007 | 0.0995 | -4.808 | -175.592
13.745 | 37.341 | -186.072 | 0.30008 | 0.098 | -4.609 | -186.072
13.745 | 37.537 | -195.148 | 0.30007 | 0.101 | -4.413 | -195.148
13.745 | 37.739 | -203.71 | 0.30009 | 0.101 | -4.211 | -203.71
13.745 | 37.941 | -211.543 | 0.30007 | 0.1005 | -4.009 | -211.543
13.745 | 38.142 | -218.036 | 0.30008 | 0.099 | -3.808 | -218.036
13.745 | 38.34 | -222.929 | 0.30007 | 0.1005 | -3.61 | -222.929
13.745 | 38.541 | -226.035 | 0.30008 | 0.1015 | -3.409 | -226.035
13.745 | 38.744 | -227.752 | 0.30008 | 0.1 | -3.206 | -227.752
13.745 | 38.944 | -228.112 | 0.30007 | 0.0985 | -3.006 | -228.112
13.745 | 39.141 | -226.777 | 0.30008 | 0.1 | -2.809 | -226.777
13.745 | 39.341 | -223.499 | 0.30008 | 0.1005 | -2.609 | -223.499
13.745 | 39.542 | -217.642 | 0.30007 | 0.0995 | -2.408 | -217.642
13.745 | 39.741 | -210.541 | 0.30008 | 0.0985 | -2.209 | -210.541
13.745 | 39.938 | -201.049 | 0.30009 | 0.1005 | -2.012 | -201.049
13.745 | 40.139 | -189.214 | 0.3001 | 0.101 | -1.811 | -189.214
13.745 | 40.341 | -176.046 | 0.30009 | 0.1005 | -1.609 | -176.046
13.745 | 40.542 | -160.194 | 0.30008 | 0.0995 | -1.408 | -160.194
13.745 | 40.741 | -143.077 | 0.30008 | 0.1 | -1.209 | -143.077
13.745 | 40.941 | -124.089 | 0.30009 | 0.1015 | -1.009 | -124.089
13.745 | 41.144 | -102.747 | 0.30008 | 0.1 | -0.806 | -102.747
13.745 | 41.344 | -80.478 | 0.3001 | 0.099 | -0.606 | -80.478
13.745 | 41.542 | -57.998 | 0.30009 | 0.0995 | -0.408 | -57.998
13.745 | 41.741 | -33.706 | 0.30007 | 0.1005 | -0.209 | -33.706
13.745 | 41.942 | -9.03 | 0.30008 | 0.1 | -0.008 | -9.03
13.745 | 42.142 | 15.603 | 0.30007 | 0.098 | 0.192 | 15.603
13.745 | 42.338 | 39.055 | 0.30007 | 0.1005 | 0.388 | 39.055
13.745 | 42.539 | 63.362 | 0.30009 | 0.101 | 0.589 | 63.362
13.745 | 42.741 | 86.673 | 0.30008 | 0.1005 | 0.791 | 86.673
13.745 | 42.942 | 108.527 | 0.30009 | 0.099 | 0.992 | 108.527
13.745 | 43.14 | 128.531 | 0.30008 | 0.1005 | 1.19 | 128.531
13.745 | 43.341 | 147.344 | 0.30008 | 0.101 | 1.391 | 147.344
13.745 | 43.543 | 164.317 | 0.30007 | 0.101 | 1.593 | 164.317
13.745 | 43.745 | 179.084 | 0.30007 | 0.0985 | 1.795 | 179.084
13.745 | 43.942 | 192.139 | 0.30009 | 0.1 | 1.992 | 192.139
13.745 | 44.142 | 202.968 | 0.30008 | 0.1005 | 2.192 | 202.968
13.745 | 44.343 | 212.089 | 0.30007 | 0.1 | 2.393 | 212.089
13.745 | 44.543 | 219.305 | 0.30008 | 0.0985 | 2.593 | 219.305
13.745 | 44.74 | 223.763 | 0.30008 | 0.1 | 2.79 | 223.763
13.745 | 44.94 | 226.96 | 0.30008 | 0.1005 | 2.99 | 226.96
13.745 | 45.141 | 228.347 | 0.30008 | 0.1005 | 3.191 | 228.347
13.745 | 45.342 | 228.212 | 0.30008 | 0.0995 | 3.392 | 228.212
13.745 | 45.541 | 225.825 | 0.30008 | 0.1 | 3.591 | 225.825
13.745 | 45.741 | 222.03 | 0.30008 | 0.1015 | 3.791 | 222.03
13.745 | 45.944 | 216.623 | 0.30011 | 0.1 | 3.994 | 216.623
13.745 | 46.144 | 210.176 | 0.3001 | 0.0995 | 4.194 | 210.176
13.745 | 46.343 | 202.469 | 0.30011 | 0.0995 | 4.393 | 202.469
13.745 | 46.542 | 193.543 | 0.30009 | 0.1005 | 4.592 | 193.543
13.745 | 46.743 | 183.782 | 0.30008 | 0.1 | 4.793 | 183.782
13.745 | 46.943 | 173.329 | 0.30007 | 0.0985 | 4.993 | 173.329
13.745 | 47.14 | 162.562 | 0.30008 | 0.1 | 5.19 | 162.562
13.745 | 47.34 | 151.658 | 0.30008 | 0.1005 | 5.39 | 151.658
13.745 | 47.541 | 140.652 | 0.30007 | 0.1005 | 5.591 | 140.652
13.745 | 47.742 | 129.676 | 0.30008 | 0.099 | 5.792 | 129.676
13.745 | 47.94 | 119.288 | 0.30008 | 0.1005 | 5.99 | 119.288
13.745 | 48.141 | 108.927 | 0.30008 | 0.101 | 6.191 | 108.927
13.745 | 48.343 | 98.772 | 0.30008 | 0.1005 | 6.393 | 98.772
13.745 | 48.544 | 90.036 | 0.30008 | 0.099 | 6.594 | 90.036
13.745 | 48.742 | 81.67 | 0.30008 | 0.1 | 6.792 | 81.67
13.745 | 48.942 | 73.916 | 0.30008 | 0.101 | 6.992 | 73.916
13.745 | 49.144 | 66.643 | 0.30007 | 0.1 | 7.194 | 66.643
13.745 | 49.344 | 59.682 | 0.30007 | 0.0985 | 7.394 | 59.682
13.745 | 49.541 | 53.891 | 0.30008 | 0.1 | 7.591 | 53.891
13.745 | 49.741 | 48.539 | 0.3001 | 0.1005 | 7.791 | 48.539
13.745 | 49.942 | 44.189 | 0.30008 | 0.1 | 7.992 | 44.189
13.745 | 50.142 | 39.912 | 0.30008 | 0.099 | 8.192 | 39.912
13.745 | 50.34 | 35.862 | 0.30008 | 0.1 | 8.39 | 35.862
13.745 | 50.54 | 32.716 | 0.30007 | 0.101 | 8.59 | 32.716
13.745 | 50.742 | 29.645 | 0.30007 | 0.101 | 8.792 | 29.645
13.745 | 50.944 | 26.694 | 0.30007 | 0.099 | 8.994 | 26.694
13.745 | 51.142 | 24.074 | 0.30008 | 0.1 | 9.192 | 24.074
13.745 | 51.342 | 22.194 | 0.30007 | 0.101 | 9.392 | 22.194
13.745 | 51.544 | 20.294 | 0.30007 | 0.1 | 9.594 | 20.294
13.745 | 51.744 | 18.487 | 0.30007 | 0.099 | 9.794 | 18.487
13.745 | 51.942 | 16.842 | 0.30008 | 0.2505 | 9.992 | 16.842
13.745 | 52.443 | 13.227 | 0.30008 | 0.2485 | 10.493 | 13.227
13.745 | 52.94 | 10.789 | 0.30009 | 0.2515 | 10.99 | 10.789
13.745 | 53.443 | 9.179 | 0.30007 | 0.2505 | 11.493 | 9.179
13.745 | 53.944 | 7.491 | 0.30008 | 0.2485 | 11.994 | 7.491
13.745 | 54.441 | 6.136 | 0.30008 | 0.251 | 12.491 | 6.136
13.745 | 54.943 | 5.531 | 0.30009 | 0.249 | 12.993 | 5.531
13.745 | 55.441 | 4.809 | 0.30008 | 0.251 | 13.491 | 4.809
13.745 | 55.943 | 4.123 | 0.30008 | 0.251 | 13.993 | 4.123
13.745 | 56.445 | 3.551 | 0.30008 | 0.2485 | 14.495 | 3.551
13.745 | 56.942 | 3.084 | 0.30007 | 0.2505 | 14.992 | 3.084
13.745 | 57.443 | 2.683 | 0.30007 | 0.2495 | 15.493 | 2.683
13.745 | 57.942 | 2.354 | 0.30008 | 0.2495 | 15.992 | 2.354
13.745 | 58.441 | 2.076 | 0.30009 | 0.252 | 16.491 | 2.076
13.745 | 58.945 | 1.843 | 0.30009 | 0.249 | 16.995 | 1.843
13.745 | 59.443 | 1.647 | 0.30008 | 0.2495 | 17.493 | 1.647
13.745 | 59.942 | 1.471 | 0.30008 | 0.2505 | 17.992 | 1.471
13.745 | 60.443 | 1.332 | 0.3001 | 0.249 | 18.493 | 1.332
13.745 | 60.941 | 1.204 | 0.30008 | 0.252 | 18.991 | 1.204
13.745 | 61.445 | 1.098 | 0.30009 | 0.2495 | 19.495 | 1.098
13.745 | 61.944 | 1.008 | 0.30009 | 0.2485 | 19.994 | 1.008
13.745 | 62.441 | 0.924 | 0.30007 | 0.2505 | 20.491 | 0.924
13.745 | 62.942 | 0.857 | 0.30008 | 0.2495 | 20.992 | 0.857
13.745 | 63.441 | 0.792 | 0.30008 | 0.251 | 21.491 | 0.792
13.745 | 63.943 | 0.736 | 0.30006 | 0.25 | 21.993 | 0.736
13.745 | 64.443 | 0.687 | 0.30007 | 0.249 | 22.493 | 0.687
13.745 | 64.941 | 0.64 | 0.30007 | 0.2505 | 22.991 | 0.64
13.745 | 65.442 | 0.603 | 0.30006 | 0.2495 | 23.492 | 0.603
13.745 | 65.941 | 0.571 | 0.30006 | 0.25 | 23.991 | 0.571
13.745 | 66.441 | 0.542 | 0.30007 | 0.2515 | 24.491 | 0.542
13.745 | 66.944 | 0.512 | 0.30007 | 0.25 | 24.994 | 0.512

## CW Run 3A (Field Out)
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002021.stp | ! data file
2020-08-13 | ! date
9:06 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked (<------>), Hyst x 4 (+/-4A) I=3A
! end File Header |  |  |  |  | I | 3.0008 | A
! begin Step Header |  |  |  |  | BdL | -2.3838 | g-cm
0 | ! set current |  |  |  | B_center | 9.208 | g
0 | ! actual current at start of step |  |  |  | B_peak1 | -228.212 | g
0 | ! actual current at end of step |  |  |  | B_Peak2 | 228.13 | g
! end Step Header |  |  |  |  | Eff_L | -0.25889 | cm
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.745 | 16.953 | 0.29 | 0.30009 | 0.246 | -24.997 | 0.29
13.745 | 17.445 | 0.296 | 0.30007 | 0.25 | -24.505 | 0.296
13.745 | 17.945 | 0.318 | 0.30008 | 0.25 | -24.005 | 0.318
13.745 | 18.445 | 0.352 | 0.30009 | 0.2505 | -23.505 | 0.352
13.745 | 18.946 | 0.388 | 0.30007 | 0.249 | -23.004 | 0.388
13.745 | 19.444 | 0.429 | 0.30007 | 0.249 | -22.506 | 0.429
13.745 | 19.942 | 0.475 | 0.30009 | 0.252 | -22.008 | 0.475
13.745 | 20.446 | 0.526 | 0.30009 | 0.2495 | -21.504 | 0.526
13.745 | 20.945 | 0.585 | 0.30009 | 0.2495 | -21.005 | 0.585
13.745 | 21.444 | 0.649 | 0.30007 | 0.2505 | -20.506 | 0.649
13.745 | 21.945 | 0.724 | 0.30008 | 0.249 | -20.005 | 0.724
13.745 | 22.443 | 0.805 | 0.30008 | 0.252 | -19.507 | 0.805
13.745 | 22.947 | 0.895 | 0.30007 | 0.249 | -19.003 | 0.895
13.745 | 23.445 | 1.001 | 0.30007 | 0.2485 | -18.505 | 1.001
13.745 | 23.942 | 1.124 | 0.30007 | 0.2515 | -18.008 | 1.124
13.745 | 24.445 | 1.267 | 0.30008 | 0.2485 | -17.505 | 1.267
13.745 | 24.942 | 1.424 | 0.30007 | 0.251 | -17.008 | 1.424
13.745 | 25.444 | 1.612 | 0.30007 | 0.251 | -16.506 | 1.612
13.745 | 25.946 | 1.837 | 0.30008 | 0.2485 | -16.004 | 1.837
13.745 | 26.443 | 2.085 | 0.30008 | 0.251 | -15.507 | 2.085
13.745 | 26.945 | 2.414 | 0.30008 | 0.249 | -15.005 | 2.414
13.745 | 27.443 | 2.777 | 0.30007 | 0.2495 | -14.507 | 2.777
13.745 | 27.942 | 3.202 | 0.30007 | 0.252 | -14.008 | 3.202
13.745 | 28.446 | 3.76 | 0.30007 | 0.2485 | -13.504 | 3.76
13.745 | 28.943 | 4.406 | 0.30008 | 0.25 | -13.007 | 4.406
13.745 | 29.443 | 5.231 | 0.30009 | 0.25 | -12.507 | 5.231
13.745 | 29.943 | 6.697 | 0.30007 | 0.249 | -12.007 | 6.697
13.745 | 30.441 | 8.164 | 0.30007 | 0.2525 | -11.509 | 8.164
13.745 | 30.946 | 9.779 | 0.30008 | 0.249 | -11.004 | 9.779
13.745 | 31.444 | 12.344 | 0.30008 | 0.2485 | -10.506 | 12.344
13.745 | 31.941 | 15.063 | 0.30009 | 0.1005 | -10.009 | 15.063
13.745 | 32.142 | 16.656 | 0.30009 | 0.1005 | -9.808 | 16.656
13.745 | 32.343 | 18.172 | 0.3001 | 0.1 | -9.607 | 18.172
13.745 | 32.543 | 19.915 | 0.30008 | 0.099 | -9.407 | 19.915
13.745 | 32.741 | 21.943 | 0.3001 | 0.1005 | -9.209 | 21.943
13.745 | 32.942 | 24.564 | 0.30007 | 0.1015 | -9.008 | 24.564
13.745 | 33.145 | 27.147 | 0.30007 | 0.1005 | -8.805 | 27.147
13.745 | 33.346 | 30.046 | 0.30007 | 0.099 | -8.604 | 30.046
13.745 | 33.544 | 32.869 | 0.30008 | 0.1 | -8.406 | 32.869
13.745 | 33.744 | 36.623 | 0.30009 | 0.1005 | -8.206 | 36.623
13.745 | 33.945 | 40.436 | 0.30007 | 0.1 | -8.005 | 40.436
13.745 | 34.145 | 44.699 | 0.30008 | 0.0985 | -7.805 | 44.699
13.745 | 34.342 | 50.127 | 0.30007 | 0.1 | -7.608 | 50.127
13.745 | 34.542 | 55.653 | 0.30007 | 0.1005 | -7.408 | 55.653
13.745 | 34.743 | 61.765 | 0.30007 | 0.0995 | -7.207 | 61.765
13.745 | 34.942 | 68.168 | 0.30008 | 0.099 | -7.008 | 68.168
13.745 | 35.14 | 75.646 | 0.30008 | 0.1005 | -6.81 | 75.646
13.745 | 35.341 | 83.635 | 0.30008 | 0.1015 | -6.609 | 83.635
13.745 | 35.544 | 92.529 | 0.30009 | 0.1005 | -6.406 | 92.529
13.745 | 35.745 | 101.681 | 0.30008 | 0.0985 | -6.205 | 101.681
13.745 | 35.942 | 111.261 | 0.30009 | 0.101 | -6.008 | 111.261
13.745 | 36.144 | 121.822 | 0.30008 | 0.1005 | -5.806 | 121.822
13.745 | 36.345 | 132.402 | 0.30007 | 0.1005 | -5.605 | 132.402
13.745 | 36.546 | 143.443 | 0.30007 | 0.098 | -5.404 | 143.443
13.745 | 36.742 | 154.303 | 0.30007 | 0.1 | -5.208 | 154.303
13.745 | 36.942 | 165.276 | 0.30008 | 0.1005 | -5.008 | 165.276
13.745 | 37.143 | 175.78 | 0.30007 | 0.1 | -4.807 | 175.78
13.745 | 37.343 | 185.978 | 0.30007 | 0.0975 | -4.607 | 185.978
13.745 | 37.538 | 195.391 | 0.30009 | 0.101 | -4.412 | 195.391
13.745 | 37.74 | 204.01 | 0.30008 | 0.101 | -4.21 | 204.01
13.745 | 37.942 | 211.393 | 0.30008 | 0.101 | -4.008 | 211.393
13.745 | 38.144 | 218.081 | 0.30009 | 0.099 | -3.806 | 218.081
13.745 | 38.342 | 222.612 | 0.30008 | 0.1 | -3.608 | 222.612
13.745 | 38.542 | 225.941 | 0.30007 | 0.1015 | -3.408 | 225.941
13.745 | 38.745 | 227.809 | 0.30008 | 0.1 | -3.205 | 227.809
13.745 | 38.945 | 228.13 | 0.3001 | 0.0985 | -3.005 | 228.13
13.745 | 39.142 | 226.756 | 0.30009 | 0.1 | -2.808 | 226.756
13.745 | 39.342 | 223.442 | 0.30008 | 0.1005 | -2.608 | 223.442
13.745 | 39.543 | 217.89 | 0.30007 | 0.1 | -2.407 | 217.89
13.745 | 39.743 | 210.219 | 0.30007 | 0.098 | -2.207 | 210.219
13.745 | 39.939 | 201.04 | 0.30007 | 0.1 | -2.011 | 201.04
13.745 | 40.139 | 189.538 | 0.30007 | 0.1015 | -1.811 | 189.538
13.745 | 40.342 | 175.478 | 0.30008 | 0.1005 | -1.608 | 175.478
13.745 | 40.543 | 160.269 | 0.30008 | 0.0995 | -1.407 | 160.269
13.745 | 40.742 | 142.923 | 0.30008 | 0.1005 | -1.208 | 142.923
13.745 | 40.943 | 123.491 | 0.30008 | 0.101 | -1.007 | 123.491
13.745 | 41.145 | 102.541 | 0.30007 | 0.1 | -0.805 | 102.541
13.745 | 41.345 | 80.3 | 0.30007 | 0.099 | -0.605 | 80.3
13.745 | 41.543 | 57.648 | 0.30008 | 0.1 | -0.407 | 57.648
13.745 | 41.743 | 33.626 | 0.30007 | 0.1005 | -0.207 | 33.626
13.745 | 41.944 | 9.208 | 0.30008 | 0.1 | -0.006 | 9.208
13.745 | 42.144 | -15.852 | 0.3001 | 0.098 | 0.194 | -15.852
13.745 | 42.34 | -39.559 | 0.30008 | 0.1 | 0.39 | -39.559
13.745 | 42.54 | -63.492 | 0.30008 | 0.101 | 0.59 | -63.492
13.745 | 42.742 | -86.796 | 0.30007 | 0.1005 | 0.792 | -86.796
13.745 | 42.943 | -108.594 | 0.30008 | 0.099 | 0.993 | -108.594
13.745 | 43.141 | -128.175 | 0.30008 | 0.1005 | 1.191 | -128.175
13.745 | 43.342 | -147.326 | 0.30009 | 0.1015 | 1.392 | -147.326
13.745 | 43.545 | -164.529 | 0.30008 | 0.1005 | 1.595 | -164.529
13.745 | 43.746 | -179.165 | 0.30008 | 0.0985 | 1.796 | -179.165
13.745 | 43.943 | -192.139 | 0.30008 | 0.1 | 1.993 | -192.139
13.745 | 44.143 | -203.142 | 0.30009 | 0.1005 | 2.193 | -203.142
13.745 | 44.344 | -212.097 | 0.30007 | 0.1 | 2.394 | -212.097
13.745 | 44.544 | -219.135 | 0.30009 | 0.0985 | 2.594 | -219.135
13.745 | 44.741 | -223.782 | 0.30009 | 0.0995 | 2.791 | -223.782
13.745 | 44.94 | -226.643 | 0.3001 | 0.101 | 2.99 | -226.643
13.745 | 45.142 | -228.212 | 0.30008 | 0.1005 | 3.192 | -228.212
13.745 | 45.343 | -228.07 | 0.30008 | 0.0995 | 3.393 | -228.07
13.745 | 45.542 | -226.102 | 0.30008 | 0.1 | 3.592 | -226.102
13.745 | 45.742 | -221.916 | 0.30008 | 0.1015 | 3.792 | -221.916
13.745 | 45.945 | -216.57 | 0.30009 | 0.1005 | 3.995 | -216.57
13.745 | 46.146 | -210.025 | 0.30009 | 0.099 | 4.196 | -210.025
13.745 | 46.344 | -202.144 | 0.30008 | 0.0995 | 4.394 | -202.144
13.745 | 46.543 | -192.964 | 0.30008 | 0.101 | 4.593 | -192.964
13.745 | 46.745 | -183.312 | 0.3001 | 0.1 | 4.795 | -183.312
13.745 | 46.945 | -172.769 | 0.30008 | 0.0985 | 4.995 | -172.769
13.745 | 47.142 | -162.437 | 0.3001 | 0.0995 | 5.192 | -162.437
13.745 | 47.341 | -151.149 | 0.30007 | 0.1005 | 5.391 | -151.149
13.745 | 47.542 | -140.072 | 0.30011 | 0.1005 | 5.592 | -140.072
13.745 | 47.743 | -129.504 | 0.30007 | 0.099 | 5.793 | -129.504
13.745 | 47.941 | -118.86 | 0.30008 | 0.1005 | 5.991 | -118.86
13.745 | 48.142 | -108.649 | 0.30008 | 0.101 | 6.192 | -108.649
13.745 | 48.344 | -98.677 | 0.30007 | 0.1005 | 6.394 | -98.677
13.745 | 48.545 | -89.789 | 0.30009 | 0.099 | 6.595 | -89.789
13.745 | 48.743 | -81.343 | 0.30007 | 0.1 | 6.793 | -81.343
13.745 | 48.943 | -73.525 | 0.30008 | 0.101 | 6.993 | -73.525
13.745 | 49.145 | -65.972 | 0.30008 | 0.1 | 7.195 | -65.972
13.745 | 49.345 | -59.626 | 0.30008 | 0.0985 | 7.395 | -59.626
13.745 | 49.542 | -53.648 | 0.30008 | 0.1 | 7.592 | -53.648
13.745 | 49.742 | -48.767 | 0.30008 | 0.1005 | 7.792 | -48.767
13.745 | 49.943 | -43.746 | 0.30008 | 0.1 | 7.993 | -43.746
13.745 | 50.143 | -39.586 | 0.30008 | 0.099 | 8.193 | -39.586
13.745 | 50.341 | -35.606 | 0.3001 | 0.1 | 8.391 | -35.606
13.745 | 50.541 | -31.88 | 0.30009 | 0.101 | 8.591 | -31.88
13.745 | 50.743 | -29.074 | 0.30007 | 0.101 | 8.793 | -29.074
13.745 | 50.945 | -26.425 | 0.30007 | 0.099 | 8.995 | -26.425
13.745 | 51.143 | -23.684 | 0.30008 | 0.1 | 9.193 | -23.684
13.745 | 51.343 | -21.863 | 0.30009 | 0.101 | 9.393 | -21.863
13.745 | 51.545 | -19.449 | 0.30008 | 0.1005 | 9.595 | -19.449
13.745 | 51.746 | -17.864 | 0.3001 | 0.0985 | 9.796 | -17.864
13.745 | 51.943 | -16.375 | 0.30008 | 0.2505 | 9.993 | -16.375
13.745 | 52.444 | -13.344 | 0.3001 | 0.2485 | 10.494 | -13.344
13.745 | 52.941 | -10.659 | 0.30009 | 0.2515 | 10.991 | -10.659
13.745 | 53.444 | -8.816 | 0.30007 | 0.25 | 11.494 | -8.816
13.745 | 53.944 | -7.125 | 0.30008 | 0.2495 | 11.994 | -7.125
13.745 | 54.443 | -5.795 | 0.30009 | 0.2505 | 12.493 | -5.795
13.745 | 54.944 | -5.213 | 0.30008 | 0.249 | 12.994 | -5.213
13.745 | 55.442 | -4.494 | 0.30009 | 0.251 | 13.492 | -4.494
13.745 | 55.944 | -3.8 | 0.30008 | 0.251 | 13.994 | -3.8
13.745 | 56.446 | -3.227 | 0.30008 | 0.2485 | 14.496 | -3.227
13.745 | 56.943 | -2.745 | 0.30008 | 0.2505 | 14.993 | -2.745
13.745 | 57.444 | -2.356 | 0.30007 | 0.2495 | 15.494 | -2.356
13.745 | 57.943 | -2.012 | 0.30008 | 0.2495 | 15.993 | -2.012
13.745 | 58.442 | -1.73 | 0.30009 | 0.252 | 16.492 | -1.73
13.745 | 58.946 | -1.498 | 0.30008 | 0.249 | 16.996 | -1.498
13.745 | 59.444 | -1.291 | 0.30009 | 0.2495 | 17.494 | -1.291
13.745 | 59.943 | -1.119 | 0.30009 | 0.25 | 17.993 | -1.119
13.745 | 60.443 | -0.967 | 0.30008 | 0.2495 | 18.493 | -0.967
13.745 | 60.942 | -0.84 | 0.30007 | 0.252 | 18.992 | -0.84
13.745 | 61.446 | -0.733 | 0.30007 | 0.249 | 19.496 | -0.733
13.745 | 61.944 | -0.638 | 0.30008 | 0.249 | 19.994 | -0.638
13.745 | 62.442 | -0.552 | 0.30009 | 0.2505 | 20.492 | -0.552
13.745 | 62.943 | -0.474 | 0.30007 | 0.2495 | 20.993 | -0.474
13.745 | 63.442 | -0.412 | 0.30008 | 0.251 | 21.492 | -0.412
13.745 | 63.944 | -0.353 | 0.30008 | 0.25 | 21.994 | -0.353
13.745 | 64.444 | -0.297 | 0.30011 | 0.249 | 22.494 | -0.297
13.745 | 64.942 | -0.252 | 0.30007 | 0.2505 | 22.992 | -0.252
13.745 | 65.443 | -0.213 | 0.30007 | 0.25 | 23.493 | -0.213
13.745 | 65.943 | -0.178 | 0.30008 | 0.2495 | 23.993 | -0.178
13.745 | 66.442 | -0.141 | 0.30009 | 0.2515 | 24.492 | -0.141
13.745 | 66.945 | -0.114 | 0.30008 | 0.25 | 24.995 | -0.114

## SW Run 4A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002005.stp | ! data file
2020-08-11 | ! date
9:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 4 (+/-4A), I=4A
! end File Header |  |  |  |  | I | 3.9995 | A
! begin Step Header |  |  |  |  | BdL | -2703.9951 | g-cm
0 | ! set current |  |  |  | B_center | -467.71 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.7813 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.942 | -1.365 | 0.39998 | 0.247 | -25.008 | -1.365
13.744 | 17.436 | -1.225 | 0.39996 | 0.2485 | -24.514 | -1.225
13.744 | 17.933 | -1.225 | 0.39997 | 0.2515 | -24.017 | -1.225
13.744 | 18.436 | -1.284 | 0.39997 | 0.2495 | -23.514 | -1.284
13.744 | 18.935 | -1.356 | 0.39995 | 0.2485 | -23.015 | -1.356
13.744 | 19.432 | -1.449 | 0.39995 | 0.251 | -22.518 | -1.449
13.744 | 19.934 | -1.547 | 0.39995 | 0.25 | -22.016 | -1.547
13.744 | 20.434 | -1.652 | 0.39995 | 0.2495 | -21.516 | -1.652
13.744 | 20.933 | -1.769 | 0.39994 | 0.251 | -21.017 | -1.769
13.744 | 21.435 | -1.9 | 0.39995 | 0.249 | -20.515 | -1.9
13.744 | 21.933 | -2.053 | 0.39995 | 0.25 | -20.017 | -2.053
13.744 | 22.433 | -2.217 | 0.39997 | 0.2505 | -19.517 | -2.217
13.744 | 22.934 | -2.409 | 0.39995 | 0.2495 | -19.016 | -2.409
13.744 | 23.433 | -2.619 | 0.39996 | 0.251 | -18.517 | -2.619
13.744 | 23.935 | -2.843 | 0.39995 | 0.2495 | -18.015 | -2.843
13.744 | 24.434 | -3.113 | 0.39997 | 0.2485 | -17.516 | -3.113
13.744 | 24.931 | -3.407 | 0.39996 | 0.252 | -17.019 | -3.407
13.744 | 25.435 | -3.76 | 0.39995 | 0.2495 | -16.515 | -3.76
13.744 | 25.934 | -4.163 | 0.39997 | 0.2505 | -16.016 | -4.163
13.744 | 26.435 | -4.637 | 0.39995 | 0.25 | -15.515 | -4.637
13.744 | 26.935 | -5.177 | 0.39995 | 0.248 | -15.015 | -5.177
13.744 | 27.431 | -5.795 | 0.39996 | 0.252 | -14.519 | -5.795
13.744 | 27.935 | -6.568 | 0.39996 | 0.249 | -14.015 | -6.568
13.744 | 28.433 | -7.944 | 0.39995 | 0.25 | -13.517 | -7.944
13.744 | 28.933 | -9.306 | 0.39994 | 0.2505 | -13.017 | -9.306
13.744 | 29.434 | -10.62 | 0.39995 | 0.2485 | -12.516 | -10.62
13.744 | 29.931 | -12.162 | 0.39995 | 0.251 | -12.019 | -12.162
13.744 | 30.433 | -14.756 | 0.39995 | 0.251 | -11.517 | -14.756
13.744 | 30.935 | -17.335 | 0.39996 | 0.2485 | -11.015 | -17.335
13.744 | 31.432 | -21.232 | 0.39997 | 0.251 | -10.518 | -21.232
13.744 | 31.934 | -25.893 | 0.39995 | 0.0985 | -10.016 | -25.893
13.744 | 32.131 | -28.212 | 0.39995 | 0.0995 | -9.819 | -28.212
13.744 | 32.33 | -30.616 | 0.39996 | 0.101 | -9.62 | -30.616
13.744 | 32.532 | -33.328 | 0.39996 | 0.1005 | -9.418 | -33.328
13.744 | 32.733 | -36.296 | 0.39995 | 0.099 | -9.217 | -36.296
13.744 | 32.931 | -39.611 | 0.39996 | 0.1005 | -9.019 | -39.611
13.744 | 33.132 | -43.663 | 0.39995 | 0.101 | -8.818 | -43.663
13.744 | 33.334 | -48.156 | 0.39995 | 0.1005 | -8.616 | -48.156
13.744 | 33.535 | -52.91 | 0.39995 | 0.099 | -8.415 | -52.91
13.744 | 33.733 | -58.216 | 0.39996 | 0.0995 | -8.217 | -58.216
13.744 | 33.932 | -64.391 | 0.39994 | 0.1005 | -8.018 | -64.391
13.744 | 34.133 | -71.307 | 0.39996 | 0.1005 | -7.817 | -71.307
13.744 | 34.334 | -78.751 | 0.39995 | 0.0985 | -7.616 | -78.751
13.744 | 34.531 | -87.449 | 0.39995 | 0.0995 | -7.419 | -87.449
13.744 | 34.73 | -96.611 | 0.39995 | 0.1005 | -7.22 | -96.611
13.744 | 34.931 | -106.9 | 0.39996 | 0.1005 | -7.019 | -106.9
13.744 | 35.132 | -118.438 | 0.39995 | 0.099 | -6.818 | -118.438
13.744 | 35.33 | -130.795 | 0.39995 | 0.1005 | -6.62 | -130.795
13.744 | 35.531 | -144.508 | 0.39995 | 0.101 | -6.419 | -144.508
13.744 | 35.733 | -159.103 | 0.39996 | 0.101 | -6.217 | -159.103
13.744 | 35.935 | -174.708 | 0.39996 | 0.099 | -6.015 | -174.708
13.744 | 36.133 | -191.017 | 0.39996 | 0.0995 | -5.817 | -191.017
13.744 | 36.332 | -207.9 | 0.39995 | 0.101 | -5.618 | -207.9
13.744 | 36.534 | -226.312 | 0.39996 | 0.1 | -5.416 | -226.312
13.744 | 36.734 | -244.85 | 0.39995 | 0.099 | -5.216 | -244.85
13.744 | 36.932 | -262.977 | 0.39995 | 0.0995 | -5.018 | -262.977
13.744 | 37.131 | -281.542 | 0.39994 | 0.1005 | -4.819 | -281.542
13.744 | 37.332 | -299.593 | 0.39996 | 0.0995 | -4.618 | -299.593
13.744 | 37.531 | -317.714 | 0.39994 | 0.0985 | -4.419 | -317.714
13.744 | 37.728 | -334.788 | 0.39996 | 0.1005 | -4.222 | -334.788
13.744 | 37.929 | -351.561 | 0.39995 | 0.1015 | -4.021 | -351.561
13.744 | 38.132 | -367.121 | 0.39995 | 0.101 | -3.818 | -367.121
13.744 | 38.334 | -381.596 | 0.39996 | 0.099 | -3.616 | -381.596
13.744 | 38.532 | -394.754 | 0.39996 | 0.1 | -3.418 | -394.754
13.744 | 38.732 | -406.521 | 0.39996 | 0.1005 | -3.218 | -406.521
13.744 | 38.933 | -417.209 | 0.39996 | 0.1005 | -3.017 | -417.209
13.744 | 39.134 | -426.874 | 0.39996 | 0.0985 | -2.816 | -426.874
13.744 | 39.331 | -434.552 | 0.39995 | 0.0995 | -2.619 | -434.552
13.744 | 39.53 | -441.607 | 0.39996 | 0.1005 | -2.42 | -441.607
13.744 | 39.731 | -447.493 | 0.39996 | 0.1 | -2.219 | -447.493
13.744 | 39.931 | -452.984 | 0.39996 | 0.0985 | -2.019 | -452.984
13.744 | 40.128 | -456.819 | 0.39996 | 0.1005 | -1.822 | -456.819
13.744 | 40.329 | -459.885 | 0.39995 | 0.101 | -1.621 | -459.885
13.744 | 40.531 | -462.66 | 0.39996 | 0.1015 | -1.419 | -462.66
13.744 | 40.734 | -464.182 | 0.39996 | 0.099 | -1.216 | -464.182
13.744 | 40.932 | -465.725 | 0.39997 | 0.1 | -1.018 | -465.725
13.744 | 41.132 | -466.269 | 0.39996 | 0.101 | -0.818 | -466.269
13.744 | 41.334 | -466.872 | 0.39995 | 0.1005 | -0.616 | -466.872
13.744 | 41.535 | -467.318 | 0.39994 | 0.099 | -0.415 | -467.318
13.744 | 41.733 | -467.587 | 0.39995 | 0.099 | -0.217 | -467.587
13.744 | 41.931 | -467.71 | 0.39995 | 0.1005 | -0.019 | -467.71
13.744 | 42.132 | -467.713 | 0.39996 | 0.1 | 0.182 | -467.713
13.744 | 42.332 | -467.605 | 0.39996 | 0.099 | 0.382 | -467.605
13.744 | 42.53 | -467.373 | 0.39994 | 0.0995 | 0.58 | -467.373
13.744 | 42.729 | -467.004 | 0.39994 | 0.1015 | 0.779 | -467.004
13.744 | 42.932 | -466.437 | 0.39994 | 0.1005 | 0.982 | -466.437
13.744 | 43.133 | -465.025 | 0.39995 | 0.1 | 1.183 | -465.025
13.744 | 43.333 | -463.539 | 0.39995 | 0.0995 | 1.383 | -463.539
13.744 | 43.532 | -461.808 | 0.39995 | 0.101 | 1.582 | -461.808
13.744 | 43.734 | -458.947 | 0.39997 | 0.1005 | 1.784 | -458.947
13.744 | 43.935 | -455.32 | 0.39995 | 0.099 | 1.985 | -455.32
13.744 | 44.133 | -451.352 | 0.39995 | 0.099 | 2.183 | -451.352
13.744 | 44.331 | -446.029 | 0.39995 | 0.101 | 2.381 | -446.029
13.744 | 44.533 | -439.886 | 0.39995 | 0.1 | 2.583 | -439.886
13.744 | 44.733 | -432.721 | 0.39994 | 0.0985 | 2.783 | -432.721
13.744 | 44.93 | -424.221 | 0.39994 | 0.0995 | 2.98 | -424.221
13.744 | 45.129 | -414.39 | 0.39995 | 0.101 | 3.179 | -414.39
13.744 | 45.331 | -403.538 | 0.39994 | 0.101 | 3.381 | -403.538
13.744 | 45.533 | -390.876 | 0.39995 | 0.0995 | 3.583 | -390.876
13.744 | 45.732 | -377.69 | 0.39995 | 0.1 | 3.782 | -377.69
13.744 | 45.932 | -363.307 | 0.39995 | 0.101 | 3.982 | -363.307
13.744 | 46.134 | -347.118 | 0.39995 | 0.1005 | 4.184 | -347.118
13.744 | 46.335 | -330.487 | 0.39994 | 0.099 | 4.385 | -330.487
13.744 | 46.533 | -313.432 | 0.39994 | 0.0995 | 4.583 | -313.432
13.744 | 46.732 | -295.251 | 0.39995 | 0.1005 | 4.782 | -295.251
13.744 | 46.933 | -276.829 | 0.39994 | 0.1 | 4.983 | -276.829
13.744 | 47.133 | -258.634 | 0.39994 | 0.099 | 5.183 | -258.634
13.744 | 47.331 | -239.898 | 0.39995 | 0.099 | 5.381 | -239.898
13.744 | 47.529 | -221.755 | 0.39996 | 0.101 | 5.579 | -221.755
13.744 | 47.731 | -203.796 | 0.39995 | 0.101 | 5.781 | -203.796
13.744 | 47.933 | -186.572 | 0.39994 | 0.0995 | 5.983 | -186.572
13.744 | 48.132 | -170.473 | 0.39994 | 0.0995 | 6.182 | -170.473
13.744 | 48.331 | -155.107 | 0.39994 | 0.101 | 6.381 | -155.107
13.744 | 48.533 | -140.762 | 0.39996 | 0.101 | 6.583 | -140.762
13.744 | 48.735 | -127.26 | 0.39995 | 0.0995 | 6.785 | -127.26
13.744 | 48.934 | -115.532 | 0.39995 | 0.099 | 6.984 | -115.532
13.744 | 49.132 | -104.359 | 0.39994 | 0.1005 | 7.182 | -104.359
13.744 | 49.333 | -94.51 | 0.39994 | 0.1005 | 7.383 | -94.51
13.744 | 49.534 | -85.476 | 0.39994 | 0.099 | 7.584 | -85.476
13.744 | 49.732 | -76.883 | 0.39994 | 0.099 | 7.782 | -76.883
13.744 | 49.93 | -69.64 | 0.39996 | 0.1005 | 7.98 | -69.64
13.744 | 50.131 | -63.253 | 0.39995 | 0.101 | 8.181 | -63.253
13.744 | 50.333 | -57.012 | 0.39993 | 0.099 | 8.383 | -57.012
13.744 | 50.531 | -51.621 | 0.39994 | 0.0995 | 8.581 | -51.621
13.744 | 50.73 | -47.358 | 0.39995 | 0.101 | 8.78 | -47.358
13.744 | 50.932 | -43.03 | 0.39994 | 0.101 | 8.982 | -43.03
13.744 | 51.134 | -38.908 | 0.39996 | 0.0995 | 9.184 | -38.908
13.744 | 51.333 | -35.894 | 0.39996 | 0.0995 | 9.383 | -35.894
13.744 | 51.532 | -32.903 | 0.39998 | 0.1005 | 9.582 | -32.903
13.744 | 51.733 | -29.986 | 0.39995 | 0.1005 | 9.783 | -29.986
13.744 | 51.934 | -27.285 | 0.39996 | 0.2485 | 9.984 | -27.285
13.744 | 52.431 | -22.337 | 0.39995 | 0.25 | 10.481 | -22.337
13.744 | 52.931 | -18.659 | 0.39995 | 0.251 | 10.981 | -18.659
13.744 | 53.433 | -15.673 | 0.39996 | 0.2495 | 11.483 | -15.673
13.744 | 53.932 | -13.014 | 0.39995 | 0.2515 | 11.982 | -13.014
13.744 | 54.435 | -11.284 | 0.39995 | 0.2485 | 12.485 | -11.284
13.744 | 54.932 | -9.651 | 0.39995 | 0.249 | 12.982 | -9.651
13.744 | 55.43 | -8.258 | 0.39995 | 0.252 | 13.48 | -8.258
13.744 | 55.934 | -7.549 | 0.39995 | 0.2495 | 13.984 | -7.549
13.744 | 56.433 | -6.723 | 0.39996 | 0.2505 | 14.483 | -6.723
13.744 | 56.934 | -5.97 | 0.39995 | 0.2495 | 14.984 | -5.97
13.744 | 57.433 | -5.271 | 0.39994 | 0.2485 | 15.483 | -5.271
13.744 | 57.93 | -4.692 | 0.39996 | 0.252 | 15.98 | -4.692
13.744 | 58.434 | -4.183 | 0.39995 | 0.25 | 16.484 | -4.183
13.744 | 58.934 | -3.752 | 0.39995 | 0.249 | 16.984 | -3.752
13.744 | 59.432 | -3.373 | 0.39995 | 0.251 | 17.482 | -3.373
13.744 | 59.934 | -3.052 | 0.39995 | 0.2485 | 17.984 | -3.052
13.744 | 60.431 | -2.768 | 0.39996 | 0.2515 | 18.481 | -2.768
13.744 | 60.934 | -2.516 | 0.39995 | 0.25 | 18.984 | -2.516
13.744 | 61.434 | -2.287 | 0.39996 | 0.249 | 19.484 | -2.287
13.744 | 61.932 | -2.09 | 0.39996 | 0.251 | 19.982 | -2.09
13.744 | 62.434 | -1.912 | 0.39995 | 0.2485 | 20.484 | -1.912
13.744 | 62.931 | -1.755 | 0.39995 | 0.2495 | 20.981 | -1.755
13.744 | 63.43 | -1.607 | 0.39996 | 0.252 | 21.48 | -1.607
13.744 | 63.934 | -1.478 | 0.39997 | 0.2485 | 21.984 | -1.478
13.744 | 64.431 | -1.358 | 0.39996 | 0.251 | 22.481 | -1.358
13.744 | 64.933 | -1.258 | 0.39997 | 0.2495 | 22.983 | -1.258
13.744 | 65.432 | -1.159 | 0.39998 | 0.249 | 23.482 | -1.159
13.744 | 65.93 | -1.07 | 0.39994 | 0.252 | 23.98 | -1.07
13.744 | 66.434 | -1.001 | 0.39994 | 0.249 | 24.484 | -1.001
13.744 | 66.932 | -0.921 | 0.39995 | 0.25 | 24.982 | -0.921

## SW Run 3A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002006.stp | ! data file
2020-08-11 | ! date
10:09 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=3A
! end File Header |  |  |  |  | I | 3.0008 | A
! begin Step Header |  |  |  |  | BdL | -2031.2099 | g-cm
0 | ! set current |  |  |  | B_center | -353.928 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.739 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.943 | -0.878 | 0.30008 | 0.2465 | -25.007 | -0.878
13.744 | 17.436 | -0.75 | 0.30009 | 0.2485 | -24.514 | -0.75
13.744 | 17.933 | -0.741 | 0.30008 | 0.2505 | -24.017 | -0.741
13.744 | 18.434 | -0.773 | 0.30008 | 0.25 | -23.516 | -0.773
13.744 | 18.934 | -0.811 | 0.30008 | 0.248 | -23.016 | -0.811
13.744 | 19.43 | -0.86 | 0.30007 | 0.251 | -22.52 | -0.86
13.744 | 19.932 | -0.92 | 0.3001 | 0.2505 | -22.018 | -0.92
13.744 | 20.433 | -0.985 | 0.30008 | 0.249 | -21.517 | -0.985
13.744 | 20.931 | -1.055 | 0.30007 | 0.2515 | -21.019 | -1.055
13.744 | 21.434 | -1.139 | 0.30007 | 0.2485 | -20.516 | -1.139
13.744 | 21.931 | -1.229 | 0.30007 | 0.251 | -20.019 | -1.229
13.744 | 22.433 | -1.328 | 0.30008 | 0.2505 | -19.517 | -1.328
13.744 | 22.934 | -1.45 | 0.30008 | 0.249 | -19.016 | -1.45
13.744 | 23.432 | -1.586 | 0.30008 | 0.251 | -18.518 | -1.586
13.744 | 23.934 | -1.738 | 0.30008 | 0.249 | -18.016 | -1.738
13.744 | 24.432 | -1.914 | 0.30008 | 0.249 | -17.518 | -1.914
13.744 | 24.93 | -2.108 | 0.30009 | 0.252 | -17.02 | -2.108
13.744 | 25.434 | -2.343 | 0.30009 | 0.249 | -16.516 | -2.343
13.744 | 25.932 | -2.613 | 0.3001 | 0.2505 | -16.018 | -2.613
13.744 | 26.433 | -2.939 | 0.30008 | 0.25 | -15.517 | -2.939
13.744 | 26.933 | -3.308 | 0.30008 | 0.248 | -15.017 | -3.308
13.744 | 27.429 | -3.743 | 0.30009 | 0.2525 | -14.521 | -3.743
13.744 | 27.934 | -4.274 | 0.30008 | 0.2495 | -14.016 | -4.274
13.744 | 28.433 | -4.898 | 0.30007 | 0.249 | -13.517 | -4.898
13.744 | 28.931 | -5.685 | 0.30008 | 0.251 | -13.019 | -5.685
13.744 | 29.433 | -7.122 | 0.30009 | 0.2485 | -12.517 | -7.122
13.744 | 29.93 | -8.458 | 0.30008 | 0.251 | -12.02 | -8.458
13.744 | 30.432 | -10.012 | 0.30008 | 0.251 | -11.518 | -10.012
13.744 | 30.934 | -11.822 | 0.30008 | 0.2485 | -11.016 | -11.822
13.744 | 31.431 | -14.648 | 0.3001 | 0.251 | -10.519 | -14.648
13.744 | 31.933 | -18.432 | 0.30009 | 0.0985 | -10.017 | -18.432
13.744 | 32.13 | -19.901 | 0.30008 | 0.0995 | -9.82 | -19.901
13.744 | 32.329 | -21.575 | 0.30008 | 0.101 | -9.621 | -21.575
13.744 | 32.531 | -24.041 | 0.30008 | 0.1005 | -9.419 | -24.041
13.744 | 32.732 | -26.013 | 0.30009 | 0.099 | -9.218 | -26.013
13.744 | 32.93 | -28.926 | 0.30007 | 0.101 | -9.02 | -28.926
13.744 | 33.132 | -31.743 | 0.30009 | 0.101 | -8.818 | -31.743
13.744 | 33.334 | -34.849 | 0.3001 | 0.1005 | -8.616 | -34.849
13.744 | 33.535 | -38.906 | 0.30008 | 0.0985 | -8.415 | -38.906
13.744 | 33.732 | -43.062 | 0.30008 | 0.1 | -8.218 | -43.062
13.744 | 33.932 | -47.508 | 0.30007 | 0.1005 | -8.018 | -47.508
13.744 | 34.133 | -52.834 | 0.30008 | 0.1 | -7.817 | -52.834
13.744 | 34.333 | -58.461 | 0.30008 | 0.0985 | -7.617 | -58.461
13.744 | 34.53 | -64.648 | 0.30008 | 0.0995 | -7.42 | -64.648
13.744 | 34.729 | -71.802 | 0.30008 | 0.101 | -7.221 | -71.802
13.744 | 34.931 | -79.951 | 0.30007 | 0.1005 | -7.019 | -79.951
13.744 | 35.132 | -88.172 | 0.30009 | 0.099 | -6.818 | -88.172
13.744 | 35.33 | -97.862 | 0.3001 | 0.1 | -6.62 | -97.862
13.744 | 35.53 | -107.891 | 0.30008 | 0.1015 | -6.42 | -107.891
13.744 | 35.733 | -119.337 | 0.30008 | 0.1005 | -6.217 | -119.337
13.744 | 35.934 | -131.092 | 0.30008 | 0.099 | -6.016 | -131.092
13.744 | 36.132 | -143.683 | 0.30008 | 0.1 | -5.818 | -143.683
13.744 | 36.332 | -156.838 | 0.30008 | 0.101 | -5.618 | -156.838
13.744 | 36.534 | -170.187 | 0.30008 | 0.1 | -5.416 | -170.187
13.744 | 36.734 | -184.343 | 0.30009 | 0.0985 | -5.216 | -184.343
13.744 | 36.931 | -198.459 | 0.30009 | 0.0995 | -5.019 | -198.459
13.744 | 37.13 | -212.573 | 0.30009 | 0.1005 | -4.82 | -212.573
13.744 | 37.331 | -226.529 | 0.30008 | 0.0995 | -4.619 | -226.529
13.744 | 37.53 | -240.035 | 0.30007 | 0.099 | -4.42 | -240.035
13.744 | 37.728 | -253.256 | 0.30008 | 0.1005 | -4.222 | -253.256
13.744 | 37.929 | -265.745 | 0.30011 | 0.1015 | -4.021 | -265.745
13.744 | 38.132 | -277.326 | 0.30008 | 0.1005 | -3.818 | -277.326
13.744 | 38.333 | -288.612 | 0.30008 | 0.099 | -3.617 | -288.612
13.744 | 38.531 | -298.555 | 0.30009 | 0.1 | -3.419 | -298.555
13.744 | 38.731 | -307.67 | 0.30011 | 0.101 | -3.219 | -307.67
13.744 | 38.933 | -315.635 | 0.30007 | 0.1005 | -3.017 | -315.635
13.744 | 39.134 | -322.744 | 0.30008 | 0.098 | -2.816 | -322.744
13.744 | 39.33 | -329.238 | 0.30008 | 0.1 | -2.62 | -329.238
13.744 | 39.53 | -334.251 | 0.30009 | 0.1005 | -2.42 | -334.251
13.744 | 39.731 | -339.274 | 0.30009 | 0.0995 | -2.219 | -339.274
13.744 | 39.93 | -342.96 | 0.30008 | 0.0985 | -2.02 | -342.96
13.744 | 40.127 | -345.688 | 0.30007 | 0.1005 | -1.823 | -345.688
13.744 | 40.328 | -348.297 | 0.30008 | 0.1015 | -1.622 | -348.297
13.744 | 40.531 | -350.005 | 0.30009 | 0.101 | -1.419 | -350.005
13.744 | 40.733 | -351.489 | 0.30008 | 0.0995 | -1.217 | -351.489
13.744 | 40.932 | -352.117 | 0.30008 | 0.1 | -1.018 | -352.117
13.744 | 41.132 | -352.894 | 0.3001 | 0.1005 | -0.818 | -352.894
13.744 | 41.333 | -353.355 | 0.30008 | 0.1005 | -0.617 | -353.355
13.744 | 41.534 | -353.671 | 0.3001 | 0.099 | -0.416 | -353.671
13.744 | 41.732 | -353.876 | 0.30007 | 0.099 | -0.218 | -353.876
13.744 | 41.93 | -353.928 | 0.30008 | 0.101 | -0.02 | -353.928
13.744 | 42.132 | -353.898 | 0.30009 | 0.1 | 0.182 | -353.898
13.744 | 42.332 | -353.769 | 0.30009 | 0.0985 | 0.382 | -353.769
13.744 | 42.529 | -353.579 | 0.30008 | 0.1 | 0.579 | -353.579
13.744 | 42.729 | -353.217 | 0.30009 | 0.101 | 0.779 | -353.217
13.744 | 42.931 | -352.773 | 0.30007 | 0.101 | 0.981 | -352.773
13.744 | 43.133 | -352.121 | 0.3001 | 0.0995 | 1.183 | -352.121
13.744 | 43.332 | -350.53 | 0.30009 | 0.0995 | 1.382 | -350.53
13.744 | 43.531 | -349.033 | 0.30009 | 0.101 | 1.581 | -349.033
13.744 | 43.733 | -347.271 | 0.30008 | 0.1005 | 1.783 | -347.271
13.744 | 43.934 | -344.48 | 0.3001 | 0.0995 | 1.984 | -344.48
13.744 | 44.133 | -341 | 0.30009 | 0.099 | 2.183 | -341
13.744 | 44.331 | -337.173 | 0.30009 | 0.1005 | 2.381 | -337.173
13.744 | 44.532 | -332.666 | 0.30008 | 0.1005 | 2.582 | -332.666
13.744 | 44.733 | -326.879 | 0.30008 | 0.0985 | 2.783 | -326.879
13.744 | 44.93 | -320.373 | 0.30007 | 0.0995 | 2.98 | -320.373
13.744 | 45.129 | -313.214 | 0.30009 | 0.101 | 3.179 | -313.214
13.744 | 45.331 | -304.632 | 0.30008 | 0.101 | 3.381 | -304.632
13.744 | 45.533 | -295.311 | 0.30008 | 0.0995 | 3.583 | -295.311
13.744 | 45.732 | -285.283 | 0.30009 | 0.0995 | 3.782 | -285.283
13.744 | 45.931 | -274.564 | 0.30009 | 0.101 | 3.981 | -274.564
13.744 | 46.133 | -262.24 | 0.30008 | 0.101 | 4.183 | -262.24
13.744 | 46.335 | -249.9 | 0.30008 | 0.099 | 4.385 | -249.9
13.744 | 46.533 | -236.511 | 0.30007 | 0.0995 | 4.583 | -236.511
13.744 | 46.732 | -222.832 | 0.30007 | 0.1005 | 4.782 | -222.832
13.744 | 46.933 | -208.564 | 0.30008 | 0.1 | 4.983 | -208.564
13.744 | 47.133 | -194.841 | 0.30008 | 0.099 | 5.183 | -194.841
13.744 | 47.331 | -180.544 | 0.3001 | 0.099 | 5.381 | -180.544
13.744 | 47.529 | -166.847 | 0.30008 | 0.101 | 5.579 | -166.847
13.744 | 47.731 | -153.516 | 0.30009 | 0.1005 | 5.781 | -153.516
13.744 | 47.932 | -140.342 | 0.30008 | 0.1 | 5.982 | -140.342
13.744 | 48.132 | -128.06 | 0.30008 | 0.0995 | 6.182 | -128.06
13.744 | 48.331 | -116.515 | 0.30007 | 0.101 | 6.381 | -116.515
13.744 | 48.533 | -105.338 | 0.30008 | 0.101 | 6.583 | -105.338
13.744 | 48.735 | -95.606 | 0.30007 | 0.0995 | 6.785 | -95.606
13.744 | 48.934 | -86.324 | 0.30008 | 0.099 | 6.984 | -86.324
13.744 | 49.132 | -77.676 | 0.30008 | 0.1005 | 7.182 | -77.676
13.744 | 49.333 | -69.947 | 0.30008 | 0.1005 | 7.383 | -69.947
13.744 | 49.534 | -62.978 | 0.30008 | 0.099 | 7.584 | -62.978
13.744 | 49.732 | -57.048 | 0.30008 | 0.099 | 7.782 | -57.048
13.744 | 49.93 | -51.473 | 0.30008 | 0.1005 | 7.98 | -51.473
13.744 | 50.131 | -46.484 | 0.30008 | 0.1005 | 8.181 | -46.484
13.744 | 50.332 | -42.115 | 0.30009 | 0.0995 | 8.382 | -42.115
13.744 | 50.531 | -37.98 | 0.30008 | 0.0995 | 8.581 | -37.98
13.744 | 50.73 | -34.737 | 0.30008 | 0.101 | 8.78 | -34.737
13.744 | 50.932 | -31.068 | 0.30008 | 0.101 | 8.982 | -31.068
13.744 | 51.134 | -28.324 | 0.30009 | 0.0995 | 9.184 | -28.324
13.744 | 51.333 | -25.624 | 0.30007 | 0.099 | 9.383 | -25.624
13.744 | 51.531 | -23.828 | 0.30009 | 0.101 | 9.581 | -23.828
13.744 | 51.733 | -21.335 | 0.30008 | 0.1005 | 9.783 | -21.335
13.744 | 51.934 | -19.735 | 0.30008 | 0.2485 | 9.984 | -19.735
13.744 | 52.431 | -15.918 | 0.30008 | 0.25 | 10.481 | -15.918
13.744 | 52.931 | -13.23 | 0.30008 | 0.251 | 10.981 | -13.23
13.744 | 53.433 | -10.657 | 0.30008 | 0.2495 | 11.483 | -10.657
13.744 | 53.932 | -9.138 | 0.30008 | 0.2515 | 11.982 | -9.138
13.744 | 54.435 | -7.628 | 0.30008 | 0.2485 | 12.485 | -7.628
13.744 | 54.932 | -6.312 | 0.30008 | 0.249 | 12.982 | -6.312
13.744 | 55.43 | -5.769 | 0.30008 | 0.252 | 13.48 | -5.769
13.744 | 55.934 | -5.073 | 0.30007 | 0.249 | 13.984 | -5.073
13.744 | 56.432 | -4.428 | 0.30007 | 0.251 | 14.482 | -4.428
13.744 | 56.934 | -3.848 | 0.30009 | 0.2495 | 14.984 | -3.848
13.744 | 57.433 | -3.372 | 0.30008 | 0.248 | 15.483 | -3.372
13.744 | 57.929 | -2.96 | 0.30008 | 0.2525 | 15.979 | -2.96
13.744 | 58.434 | -2.619 | 0.30008 | 0.2495 | 16.484 | -2.619
13.744 | 58.933 | -2.317 | 0.30008 | 0.2495 | 16.983 | -2.317
13.744 | 59.432 | -2.062 | 0.30008 | 0.251 | 17.482 | -2.062
13.744 | 59.934 | -1.85 | 0.30008 | 0.2485 | 17.984 | -1.85
13.744 | 60.431 | -1.654 | 0.30007 | 0.251 | 18.481 | -1.654
13.744 | 60.933 | -1.487 | 0.30008 | 0.2505 | 18.983 | -1.487
13.744 | 61.434 | -1.339 | 0.30007 | 0.249 | 19.484 | -1.339
13.744 | 61.932 | -1.212 | 0.30008 | 0.251 | 19.982 | -1.212
13.744 | 62.434 | -1.098 | 0.3001 | 0.248 | 20.484 | -1.098
13.744 | 62.93 | -0.998 | 0.30009 | 0.25 | 20.98 | -0.998
13.744 | 63.43 | -0.906 | 0.30008 | 0.2515 | 21.48 | -0.906
13.744 | 63.933 | -0.819 | 0.30009 | 0.249 | 21.983 | -0.819
13.744 | 64.431 | -0.745 | 0.30008 | 0.251 | 22.481 | -0.745
13.744 | 64.933 | -0.676 | 0.30008 | 0.2495 | 22.983 | -0.676
13.744 | 65.432 | -0.62 | 0.30009 | 0.249 | 23.482 | -0.62
13.744 | 65.93 | -0.565 | 0.30008 | 0.2515 | 23.98 | -0.565
13.744 | 66.433 | -0.513 | 0.30011 | 0.2495 | 24.483 | -0.513
13.744 | 66.932 | -0.466 | 0.30009 | 0.25 | 24.982 | -0.466

## SW Run 2.8A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002007.stp | ! data file
2020-08-11 | ! date
10:33 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=2.8A
! end File Header |  |  |  |  | I | 2.7994 | A
! begin Step Header |  |  |  |  | BdL | -1895.3366 | g-cm
0 | ! set current |  |  |  | B_center | -330.717 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.731 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.943 | -0.751 | 0.27995 | 0.246 | -25.007 | -0.751
13.744 | 17.435 | -0.658 | 0.27994 | 0.249 | -24.515 | -0.658
13.744 | 17.933 | -0.653 | 0.27995 | 0.251 | -24.017 | -0.653
13.744 | 18.435 | -0.68 | 0.27996 | 0.2495 | -23.515 | -0.68
13.744 | 18.934 | -0.717 | 0.27994 | 0.2475 | -23.016 | -0.717
13.744 | 19.429 | -0.761 | 0.27995 | 0.2515 | -22.521 | -0.761
13.744 | 19.932 | -0.815 | 0.27996 | 0.25 | -22.018 | -0.815
13.744 | 20.432 | -0.873 | 0.27994 | 0.2495 | -21.518 | -0.873
13.744 | 20.931 | -0.943 | 0.27995 | 0.2515 | -21.019 | -0.943
13.744 | 21.434 | -1.015 | 0.27995 | 0.2485 | -20.516 | -1.015
13.744 | 21.931 | -1.1 | 0.27995 | 0.251 | -20.019 | -1.1
13.744 | 22.433 | -1.191 | 0.27995 | 0.2505 | -19.517 | -1.191
13.744 | 22.934 | -1.298 | 0.27994 | 0.2485 | -19.016 | -1.298
13.744 | 23.431 | -1.418 | 0.27994 | 0.2515 | -18.519 | -1.418
13.744 | 23.934 | -1.555 | 0.27995 | 0.249 | -18.016 | -1.555
13.744 | 24.432 | -1.714 | 0.27995 | 0.2485 | -17.518 | -1.714
13.744 | 24.929 | -1.892 | 0.27994 | 0.2525 | -17.021 | -1.892
13.744 | 25.434 | -2.095 | 0.27995 | 0.249 | -16.516 | -2.095
13.744 | 25.932 | -2.344 | 0.27993 | 0.2505 | -16.018 | -2.344
13.744 | 26.433 | -2.636 | 0.27993 | 0.25 | -15.517 | -2.636
13.744 | 26.933 | -2.988 | 0.27995 | 0.248 | -15.017 | -2.988
13.744 | 27.429 | -3.387 | 0.27994 | 0.252 | -14.521 | -3.387
13.744 | 27.933 | -3.861 | 0.27994 | 0.2495 | -14.017 | -3.861
13.744 | 28.432 | -4.455 | 0.27994 | 0.2495 | -13.518 | -4.455
13.744 | 28.931 | -5.155 | 0.27993 | 0.2505 | -13.019 | -5.155
13.744 | 29.432 | -6.535 | 0.27994 | 0.2485 | -12.518 | -6.535
13.744 | 29.929 | -7.257 | 0.27994 | 0.2515 | -12.021 | -7.257
13.744 | 30.432 | -9.006 | 0.27994 | 0.2505 | -11.518 | -9.006
13.744 | 30.933 | -10.857 | 0.27993 | 0.249 | -11.017 | -10.857
13.744 | 31.431 | -13.603 | 0.27994 | 0.251 | -10.519 | -13.603
13.744 | 31.933 | -16.594 | 0.27995 | 0.0985 | -10.017 | -16.594
13.744 | 32.13 | -18.262 | 0.27995 | 0.0995 | -9.82 | -18.262
13.744 | 32.329 | -19.861 | 0.27994 | 0.101 | -9.621 | -19.861
13.744 | 32.531 | -22.313 | 0.27994 | 0.1005 | -9.419 | -22.313
13.744 | 32.732 | -24.056 | 0.27994 | 0.099 | -9.218 | -24.056
13.744 | 32.93 | -26.822 | 0.27995 | 0.1005 | -9.02 | -26.822
13.744 | 33.131 | -29.336 | 0.27994 | 0.101 | -8.819 | -29.336
13.744 | 33.333 | -32.501 | 0.27993 | 0.1005 | -8.617 | -32.501
13.744 | 33.534 | -36.23 | 0.27996 | 0.099 | -8.416 | -36.23
13.744 | 33.732 | -39.686 | 0.27994 | 0.0995 | -8.218 | -39.686
13.744 | 33.931 | -43.893 | 0.27993 | 0.101 | -8.019 | -43.893
13.744 | 34.133 | -49.093 | 0.27994 | 0.1 | -7.817 | -49.093
13.744 | 34.333 | -54.549 | 0.27994 | 0.0985 | -7.617 | -54.549
13.744 | 34.53 | -60.441 | 0.27995 | 0.0995 | -7.42 | -60.441
13.744 | 34.729 | -67.108 | 0.27994 | 0.1005 | -7.221 | -67.108
13.744 | 34.93 | -74.49 | 0.27996 | 0.1005 | -7.02 | -74.49
13.744 | 35.131 | -82.357 | 0.27995 | 0.099 | -6.819 | -82.357
13.744 | 35.329 | -91.055 | 0.27995 | 0.1005 | -6.621 | -91.055
13.744 | 35.53 | -100.568 | 0.27994 | 0.1015 | -6.42 | -100.568
13.744 | 35.733 | -111.105 | 0.27994 | 0.1005 | -6.217 | -111.105
13.744 | 35.934 | -122.372 | 0.27994 | 0.099 | -6.016 | -122.372
13.744 | 36.132 | -133.823 | 0.27996 | 0.1 | -5.818 | -133.823
13.744 | 36.332 | -145.966 | 0.27994 | 0.1005 | -5.618 | -145.966
13.744 | 36.533 | -159.193 | 0.27993 | 0.1005 | -5.417 | -159.193
13.744 | 36.734 | -172.108 | 0.27994 | 0.098 | -5.216 | -172.108
13.744 | 36.93 | -184.904 | 0.27994 | 0.1 | -5.02 | -184.904
13.744 | 37.13 | -198.385 | 0.27995 | 0.1005 | -4.82 | -198.385
13.744 | 37.331 | -211.65 | 0.27995 | 0.0995 | -4.619 | -211.65
13.744 | 37.53 | -224.203 | 0.27996 | 0.0985 | -4.42 | -224.203
13.744 | 37.727 | -236.6 | 0.27996 | 0.1005 | -4.223 | -236.6
13.744 | 37.928 | -248.364 | 0.27994 | 0.1015 | -4.022 | -248.364
13.744 | 38.131 | -259.157 | 0.27994 | 0.101 | -3.819 | -259.157
13.744 | 38.333 | -269.419 | 0.27993 | 0.099 | -3.617 | -269.419
13.744 | 38.531 | -278.913 | 0.27993 | 0.1 | -3.419 | -278.913
13.744 | 38.731 | -287.419 | 0.27993 | 0.1005 | -3.219 | -287.419
13.744 | 38.932 | -294.942 | 0.27994 | 0.1005 | -3.018 | -294.942
13.744 | 39.133 | -302.042 | 0.27995 | 0.0985 | -2.817 | -302.042
13.744 | 39.33 | -307.678 | 0.27994 | 0.0995 | -2.62 | -307.678
13.744 | 39.529 | -312.31 | 0.27994 | 0.1005 | -2.421 | -312.31
13.744 | 39.73 | -316.584 | 0.27994 | 0.1 | -2.22 | -316.584
13.744 | 39.93 | -320.341 | 0.27993 | 0.0985 | -2.02 | -320.341
13.744 | 40.127 | -323.166 | 0.27993 | 0.1005 | -1.823 | -323.166
13.744 | 40.328 | -325.602 | 0.27994 | 0.1015 | -1.622 | -325.602
13.744 | 40.531 | -327.065 | 0.27994 | 0.101 | -1.419 | -327.065
13.744 | 40.733 | -328.436 | 0.27994 | 0.099 | -1.217 | -328.436
13.744 | 40.931 | -329.051 | 0.27994 | 0.1 | -1.019 | -329.051
13.744 | 41.131 | -329.689 | 0.27995 | 0.101 | -0.819 | -329.689
13.744 | 41.333 | -330.165 | 0.27995 | 0.1005 | -0.617 | -330.165
13.744 | 41.534 | -330.485 | 0.27993 | 0.099 | -0.416 | -330.485
13.744 | 41.732 | -330.654 | 0.27993 | 0.099 | -0.218 | -330.654
13.744 | 41.93 | -330.717 | 0.27993 | 0.1005 | -0.02 | -330.717
13.744 | 42.131 | -330.686 | 0.27994 | 0.1 | 0.181 | -330.686
13.744 | 42.331 | -330.573 | 0.27993 | 0.099 | 0.381 | -330.573
13.744 | 42.529 | -330.388 | 0.27994 | 0.0995 | 0.579 | -330.388
13.744 | 42.728 | -330.072 | 0.27993 | 0.1015 | 0.778 | -330.072
13.744 | 42.931 | -329.636 | 0.27994 | 0.101 | 0.981 | -329.636
13.744 | 43.133 | -329.011 | 0.27993 | 0.0995 | 1.183 | -329.011
13.744 | 43.332 | -327.592 | 0.27993 | 0.0995 | 1.382 | -327.592
13.744 | 43.531 | -326.172 | 0.27994 | 0.101 | 1.581 | -326.172
13.744 | 43.733 | -324.466 | 0.27994 | 0.1005 | 1.783 | -324.466
13.744 | 43.934 | -321.859 | 0.27995 | 0.099 | 1.984 | -321.859
13.744 | 44.132 | -318.994 | 0.27994 | 0.099 | 2.182 | -318.994
13.744 | 44.33 | -314.862 | 0.27994 | 0.101 | 2.38 | -314.862
13.744 | 44.532 | -310.636 | 0.27998 | 0.1 | 2.582 | -310.636
13.744 | 44.732 | -305.386 | 0.27995 | 0.0985 | 2.782 | -305.386
13.744 | 44.929 | -299.651 | 0.27995 | 0.0995 | 2.979 | -299.651
13.744 | 45.128 | -292.719 | 0.27993 | 0.101 | 3.178 | -292.719
13.744 | 45.33 | -284.637 | 0.27995 | 0.101 | 3.38 | -284.637
13.744 | 45.532 | -276.149 | 0.27993 | 0.0995 | 3.582 | -276.149
13.744 | 45.731 | -266.807 | 0.27994 | 0.1 | 3.781 | -266.807
13.744 | 45.931 | -256.128 | 0.27994 | 0.101 | 3.981 | -256.128
13.744 | 46.133 | -244.825 | 0.27993 | 0.1005 | 4.183 | -244.825
13.744 | 46.334 | -233.421 | 0.27993 | 0.0995 | 4.384 | -233.421
13.744 | 46.533 | -220.826 | 0.27995 | 0.099 | 4.583 | -220.826
13.744 | 46.731 | -208.212 | 0.27994 | 0.1005 | 4.781 | -208.212
13.744 | 46.932 | -194.933 | 0.27995 | 0.1005 | 4.982 | -194.933
13.744 | 47.133 | -181.88 | 0.27995 | 0.0985 | 5.183 | -181.88
13.744 | 47.33 | -168.755 | 0.27994 | 0.099 | 5.38 | -168.755
13.744 | 47.528 | -155.75 | 0.27994 | 0.101 | 5.578 | -155.75
13.744 | 47.73 | -143.183 | 0.27994 | 0.101 | 5.78 | -143.183
13.744 | 47.932 | -131.064 | 0.27994 | 0.1 | 5.982 | -131.064
13.744 | 48.132 | -119.604 | 0.27996 | 0.099 | 6.182 | -119.604
13.744 | 48.33 | -108.802 | 0.27994 | 0.1015 | 6.38 | -108.802
13.744 | 48.533 | -98.365 | 0.27993 | 0.1005 | 6.583 | -98.365
13.744 | 48.734 | -89.057 | 0.27994 | 0.0995 | 6.784 | -89.057
13.744 | 48.933 | -80.465 | 0.27993 | 0.099 | 6.983 | -80.465
13.744 | 49.131 | -72.63 | 0.27996 | 0.101 | 7.181 | -72.63
13.744 | 49.333 | -65.593 | 0.27995 | 0.1 | 7.383 | -65.593
13.744 | 49.533 | -58.857 | 0.27994 | 0.099 | 7.583 | -58.857
13.744 | 49.731 | -53.106 | 0.27994 | 0.099 | 7.781 | -53.106
13.744 | 49.929 | -47.731 | 0.27994 | 0.1005 | 7.979 | -47.731
13.744 | 50.13 | -42.959 | 0.27995 | 0.101 | 8.18 | -42.959
13.744 | 50.332 | -38.952 | 0.27994 | 0.0995 | 8.382 | -38.952
13.744 | 50.531 | -35.063 | 0.27994 | 0.0995 | 8.581 | -35.063
13.744 | 50.73 | -32.137 | 0.27995 | 0.101 | 8.78 | -32.137
13.744 | 50.932 | -29.077 | 0.27994 | 0.1005 | 8.982 | -29.077
13.744 | 51.133 | -26.347 | 0.27996 | 0.1 | 9.183 | -26.347
13.744 | 51.333 | -23.797 | 0.27993 | 0.099 | 9.383 | -23.797
13.744 | 51.531 | -22.144 | 0.27993 | 0.1005 | 9.581 | -22.144
13.744 | 51.732 | -19.692 | 0.27994 | 0.101 | 9.782 | -19.692
13.744 | 51.934 | -18.198 | 0.27994 | 0.2485 | 9.984 | -18.198
13.744 | 52.431 | -14.523 | 0.27994 | 0.25 | 10.481 | -14.523
13.744 | 52.931 | -11.941 | 0.27995 | 0.251 | 10.981 | -11.941
13.744 | 53.433 | -10.116 | 0.27994 | 0.249 | 11.483 | -10.116
13.744 | 53.931 | -8.365 | 0.27993 | 0.2515 | 11.981 | -8.365
13.744 | 54.434 | -6.908 | 0.27993 | 0.2485 | 12.484 | -6.908
13.744 | 54.931 | -6.232 | 0.27994 | 0.2495 | 12.981 | -6.232
13.744 | 55.43 | -5.397 | 0.27994 | 0.252 | 13.48 | -5.397
13.744 | 55.934 | -4.636 | 0.27994 | 0.249 | 13.984 | -4.636
13.744 | 56.432 | -4.003 | 0.27994 | 0.2505 | 14.482 | -4.003
13.744 | 56.933 | -3.469 | 0.27995 | 0.2495 | 14.983 | -3.469
13.744 | 57.432 | -3.04 | 0.27994 | 0.249 | 15.482 | -3.04
13.744 | 57.93 | -2.656 | 0.27995 | 0.252 | 15.98 | -2.656
13.744 | 58.434 | -2.344 | 0.27994 | 0.2495 | 16.484 | -2.344
13.744 | 58.933 | -2.062 | 0.27994 | 0.2495 | 16.983 | -2.062
13.744 | 59.432 | -1.838 | 0.27994 | 0.2505 | 17.482 | -1.838
13.744 | 59.933 | -1.635 | 0.27995 | 0.2485 | 17.983 | -1.635
13.744 | 60.43 | -1.463 | 0.27993 | 0.2515 | 18.48 | -1.463
13.744 | 60.933 | -1.308 | 0.27993 | 0.2505 | 18.983 | -1.308
13.744 | 61.434 | -1.169 | 0.27993 | 0.2485 | 19.484 | -1.169
13.744 | 61.931 | -1.062 | 0.27994 | 0.2515 | 19.981 | -1.062
13.744 | 62.434 | -0.955 | 0.27993 | 0.248 | 20.484 | -0.955
13.744 | 62.93 | -0.858 | 0.27994 | 0.25 | 20.98 | -0.858
13.744 | 63.43 | -0.777 | 0.27993 | 0.2515 | 21.48 | -0.777
13.744 | 63.933 | -0.702 | 0.27995 | 0.249 | 21.983 | -0.702
13.744 | 64.431 | -0.64 | 0.27993 | 0.2505 | 22.481 | -0.64
13.744 | 64.932 | -0.583 | 0.27993 | 0.2495 | 22.982 | -0.583
13.744 | 65.431 | -0.528 | 0.27995 | 0.249 | 23.481 | -0.528
13.744 | 65.929 | -0.478 | 0.27994 | 0.252 | 23.979 | -0.478
13.744 | 66.433 | -0.431 | 0.27996 | 0.2495 | 24.483 | -0.431
13.744 | 66.932 | -0.386 | 0.27994 | 0.25 | 24.982 | -0.386

## SW Run 2A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002008.stp | ! data file
2020-08-11 | ! date
10:53 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=2A
! end File Header |  |  |  |  | I | 2.0005 | A
! begin Step Header |  |  |  |  | BdL | -1355.755 | g-cm
0 | ! set current |  |  |  | B_center | -238.096 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.6942 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.942 | -0.444 | 0.20004 | 0.2465 | -25.008 | -0.444
13.744 | 17.435 | -0.377 | 0.20004 | 0.2485 | -24.515 | -0.377
13.744 | 17.932 | -0.367 | 0.20003 | 0.2505 | -24.018 | -0.367
13.744 | 18.433 | -0.381 | 0.20004 | 0.25 | -23.517 | -0.381
13.744 | 18.933 | -0.407 | 0.20005 | 0.248 | -23.017 | -0.407
13.744 | 19.429 | -0.429 | 0.20005 | 0.251 | -22.521 | -0.429
13.744 | 19.931 | -0.457 | 0.20006 | 0.2505 | -22.019 | -0.457
13.744 | 20.432 | -0.489 | 0.20005 | 0.2495 | -21.518 | -0.489
13.744 | 20.931 | -0.525 | 0.20007 | 0.251 | -21.019 | -0.525
13.744 | 21.433 | -0.571 | 0.20004 | 0.2485 | -20.517 | -0.571
13.744 | 21.93 | -0.618 | 0.20005 | 0.25 | -20.02 | -0.618
13.744 | 22.43 | -0.676 | 0.20005 | 0.251 | -19.52 | -0.676
13.744 | 22.932 | -0.736 | 0.20004 | 0.2495 | -19.018 | -0.736
13.744 | 23.431 | -0.815 | 0.20004 | 0.251 | -18.519 | -0.815
13.744 | 23.933 | -0.894 | 0.20005 | 0.249 | -18.017 | -0.894
13.744 | 24.431 | -1.006 | 0.20004 | 0.249 | -17.519 | -1.006
13.744 | 24.929 | -1.11 | 0.20004 | 0.252 | -17.021 | -1.11
13.744 | 25.433 | -1.25 | 0.20004 | 0.2495 | -16.517 | -1.25
13.744 | 25.932 | -1.405 | 0.20006 | 0.2505 | -16.018 | -1.405
13.744 | 26.433 | -1.591 | 0.20004 | 0.25 | -15.517 | -1.591
13.744 | 26.933 | -1.815 | 0.20004 | 0.248 | -15.017 | -1.815
13.744 | 27.429 | -2.072 | 0.20005 | 0.252 | -14.521 | -2.072
13.744 | 27.933 | -2.41 | 0.20006 | 0.2495 | -14.017 | -2.41
13.744 | 28.432 | -2.798 | 0.20003 | 0.249 | -13.518 | -2.798
13.744 | 28.93 | -3.268 | 0.20004 | 0.251 | -13.02 | -3.268
13.744 | 29.432 | -3.873 | 0.20004 | 0.2485 | -12.518 | -3.873
13.744 | 29.929 | -4.616 | 0.20005 | 0.251 | -12.021 | -4.616
13.744 | 30.431 | -6.113 | 0.20004 | 0.251 | -11.519 | -6.113
13.744 | 30.933 | -7.446 | 0.20005 | 0.2485 | -11.017 | -7.446
13.744 | 31.43 | -8.948 | 0.20005 | 0.251 | -10.52 | -8.948
13.744 | 31.932 | -11.58 | 0.20005 | 0.0985 | -10.018 | -11.58
13.744 | 32.129 | -12.147 | 0.20004 | 0.0995 | -9.821 | -12.147
13.744 | 32.328 | -13.63 | 0.20005 | 0.101 | -9.622 | -13.63
13.744 | 32.53 | -15.042 | 0.20004 | 0.1005 | -9.42 | -15.042
13.744 | 32.731 | -16.577 | 0.20005 | 0.099 | -9.219 | -16.577
13.744 | 32.929 | -18.309 | 0.20005 | 0.1005 | -9.021 | -18.309
13.744 | 33.13 | -20.165 | 0.20005 | 0.1015 | -8.82 | -20.165
13.744 | 33.333 | -22.704 | 0.20006 | 0.1005 | -8.617 | -22.704
13.744 | 33.534 | -25.154 | 0.20005 | 0.0985 | -8.416 | -25.154
13.744 | 33.731 | -27.914 | 0.20007 | 0.1 | -8.219 | -27.914
13.744 | 33.931 | -30.896 | 0.20006 | 0.1005 | -8.019 | -30.896
13.744 | 34.132 | -34.178 | 0.20005 | 0.1005 | -7.818 | -34.178
13.744 | 34.333 | -38.133 | 0.20005 | 0.098 | -7.617 | -38.133
13.744 | 34.529 | -42.432 | 0.20005 | 0.1 | -7.421 | -42.432
13.744 | 34.729 | -47.535 | 0.20006 | 0.1005 | -7.221 | -47.535
13.744 | 34.93 | -52.477 | 0.20004 | 0.1005 | -7.02 | -52.477
13.744 | 35.131 | -58.65 | 0.20005 | 0.099 | -6.819 | -58.65
13.744 | 35.329 | -64.992 | 0.20005 | 0.1005 | -6.621 | -64.992
13.744 | 35.53 | -71.764 | 0.20004 | 0.101 | -6.42 | -71.764
13.744 | 35.732 | -79.261 | 0.20005 | 0.101 | -6.218 | -79.261
13.744 | 35.934 | -87.353 | 0.20003 | 0.099 | -6.016 | -87.353
13.744 | 36.132 | -95.952 | 0.20006 | 0.0995 | -5.818 | -95.952
13.744 | 36.331 | -104.662 | 0.20007 | 0.101 | -5.619 | -104.662
13.744 | 36.533 | -114.117 | 0.20004 | 0.1 | -5.417 | -114.117
13.744 | 36.733 | -123.272 | 0.20004 | 0.0985 | -5.217 | -123.272
13.744 | 36.93 | -132.716 | 0.20004 | 0.0995 | -5.02 | -132.716
13.744 | 37.129 | -142.737 | 0.20004 | 0.1005 | -4.821 | -142.737
13.744 | 37.33 | -151.848 | 0.20004 | 0.0995 | -4.62 | -151.848
13.744 | 37.529 | -161.011 | 0.20004 | 0.099 | -4.421 | -161.011
13.744 | 37.727 | -170.08 | 0.20005 | 0.1005 | -4.223 | -170.08
13.744 | 37.928 | -178.517 | 0.20006 | 0.1015 | -4.022 | -178.517
13.744 | 38.131 | -186.916 | 0.20004 | 0.1005 | -3.819 | -186.916
13.744 | 38.332 | -193.885 | 0.20004 | 0.099 | -3.618 | -193.885
13.744 | 38.53 | -201.156 | 0.20005 | 0.1 | -3.42 | -201.156
13.744 | 38.73 | -207.049 | 0.20004 | 0.101 | -3.22 | -207.049
13.744 | 38.932 | -212.702 | 0.20004 | 0.1005 | -3.018 | -212.702
13.744 | 39.133 | -217.243 | 0.20005 | 0.098 | -2.817 | -217.243
13.744 | 39.329 | -221.298 | 0.20004 | 0.1 | -2.621 | -221.298
13.744 | 39.529 | -225.161 | 0.20004 | 0.1005 | -2.421 | -225.161
13.744 | 39.73 | -228.128 | 0.20004 | 0.1 | -2.22 | -228.128
13.744 | 39.93 | -230.896 | 0.20005 | 0.0985 | -2.02 | -230.896
13.744 | 40.127 | -232.522 | 0.20005 | 0.1005 | -1.823 | -232.522
13.744 | 40.328 | -234.214 | 0.20004 | 0.101 | -1.622 | -234.214
13.744 | 40.53 | -235.632 | 0.20004 | 0.1015 | -1.42 | -235.632
13.744 | 40.733 | -236.262 | 0.20004 | 0.099 | -1.217 | -236.262
13.744 | 40.931 | -236.914 | 0.20004 | 0.1 | -1.019 | -236.914
13.744 | 41.131 | -237.412 | 0.20005 | 0.1005 | -0.819 | -237.412
13.744 | 41.332 | -237.789 | 0.20005 | 0.1005 | -0.618 | -237.789
13.744 | 41.533 | -237.992 | 0.20004 | 0.099 | -0.417 | -237.992
13.744 | 41.731 | -238.085 | 0.20003 | 0.099 | -0.219 | -238.085
13.744 | 41.929 | -238.096 | 0.20005 | 0.1005 | -0.021 | -238.096
13.744 | 42.13 | -238.047 | 0.20004 | 0.1005 | 0.18 | -238.047
13.744 | 42.331 | -237.929 | 0.20003 | 0.0985 | 0.381 | -237.929
13.744 | 42.528 | -237.752 | 0.20004 | 0.1 | 0.578 | -237.752
13.744 | 42.728 | -237.516 | 0.20005 | 0.101 | 0.778 | -237.516
13.744 | 42.93 | -237.173 | 0.20006 | 0.101 | 0.98 | -237.173
13.744 | 43.132 | -236.684 | 0.20005 | 0.0995 | 1.182 | -236.684
13.744 | 43.331 | -236.028 | 0.20004 | 0.0995 | 1.381 | -236.028
13.744 | 43.53 | -234.57 | 0.20005 | 0.101 | 1.58 | -234.57
13.744 | 43.732 | -233.179 | 0.20006 | 0.101 | 1.782 | -233.179
13.744 | 43.934 | -231.627 | 0.20005 | 0.099 | 1.984 | -231.627
13.744 | 44.132 | -229.044 | 0.20005 | 0.099 | 2.182 | -229.044
13.744 | 44.33 | -226.495 | 0.20004 | 0.1005 | 2.38 | -226.495
13.744 | 44.531 | -223.478 | 0.20004 | 0.1005 | 2.581 | -223.478
13.744 | 44.732 | -219.974 | 0.20004 | 0.0985 | 2.782 | -219.974
13.744 | 44.929 | -215.606 | 0.20004 | 0.0995 | 2.979 | -215.606
13.744 | 45.128 | -210.574 | 0.20004 | 0.101 | 3.178 | -210.574
13.744 | 45.33 | -204.999 | 0.20004 | 0.101 | 3.38 | -204.999
13.744 | 45.532 | -198.218 | 0.20003 | 0.0995 | 3.582 | -198.218
13.744 | 45.731 | -191.627 | 0.20003 | 0.0995 | 3.781 | -191.627
13.744 | 45.93 | -184.013 | 0.20003 | 0.1015 | 3.98 | -184.013
13.744 | 46.133 | -175.789 | 0.20004 | 0.1005 | 4.183 | -175.789
13.744 | 46.334 | -167.613 | 0.20003 | 0.099 | 4.384 | -167.613
13.744 | 46.532 | -158.492 | 0.20004 | 0.0995 | 4.582 | -158.492
13.744 | 46.731 | -149.48 | 0.20004 | 0.1005 | 4.781 | -149.48
13.744 | 46.932 | -139.783 | 0.20004 | 0.1 | 4.982 | -139.783
13.744 | 47.132 | -130.694 | 0.20006 | 0.099 | 5.182 | -130.694
13.744 | 47.33 | -121.09 | 0.20004 | 0.099 | 5.38 | -121.09
13.744 | 47.528 | -111.482 | 0.20005 | 0.101 | 5.578 | -111.482
13.744 | 47.73 | -102.594 | 0.20004 | 0.1005 | 5.78 | -102.594
13.744 | 47.931 | -93.827 | 0.20004 | 0.1 | 5.981 | -93.827
13.744 | 48.131 | -85.254 | 0.20003 | 0.0995 | 6.181 | -85.254
13.744 | 48.33 | -77.58 | 0.20004 | 0.101 | 6.38 | -77.58
13.744 | 48.532 | -70.421 | 0.20005 | 0.1005 | 6.582 | -70.421
13.744 | 48.733 | -63.624 | 0.20003 | 0.1 | 6.783 | -63.624
13.744 | 48.933 | -57.354 | 0.20003 | 0.099 | 6.983 | -57.354
13.744 | 49.131 | -51.686 | 0.20004 | 0.1005 | 7.181 | -51.686
13.744 | 49.332 | -46.438 | 0.20004 | 0.1005 | 7.382 | -46.438
13.744 | 49.533 | -41.382 | 0.20005 | 0.099 | 7.583 | -41.382
13.744 | 49.731 | -37.384 | 0.20005 | 0.099 | 7.781 | -37.384
13.744 | 49.929 | -33.96 | 0.20003 | 0.1005 | 7.979 | -33.96
13.744 | 50.13 | -30.191 | 0.20006 | 0.1005 | 8.18 | -30.191
13.744 | 50.331 | -27.366 | 0.20004 | 0.0995 | 8.381 | -27.366
13.744 | 50.53 | -24.768 | 0.20004 | 0.0995 | 8.58 | -24.768
13.744 | 50.729 | -22.263 | 0.20004 | 0.101 | 8.779 | -22.263
13.744 | 50.931 | -20.471 | 0.20005 | 0.101 | 8.981 | -20.471
13.744 | 51.133 | -18.021 | 0.20005 | 0.0995 | 9.183 | -18.021
13.744 | 51.332 | -16.52 | 0.20004 | 0.0995 | 9.382 | -16.52
13.744 | 51.531 | -15.057 | 0.20004 | 0.1005 | 9.581 | -15.057
13.744 | 51.732 | -13.697 | 0.20004 | 0.1005 | 9.782 | -13.697
13.744 | 51.933 | -12.417 | 0.20005 | 0.2485 | 9.983 | -12.417
13.744 | 52.43 | -9.902 | 0.20004 | 0.25 | 10.48 | -9.902
13.744 | 52.93 | -8.211 | 0.20005 | 0.251 | 10.98 | -8.211
13.744 | 53.432 | -6.674 | 0.20003 | 0.2495 | 11.482 | -6.674
13.744 | 53.931 | -5.344 | 0.20004 | 0.2515 | 11.981 | -5.344
13.744 | 54.434 | -4.763 | 0.20006 | 0.2485 | 12.484 | -4.763
13.744 | 54.931 | -4.044 | 0.20006 | 0.249 | 12.981 | -4.044
13.744 | 55.429 | -3.437 | 0.20005 | 0.252 | 13.479 | -3.437
13.744 | 55.933 | -2.907 | 0.20004 | 0.2495 | 13.983 | -2.907
13.744 | 56.432 | -2.467 | 0.20004 | 0.2505 | 14.482 | -2.467
13.744 | 56.933 | -2.104 | 0.20004 | 0.2495 | 14.983 | -2.104
13.744 | 57.432 | -1.797 | 0.20006 | 0.2485 | 15.482 | -1.797
13.744 | 57.929 | -1.551 | 0.20005 | 0.252 | 15.979 | -1.551
13.744 | 58.433 | -1.343 | 0.20005 | 0.25 | 16.483 | -1.343
13.744 | 58.933 | -1.165 | 0.20006 | 0.249 | 16.983 | -1.165
13.744 | 59.431 | -1.02 | 0.20006 | 0.251 | 17.481 | -1.02
13.744 | 59.933 | -0.89 | 0.20006 | 0.248 | 17.983 | -0.89
13.744 | 60.429 | -0.777 | 0.20005 | 0.2515 | 18.479 | -0.777
13.744 | 60.932 | -0.686 | 0.20005 | 0.2505 | 18.982 | -0.686
13.744 | 61.433 | -0.601 | 0.20004 | 0.249 | 19.483 | -0.601
13.744 | 61.931 | -0.526 | 0.20004 | 0.251 | 19.981 | -0.526
13.744 | 62.433 | -0.456 | 0.20005 | 0.2485 | 20.483 | -0.456
13.744 | 62.93 | -0.399 | 0.20005 | 0.2495 | 20.98 | -0.399
13.744 | 63.429 | -0.352 | 0.20004 | 0.252 | 21.479 | -0.352
13.744 | 63.933 | -0.306 | 0.20004 | 0.2485 | 21.983 | -0.306
13.744 | 64.43 | -0.27 | 0.20005 | 0.251 | 22.48 | -0.27
13.744 | 64.932 | -0.237 | 0.20005 | 0.2495 | 22.982 | -0.237
13.744 | 65.431 | -0.202 | 0.20004 | 0.249 | 23.481 | -0.202
13.744 | 65.929 | -0.17 | 0.20005 | 0.252 | 23.979 | -0.17
13.744 | 66.433 | -0.143 | 0.20006 | 0.249 | 24.483 | -0.143
13.744 | 66.931 | -0.123 | 0.20005 | 0.25 | 24.981 | -0.123

## SW Run 1A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002009.stp | ! data file
2020-08-11 | ! date
11:09 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=1A
! end File Header |  |  |  |  | I | 0.99884 | A
! begin Step Header |  |  |  |  | BdL | -679.2538 | g-cm
0 | ! set current |  |  |  | B_center | -121.137 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.6073 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.942 | -0.117 | 0.099884 | 0.2465 | -25.008 | -0.117
13.744 | 17.435 | -0.088 | 0.099894 | 0.2485 | -24.515 | -0.088
13.744 | 17.932 | -0.082 | 0.099884 | 0.2505 | -24.018 | -0.082
13.744 | 18.433 | -0.079 | 0.099882 | 0.2495 | -23.517 | -0.079
13.744 | 18.932 | -0.077 | 0.099874 | 0.2485 | -23.018 | -0.077
13.744 | 19.429 | -0.082 | 0.099884 | 0.251 | -22.521 | -0.082
13.744 | 19.931 | -0.092 | 0.099873 | 0.25 | -22.019 | -0.092
13.744 | 20.431 | -0.096 | 0.099875 | 0.2495 | -21.519 | -0.096
13.744 | 20.93 | -0.103 | 0.099872 | 0.2515 | -21.02 | -0.103
13.744 | 21.433 | -0.111 | 0.099873 | 0.2485 | -20.517 | -0.111
13.744 | 21.93 | -0.118 | 0.09988 | 0.2495 | -20.02 | -0.118
13.744 | 22.429 | -0.132 | 0.099884 | 0.2515 | -19.521 | -0.132
13.744 | 22.932 | -0.144 | 0.099897 | 0.249 | -19.018 | -0.144
13.744 | 23.43 | -0.163 | 0.099877 | 0.2515 | -18.52 | -0.163
13.744 | 23.933 | -0.19 | 0.099867 | 0.249 | -18.017 | -0.19
13.744 | 24.431 | -0.225 | 0.099882 | 0.2485 | -17.519 | -0.225
13.744 | 24.928 | -0.263 | 0.099878 | 0.2525 | -17.022 | -0.263
13.744 | 25.433 | -0.308 | 0.099886 | 0.249 | -16.517 | -0.308
13.744 | 25.931 | -0.361 | 0.099877 | 0.2505 | -16.019 | -0.361
13.744 | 26.432 | -0.433 | 0.099874 | 0.25 | -15.518 | -0.433
13.744 | 26.932 | -0.514 | 0.099881 | 0.248 | -15.018 | -0.514
13.744 | 27.428 | -0.627 | 0.099882 | 0.252 | -14.522 | -0.627
13.744 | 27.932 | -0.761 | 0.099886 | 0.25 | -14.018 | -0.761
13.744 | 28.432 | -0.925 | 0.09989 | 0.249 | -13.518 | -0.925
13.744 | 28.93 | -1.128 | 0.099879 | 0.251 | -13.02 | -1.128
13.744 | 29.432 | -1.4 | 0.099887 | 0.248 | -12.518 | -1.4
13.744 | 29.928 | -1.735 | 0.09988 | 0.2515 | -12.022 | -1.735
13.744 | 30.431 | -2.168 | 0.099892 | 0.2505 | -11.519 | -2.168
13.744 | 30.932 | -2.75 | 0.099891 | 0.249 | -11.018 | -2.75
13.744 | 31.43 | -3.484 | 0.099885 | 0.251 | -10.52 | -3.484
13.744 | 31.932 | -4.958 | 0.099876 | 0.0985 | -10.018 | -4.958
13.744 | 32.129 | -5.327 | 0.099881 | 0.0995 | -9.821 | -5.327
13.744 | 32.328 | -5.929 | 0.09988 | 0.101 | -9.622 | -5.929
13.744 | 32.53 | -6.497 | 0.099881 | 0.1005 | -9.42 | -6.497
13.744 | 32.731 | -7.711 | 0.099877 | 0.099 | -9.219 | -7.711
13.744 | 32.929 | -8.268 | 0.099875 | 0.1005 | -9.021 | -8.268
13.744 | 33.13 | -9.53 | 0.099888 | 0.1015 | -8.82 | -9.53
13.744 | 33.333 | -10.748 | 0.099879 | 0.1005 | -8.617 | -10.748
13.744 | 33.534 | -11.974 | 0.09989 | 0.0985 | -8.416 | -11.974
13.744 | 33.731 | -13.24 | 0.099875 | 0.1 | -8.219 | -13.24
13.744 | 33.931 | -14.828 | 0.099899 | 0.1005 | -8.019 | -14.828
13.744 | 34.132 | -16.439 | 0.099884 | 0.1 | -7.818 | -16.439
13.744 | 34.332 | -18.384 | 0.09988 | 0.0985 | -7.618 | -18.384
13.744 | 34.529 | -20.456 | 0.099883 | 0.0995 | -7.421 | -20.456
13.744 | 34.728 | -23.16 | 0.099884 | 0.101 | -7.222 | -23.16
13.744 | 34.93 | -25.838 | 0.099886 | 0.1005 | -7.02 | -25.838
13.744 | 35.131 | -28.684 | 0.09988 | 0.099 | -6.819 | -28.684
13.744 | 35.329 | -32.431 | 0.099883 | 0.1 | -6.621 | -32.431
13.744 | 35.529 | -35.647 | 0.099873 | 0.1015 | -6.421 | -35.647
13.744 | 35.732 | -39.635 | 0.099894 | 0.1005 | -6.218 | -39.635
13.744 | 35.933 | -43.888 | 0.099899 | 0.099 | -6.017 | -43.888
13.744 | 36.131 | -48.046 | 0.099881 | 0.1 | -5.819 | -48.046
13.744 | 36.331 | -52.994 | 0.099884 | 0.1005 | -5.619 | -52.994
13.744 | 36.532 | -57.45 | 0.099885 | 0.1005 | -5.418 | -57.45
13.744 | 36.733 | -62.23 | 0.099882 | 0.0985 | -5.217 | -62.23
13.744 | 36.93 | -67.147 | 0.099883 | 0.0995 | -5.02 | -67.147
13.744 | 37.129 | -72.515 | 0.099888 | 0.1005 | -4.821 | -72.515
13.744 | 37.33 | -77.442 | 0.099895 | 0.0995 | -4.62 | -77.442
13.744 | 37.529 | -82.005 | 0.099881 | 0.099 | -4.421 | -82.005
13.744 | 37.727 | -86.357 | 0.099898 | 0.1 | -4.223 | -86.357
13.744 | 37.927 | -90.94 | 0.099881 | 0.1015 | -4.023 | -90.94
13.744 | 38.13 | -95.085 | 0.099899 | 0.101 | -3.82 | -95.085
13.744 | 38.332 | -98.689 | 0.099884 | 0.099 | -3.618 | -98.689
13.744 | 38.53 | -102.492 | 0.099912 | 0.1 | -3.42 | -102.492
13.744 | 38.73 | -105.293 | 0.099888 | 0.101 | -3.22 | -105.293
13.744 | 38.932 | -108.256 | 0.099878 | 0.1005 | -3.018 | -108.256
13.744 | 39.133 | -110.857 | 0.099878 | 0.098 | -2.817 | -110.857
13.744 | 39.329 | -112.753 | 0.099883 | 0.0995 | -2.621 | -112.753
13.744 | 39.528 | -114.689 | 0.099882 | 0.1005 | -2.422 | -114.689
13.744 | 39.729 | -116.214 | 0.099887 | 0.1 | -2.221 | -116.214
13.744 | 39.929 | -117.629 | 0.099889 | 0.0985 | -2.021 | -117.629
13.744 | 40.126 | -118.889 | 0.099886 | 0.1005 | -1.824 | -118.889
13.744 | 40.327 | -119.305 | 0.099881 | 0.1015 | -1.623 | -119.305
13.744 | 40.53 | -119.944 | 0.099888 | 0.101 | -1.42 | -119.944
13.744 | 40.732 | -120.393 | 0.099909 | 0.099 | -1.218 | -120.393
13.744 | 40.93 | -120.74 | 0.099887 | 0.1 | -1.02 | -120.74
13.744 | 41.13 | -121.018 | 0.099896 | 0.101 | -0.82 | -121.018
13.744 | 41.332 | -121.133 | 0.099879 | 0.1005 | -0.618 | -121.133
13.744 | 41.533 | -121.189 | 0.099875 | 0.099 | -0.417 | -121.189
13.744 | 41.731 | -121.18 | 0.09988 | 0.099 | -0.219 | -121.18
13.744 | 41.929 | -121.137 | 0.099884 | 0.1005 | -0.021 | -121.137
13.744 | 42.13 | -121.034 | 0.099886 | 0.1 | 0.18 | -121.034
13.744 | 42.33 | -120.93 | 0.099876 | 0.099 | 0.38 | -120.93
13.744 | 42.528 | -120.768 | 0.099878 | 0.1 | 0.578 | -120.768
13.744 | 42.728 | -120.604 | 0.099884 | 0.101 | 0.778 | -120.604
13.744 | 42.93 | -120.382 | 0.099879 | 0.101 | 0.98 | -120.382
13.744 | 43.132 | -120.04 | 0.099898 | 0.0995 | 1.182 | -120.04
13.744 | 43.331 | -119.694 | 0.099898 | 0.0995 | 1.381 | -119.694
13.744 | 43.53 | -119.228 | 0.09991 | 0.101 | 1.58 | -119.228
13.744 | 43.732 | -118.486 | 0.099882 | 0.1005 | 1.782 | -118.486
13.744 | 43.933 | -117.167 | 0.099881 | 0.0995 | 1.983 | -117.167
13.744 | 44.132 | -116.438 | 0.099873 | 0.099 | 2.182 | -116.438
13.744 | 44.33 | -114.902 | 0.099897 | 0.1005 | 2.38 | -114.902
13.744 | 44.531 | -113.253 | 0.099882 | 0.1 | 2.581 | -113.253
13.744 | 44.731 | -111.405 | 0.099876 | 0.0985 | 2.781 | -111.405
13.744 | 44.928 | -109.025 | 0.099879 | 0.1 | 2.978 | -109.025
13.744 | 45.128 | -106.466 | 0.099885 | 0.101 | 3.178 | -106.466
13.744 | 45.33 | -103.731 | 0.099875 | 0.101 | 3.38 | -103.731
13.744 | 45.532 | -100.582 | 0.099876 | 0.0995 | 3.582 | -100.582
13.744 | 45.731 | -96.875 | 0.099876 | 0.0995 | 3.781 | -96.875
13.744 | 45.93 | -93.037 | 0.099879 | 0.101 | 3.98 | -93.037
13.744 | 46.132 | -89.039 | 0.09987 | 0.101 | 4.182 | -89.039
13.744 | 46.334 | -84.356 | 0.099893 | 0.099 | 4.384 | -84.356
13.744 | 46.532 | -80.345 | 0.099877 | 0.099 | 4.582 | -80.345
13.744 | 46.73 | -75.184 | 0.099876 | 0.1005 | 4.78 | -75.184
13.744 | 46.931 | -70.293 | 0.099887 | 0.1005 | 4.981 | -70.293
13.744 | 47.132 | -65.919 | 0.099882 | 0.099 | 5.182 | -65.919
13.744 | 47.33 | -60.984 | 0.099887 | 0.099 | 5.38 | -60.984
13.744 | 47.528 | -55.877 | 0.099887 | 0.101 | 5.578 | -55.877
13.744 | 47.73 | -51.637 | 0.099886 | 0.1005 | 5.78 | -51.637
13.744 | 47.931 | -46.749 | 0.099896 | 0.1 | 5.981 | -46.749
13.744 | 48.131 | -42.594 | 0.099874 | 0.0995 | 6.181 | -42.594
13.744 | 48.33 | -38.701 | 0.099886 | 0.101 | 6.38 | -38.701
13.744 | 48.532 | -34.843 | 0.099892 | 0.1005 | 6.582 | -34.843
13.744 | 48.733 | -31.084 | 0.09988 | 0.0995 | 6.783 | -31.084
13.744 | 48.932 | -28.172 | 0.099887 | 0.099 | 6.982 | -28.172
13.744 | 49.13 | -25.41 | 0.099885 | 0.101 | 7.18 | -25.41
13.744 | 49.332 | -22.609 | 0.099896 | 0.1005 | 7.382 | -22.609
13.744 | 49.533 | -20.203 | 0.099891 | 0.099 | 7.583 | -20.203
13.744 | 49.731 | -18.534 | 0.099887 | 0.099 | 7.781 | -18.534
13.744 | 49.929 | -16.075 | 0.099885 | 0.1005 | 7.979 | -16.075
13.744 | 50.13 | -14.656 | 0.099874 | 0.1005 | 8.18 | -14.656
13.744 | 50.331 | -13.17 | 0.099876 | 0.1 | 8.381 | -13.17
13.744 | 50.531 | -11.683 | 0.099879 | 0.099 | 8.581 | -11.683
13.744 | 50.729 | -10.38 | 0.099877 | 0.101 | 8.779 | -10.38
13.744 | 50.931 | -9.781 | 0.099896 | 0.101 | 8.981 | -9.781
13.744 | 51.133 | -8.389 | 0.0999 | 0.0995 | 9.183 | -8.389
13.744 | 51.332 | -7.855 | 0.099901 | 0.099 | 9.382 | -7.855
13.744 | 51.53 | -7.196 | 0.099896 | 0.101 | 9.58 | -7.196
13.744 | 51.732 | -6.513 | 0.099879 | 0.1005 | 9.782 | -6.513
13.744 | 51.933 | -5.879 | 0.099881 | 0.2485 | 9.983 | -5.879
13.744 | 52.43 | -4.357 | 0.09989 | 0.25 | 10.48 | -4.357
13.744 | 52.93 | -3.684 | 0.099874 | 0.251 | 10.98 | -3.684
13.744 | 53.432 | -2.958 | 0.099875 | 0.2495 | 11.482 | -2.958
13.744 | 53.931 | -2.342 | 0.099873 | 0.251 | 11.981 | -2.342
13.744 | 54.433 | -1.846 | 0.099884 | 0.2485 | 12.483 | -1.846
13.744 | 54.93 | -1.475 | 0.099899 | 0.2495 | 12.98 | -1.475
13.744 | 55.429 | -1.154 | 0.099906 | 0.252 | 13.479 | -1.154
13.744 | 55.933 | -0.915 | 0.099885 | 0.249 | 13.983 | -0.915
13.744 | 56.431 | -0.721 | 0.099883 | 0.251 | 14.481 | -0.721
13.744 | 56.933 | -0.563 | 0.099892 | 0.2495 | 14.983 | -0.563
13.744 | 57.432 | -0.437 | 0.099875 | 0.2485 | 15.482 | -0.437
13.744 | 57.929 | -0.338 | 0.099882 | 0.252 | 15.979 | -0.338
13.744 | 58.433 | -0.253 | 0.09988 | 0.2495 | 16.483 | -0.253
13.744 | 58.932 | -0.189 | 0.099874 | 0.2495 | 16.982 | -0.189
13.744 | 59.431 | -0.129 | 0.099874 | 0.251 | 17.481 | -0.129
13.744 | 59.933 | -0.084 | 0.099889 | 0.2485 | 17.983 | -0.084
13.744 | 60.43 | -0.044 | 0.099878 | 0.251 | 18.48 | -0.044
13.744 | 60.932 | -0.01 | 0.099875 | 0.2505 | 18.982 | -0.01
13.744 | 61.433 | 0.013 | 0.099883 | 0.249 | 19.483 | 0.013
13.744 | 61.931 | 0.038 | 0.099894 | 0.251 | 19.981 | 0.038
13.744 | 62.433 | 0.062 | 0.099874 | 0.248 | 20.483 | 0.062
13.744 | 62.929 | 0.075 | 0.099885 | 0.25 | 20.979 | 0.075
13.744 | 63.429 | 0.089 | 0.099883 | 0.2515 | 21.479 | 0.089
13.744 | 63.932 | 0.107 | 0.099876 | 0.249 | 21.982 | 0.107
13.744 | 64.43 | 0.115 | 0.099877 | 0.251 | 22.48 | 0.115
13.744 | 64.932 | 0.125 | 0.099888 | 0.2495 | 22.982 | 0.125
13.744 | 65.431 | 0.137 | 0.099887 | 0.249 | 23.481 | 0.137
13.744 | 65.929 | 0.145 | 0.099881 | 0.2515 | 23.979 | 0.145
13.744 | 66.432 | 0.151 | 0.099883 | 0.2495 | 24.482 | 0.151
13.744 | 66.931 | 0.156 | 0.099885 | 0.25 | 24.981 | 0.156

## SW Run 0A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002010.stp | ! data file
2020-08-11 | ! date
11:30 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=0A (PS OFF)
! end File Header |  |  |  |  | I | -0.0030953 | A
! begin Step Header |  |  |  |  | BdL | -1.4813 | g-cm
0 | ! set current |  |  |  | B_center | -3.508 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 0.42226 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.941 | 0.206 | -0.000309 | 0.2465 | -25.009 | 0.206
13.744 | 17.434 | 0.171 | -0.00031 | 0.2485 | -24.516 | 0.171
13.744 | 17.931 | 0.171 | -0.000309 | 0.2505 | -24.019 | 0.171
13.744 | 18.432 | 0.181 | -0.000313 | 0.25 | -23.518 | 0.181
13.744 | 18.932 | 0.197 | -0.000309 | 0.248 | -23.018 | 0.197
13.744 | 19.428 | 0.219 | -0.000311 | 0.2515 | -22.522 | 0.219
13.744 | 19.931 | 0.23 | -0.000311 | 0.25 | -22.019 | 0.23
13.744 | 20.431 | 0.251 | -0.000311 | 0.2495 | -21.519 | 0.251
13.744 | 20.93 | 0.269 | -0.00031 | 0.251 | -21.02 | 0.269
13.744 | 21.432 | 0.288 | -0.000311 | 0.2485 | -20.518 | 0.288
13.744 | 21.929 | 0.314 | -0.000309 | 0.25 | -20.021 | 0.314
13.744 | 22.429 | 0.342 | -0.000309 | 0.2515 | -19.521 | 0.342
13.744 | 22.932 | 0.368 | -0.000312 | 0.249 | -19.018 | 0.368
13.744 | 23.43 | 0.392 | -0.000311 | 0.251 | -18.52 | 0.392
13.744 | 23.932 | 0.425 | -0.00031 | 0.2495 | -18.018 | 0.425
13.744 | 24.431 | 0.462 | -0.00031 | 0.2485 | -17.519 | 0.462
13.744 | 24.928 | 0.491 | -0.000309 | 0.2525 | -17.022 | 0.491
13.744 | 25.433 | 0.525 | -0.00031 | 0.249 | -16.517 | 0.525
13.744 | 25.931 | 0.562 | -0.000311 | 0.2505 | -16.019 | 0.562
13.744 | 26.432 | 0.605 | -0.000311 | 0.25 | -15.518 | 0.605
13.744 | 26.932 | 0.653 | -0.000307 | 0.248 | -15.018 | 0.653
13.744 | 27.428 | 0.698 | -0.000313 | 0.252 | -14.522 | 0.698
13.744 | 27.932 | 0.747 | -0.00031 | 0.2495 | -14.018 | 0.747
13.744 | 28.431 | 0.8 | -0.000308 | 0.2495 | -13.519 | 0.8
13.744 | 28.93 | 0.856 | -0.00031 | 0.2505 | -13.02 | 0.856
13.744 | 29.431 | 0.904 | -0.000309 | 0.2485 | -12.519 | 0.904
13.744 | 29.928 | 0.962 | -0.000313 | 0.2515 | -12.022 | 0.962
13.744 | 30.431 | 1.015 | -0.000309 | 0.2505 | -11.519 | 1.015
13.744 | 30.932 | 1.072 | -0.000309 | 0.2485 | -11.018 | 1.072
13.744 | 31.429 | 1.12 | -0.000312 | 0.2515 | -10.521 | 1.12
13.744 | 31.932 | 1.166 | -0.00031 | 0.0985 | -10.018 | 1.166
13.744 | 32.129 | 1.187 | -0.00031 | 0.0995 | -9.821 | 1.187
13.744 | 32.328 | 1.191 | -0.00031 | 0.101 | -9.622 | 1.191
13.744 | 32.53 | 1.191 | -0.000308 | 0.1005 | -9.42 | 1.191
13.744 | 32.731 | 1.189 | -0.000312 | 0.099 | -9.219 | 1.189
13.744 | 32.929 | 1.185 | -0.000308 | 0.1005 | -9.021 | 1.185
13.744 | 33.13 | 1.176 | -0.00031 | 0.101 | -8.82 | 1.176
13.744 | 33.332 | 1.156 | -0.00031 | 0.1005 | -8.618 | 1.156
13.744 | 33.533 | 1.129 | -0.00031 | 0.099 | -8.417 | 1.129
13.744 | 33.731 | 1.09 | -0.000312 | 0.0995 | -8.219 | 1.09
13.744 | 33.93 | 1.043 | -0.000311 | 0.101 | -8.02 | 1.043
13.744 | 34.132 | 0.992 | -0.00031 | 0.1 | -7.818 | 0.992
13.744 | 34.332 | 0.907 | -0.000311 | 0.098 | -7.618 | 0.907
13.744 | 34.528 | 0.83 | -0.000311 | 0.1 | -7.422 | 0.83
13.744 | 34.728 | 0.729 | -0.00031 | 0.1005 | -7.222 | 0.729
13.744 | 34.929 | 0.593 | -0.000307 | 0.1005 | -7.021 | 0.593
13.744 | 35.13 | 0.467 | -0.00031 | 0.099 | -6.82 | 0.467
13.744 | 35.328 | 0.29 | -0.000311 | 0.1005 | -6.622 | 0.29
13.744 | 35.529 | 0.132 | -0.000311 | 0.1015 | -6.421 | 0.132
13.744 | 35.732 | -0.051 | -0.000309 | 0.1005 | -6.218 | -0.051
13.744 | 35.933 | -0.293 | -0.000308 | 0.099 | -6.017 | -0.293
13.744 | 36.131 | -0.502 | -0.000308 | 0.0995 | -5.819 | -0.502
13.744 | 36.33 | -0.727 | -0.000309 | 0.101 | -5.62 | -0.727
13.744 | 36.532 | -1.023 | -0.00031 | 0.1 | -5.418 | -1.023
13.744 | 36.732 | -1.267 | -0.000308 | 0.0985 | -5.218 | -1.267
13.744 | 36.929 | -1.521 | -0.000309 | 0.1 | -5.021 | -1.521
13.744 | 37.129 | -1.838 | -0.00031 | 0.1005 | -4.821 | -1.838
13.744 | 37.33 | -2.079 | -0.000309 | 0.0995 | -4.62 | -2.079
13.744 | 37.529 | -2.38 | -0.000309 | 0.099 | -4.421 | -2.38
13.744 | 37.727 | -2.603 | -0.00031 | 0.1 | -4.223 | -2.603
13.744 | 37.927 | -2.825 | -0.000309 | 0.1015 | -4.023 | -2.825
13.744 | 38.13 | -3.08 | -0.000308 | 0.101 | -3.82 | -3.08
13.744 | 38.332 | -3.264 | -0.000311 | 0.099 | -3.618 | -3.264
13.744 | 38.53 | -3.473 | -0.000309 | 0.1 | -3.42 | -3.473
13.744 | 38.73 | -3.612 | -0.000309 | 0.1005 | -3.22 | -3.612
13.744 | 38.931 | -3.744 | -0.000309 | 0.1005 | -3.019 | -3.744
13.744 | 39.132 | -3.881 | -0.00031 | 0.0985 | -2.818 | -3.881
13.744 | 39.329 | -3.961 | -0.000308 | 0.0995 | -2.621 | -3.961
13.744 | 39.528 | -4.027 | -0.000311 | 0.1005 | -2.422 | -4.027
13.744 | 39.729 | -4.083 | -0.000311 | 0.1 | -2.221 | -4.083
13.744 | 39.929 | -4.107 | -0.00031 | 0.0985 | -2.021 | -4.107
13.744 | 40.126 | -4.11 | -0.000311 | 0.1005 | -1.824 | -4.11
13.744 | 40.327 | -4.096 | -0.000308 | 0.1015 | -1.623 | -4.096
13.744 | 40.53 | -4.071 | -0.000309 | 0.101 | -1.42 | -4.071
13.744 | 40.732 | -4.017 | -0.000308 | 0.099 | -1.218 | -4.017
13.744 | 40.93 | -3.965 | -0.000309 | 0.1 | -1.02 | -3.965
13.744 | 41.13 | -3.892 | -0.000309 | 0.101 | -0.82 | -3.892
13.744 | 41.332 | -3.797 | -0.00031 | 0.1005 | -0.618 | -3.797
13.744 | 41.533 | -3.709 | -0.00031 | 0.099 | -0.417 | -3.709
13.744 | 41.731 | -3.604 | -0.000311 | 0.099 | -0.219 | -3.604
13.744 | 41.929 | -3.508 | -0.000308 | 0.1005 | -0.021 | -3.508
13.744 | 42.13 | -3.405 | -0.000309 | 0.1 | 0.18 | -3.405
13.744 | 42.33 | -3.28 | -0.000308 | 0.099 | 0.38 | -3.28
13.744 | 42.528 | -3.181 | -0.000309 | 0.0995 | 0.578 | -3.181
13.744 | 42.727 | -3.078 | -0.000307 | 0.1015 | 0.777 | -3.078
13.744 | 42.93 | -2.964 | -0.000309 | 0.1005 | 0.98 | -2.964
13.744 | 43.131 | -2.879 | -0.000309 | 0.1 | 1.181 | -2.879
13.744 | 43.331 | -2.776 | -0.000311 | 0.0995 | 1.381 | -2.776
13.744 | 43.53 | -2.694 | -0.000309 | 0.101 | 1.58 | -2.694
13.744 | 43.732 | -2.607 | -0.00031 | 0.1005 | 1.782 | -2.607
13.744 | 43.933 | -2.507 | -0.000311 | 0.099 | 1.983 | -2.507
13.744 | 44.131 | -2.434 | -0.00031 | 0.099 | 2.181 | -2.434
13.744 | 44.329 | -2.353 | -0.000309 | 0.101 | 2.379 | -2.353
13.744 | 44.531 | -2.255 | -0.000311 | 0.1 | 2.581 | -2.255
13.744 | 44.731 | -2.18 | -0.00031 | 0.0985 | 2.781 | -2.18
13.744 | 44.928 | -2.078 | -0.000311 | 0.0995 | 2.978 | -2.078
13.744 | 45.127 | -1.992 | -0.000308 | 0.101 | 3.177 | -1.992
13.744 | 45.329 | -1.894 | -0.00031 | 0.101 | 3.379 | -1.894
13.744 | 45.531 | -1.762 | -0.000309 | 0.0995 | 3.581 | -1.762
13.744 | 45.73 | -1.652 | -0.000309 | 0.1 | 3.78 | -1.652
13.744 | 45.93 | -1.526 | -0.000308 | 0.101 | 3.98 | -1.526
13.744 | 46.132 | -1.362 | -0.000309 | 0.1005 | 4.182 | -1.362
13.744 | 46.333 | -1.225 | -0.000309 | 0.099 | 4.383 | -1.225
13.744 | 46.531 | -1.074 | -0.000311 | 0.0995 | 4.581 | -1.074
13.744 | 46.73 | -0.886 | -0.000309 | 0.1005 | 4.78 | -0.886
13.744 | 46.931 | -0.726 | -0.000311 | 0.1005 | 4.981 | -0.726
13.744 | 47.132 | -0.53 | -0.000311 | 0.0985 | 5.182 | -0.53
13.744 | 47.329 | -0.369 | -0.000311 | 0.099 | 5.379 | -0.369
13.744 | 47.527 | -0.206 | -0.000311 | 0.101 | 5.577 | -0.206
13.744 | 47.729 | -0.006 | -0.00031 | 0.101 | 5.779 | -0.006
13.744 | 47.931 | 0.146 | -0.000311 | 0.1 | 5.981 | 0.146
13.744 | 48.131 | 0.293 | -0.000309 | 0.099 | 6.181 | 0.293
13.744 | 48.329 | 0.467 | -0.000309 | 0.101 | 6.379 | 0.467
13.744 | 48.531 | 0.589 | -0.000309 | 0.101 | 6.581 | 0.589
13.744 | 48.733 | 0.731 | -0.000309 | 0.0995 | 6.783 | 0.731
13.744 | 48.932 | 0.83 | -0.00031 | 0.099 | 6.982 | 0.83
13.744 | 49.13 | 0.924 | -0.000309 | 0.1005 | 7.18 | 0.924
13.744 | 49.331 | 1.021 | -0.000309 | 0.1005 | 7.381 | 1.021
13.744 | 49.532 | 1.087 | -0.000309 | 0.099 | 7.582 | 1.087
13.744 | 49.73 | 1.144 | -0.000309 | 0.099 | 7.78 | 1.144
13.744 | 49.928 | 1.206 | -0.000309 | 0.101 | 7.978 | 1.206
13.744 | 50.13 | 1.242 | -0.00031 | 0.1005 | 8.18 | 1.242
13.744 | 50.331 | 1.279 | -0.000311 | 0.0995 | 8.381 | 1.279
13.744 | 50.53 | 1.3 | -0.000309 | 0.0995 | 8.58 | 1.3
13.744 | 50.729 | 1.317 | -0.000309 | 0.101 | 8.779 | 1.317
13.744 | 50.931 | 1.326 | -0.00031 | 0.101 | 8.981 | 1.326
13.744 | 51.133 | 1.331 | -0.000309 | 0.099 | 9.183 | 1.331
13.744 | 51.331 | 1.332 | -0.000309 | 0.0995 | 9.381 | 1.332
13.744 | 51.53 | 1.325 | -0.000309 | 0.1005 | 9.58 | 1.325
13.744 | 51.731 | 1.322 | -0.000312 | 0.1005 | 9.781 | 1.322
13.744 | 51.932 | 1.309 | -0.000308 | 0.2485 | 9.982 | 1.309
13.744 | 52.429 | 1.278 | -0.000309 | 0.2505 | 10.479 | 1.278
13.744 | 52.93 | 1.243 | -0.00031 | 0.251 | 10.98 | 1.243
13.744 | 53.432 | 1.193 | -0.000306 | 0.249 | 11.482 | 1.193
13.744 | 53.93 | 1.14 | -0.000309 | 0.2515 | 11.98 | 1.14
13.744 | 54.433 | 1.088 | -0.000307 | 0.2485 | 12.483 | 1.088
13.744 | 54.93 | 1.034 | -0.000308 | 0.249 | 12.98 | 1.034
13.744 | 55.428 | 0.987 | -0.000307 | 0.2525 | 13.478 | 0.987
13.744 | 55.933 | 0.934 | -0.00031 | 0.249 | 13.983 | 0.934
13.744 | 56.431 | 0.892 | -0.000308 | 0.251 | 14.481 | 0.892
13.744 | 56.933 | 0.845 | -0.000307 | 0.249 | 14.983 | 0.845
13.744 | 57.431 | 0.806 | -0.000309 | 0.2485 | 15.481 | 0.806
13.744 | 57.928 | 0.765 | -0.000309 | 0.2525 | 15.978 | 0.765
13.744 | 58.433 | 0.731 | -0.000309 | 0.2495 | 16.483 | 0.731
13.744 | 58.932 | 0.695 | -0.000309 | 0.2495 | 16.982 | 0.695
13.744 | 59.431 | 0.663 | -0.000307 | 0.2505 | 17.481 | 0.663
13.744 | 59.932 | 0.635 | -0.000309 | 0.2485 | 17.982 | 0.635
13.744 | 60.429 | 0.609 | -0.000311 | 0.2515 | 18.479 | 0.609
13.744 | 60.932 | 0.581 | -0.000308 | 0.2505 | 18.982 | 0.581
13.744 | 61.433 | 0.562 | -0.00031 | 0.2485 | 19.483 | 0.562
13.744 | 61.93 | 0.539 | -0.000309 | 0.251 | 19.98 | 0.539
13.744 | 62.432 | 0.519 | -0.000309 | 0.2485 | 20.482 | 0.519
13.744 | 62.929 | 0.504 | -0.000309 | 0.25 | 20.979 | 0.504
13.744 | 63.429 | 0.487 | -0.000309 | 0.2515 | 21.479 | 0.487
13.744 | 63.932 | 0.468 | -0.000309 | 0.249 | 21.982 | 0.468
13.744 | 64.43 | 0.452 | -0.000311 | 0.2505 | 22.48 | 0.452
13.744 | 64.931 | 0.439 | -0.000308 | 0.2495 | 22.981 | 0.439
13.744 | 65.43 | 0.427 | -0.000311 | 0.249 | 23.48 | 0.427
13.744 | 65.928 | 0.415 | -0.000308 | 0.252 | 23.978 | 0.415
13.744 | 66.432 | 0.406 | -0.000308 | 0.249 | 24.482 | 0.406
13.744 | 66.93 | 0.395 | -0.000309 | 0.25 | 24.98 | 0.395

## SW Run -1A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002011.stp | ! data file
2020-08-11 | ! date
13:14 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=-1A
! end File Header |  |  |  |  | I | -0.99879 | A
! begin Step Header |  |  |  |  | BdL | 672.5707 | g-cm
0 | ! set current |  |  |  | B_center | 113.723 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.9141 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.939 | 0.494 | -0.099886 | 0.247 | -25.011 | 0.494
13.744 | 17.433 | 0.412 | -0.099895 | 0.2485 | -24.517 | 0.412
13.744 | 17.93 | 0.405 | -0.099895 | 0.251 | -24.02 | 0.405
13.744 | 18.432 | 0.421 | -0.099891 | 0.25 | -23.518 | 0.421
13.744 | 18.932 | 0.453 | -0.099896 | 0.2475 | -23.018 | 0.453
13.744 | 19.427 | 0.489 | -0.099876 | 0.2515 | -22.523 | 0.489
13.744 | 19.93 | 0.524 | -0.099888 | 0.2505 | -22.02 | 0.524
13.744 | 20.431 | 0.568 | -0.099888 | 0.249 | -21.519 | 0.568
13.744 | 20.929 | 0.615 | -0.099891 | 0.251 | -21.021 | 0.615
13.744 | 21.431 | 0.667 | -0.099881 | 0.2485 | -20.519 | 0.667
13.744 | 21.928 | 0.718 | -0.099869 | 0.2505 | -20.022 | 0.718
13.744 | 22.429 | 0.78 | -0.099876 | 0.251 | -19.521 | 0.78
13.744 | 22.931 | 0.845 | -0.099888 | 0.249 | -19.019 | 0.845
13.744 | 23.429 | 0.92 | -0.099883 | 0.251 | -18.521 | 0.92
13.744 | 23.931 | 1 | -0.099883 | 0.249 | -18.019 | 1
13.744 | 24.429 | 1.088 | -0.099874 | 0.249 | -17.521 | 1.088
13.744 | 24.927 | 1.194 | -0.099887 | 0.252 | -17.023 | 1.194
13.744 | 25.431 | 1.311 | -0.099878 | 0.249 | -16.519 | 1.311
13.744 | 25.929 | 1.443 | -0.099863 | 0.2505 | -16.021 | 1.443
13.744 | 26.43 | 1.59 | -0.099892 | 0.25 | -15.52 | 1.59
13.744 | 26.93 | 1.757 | -0.099868 | 0.248 | -15.02 | 1.757
13.744 | 27.426 | 1.952 | -0.099899 | 0.2525 | -14.524 | 1.952
13.744 | 27.931 | 2.174 | -0.099887 | 0.2495 | -14.019 | 2.174
13.744 | 28.43 | 2.441 | -0.099876 | 0.249 | -13.52 | 2.441
13.744 | 28.928 | 2.746 | -0.099885 | 0.251 | -13.022 | 2.746
13.744 | 29.43 | 3.127 | -0.09988 | 0.248 | -12.52 | 3.127
13.744 | 29.926 | 3.565 | -0.099882 | 0.2515 | -12.024 | 3.565
13.744 | 30.429 | 4.118 | -0.099899 | 0.251 | -11.521 | 4.118
13.744 | 30.931 | 4.77 | -0.099879 | 0.2485 | -11.019 | 4.77
13.744 | 31.428 | 5.59 | -0.099891 | 0.251 | -10.522 | 5.59
13.744 | 31.93 | 7.14 | -0.099901 | 0.0985 | -10.02 | 7.14
13.744 | 32.127 | 7.605 | -0.09989 | 0.0995 | -9.823 | 7.605
13.744 | 32.326 | 8.087 | -0.099894 | 0.101 | -9.624 | 8.087
13.744 | 32.528 | 8.802 | -0.099887 | 0.1005 | -9.422 | 8.802
13.744 | 32.729 | 9.435 | -0.09989 | 0.099 | -9.221 | 9.435
13.744 | 32.927 | 10.8 | -0.099894 | 0.1005 | -9.023 | 10.8
13.744 | 33.128 | 11.35 | -0.099885 | 0.1015 | -8.822 | 11.35
13.744 | 33.331 | 12.718 | -0.099881 | 0.1005 | -8.619 | 12.718
13.744 | 33.532 | 13.962 | -0.099888 | 0.0985 | -8.418 | 13.962
13.744 | 33.729 | 15.276 | -0.099878 | 0.1 | -8.221 | 15.276
13.744 | 33.929 | 16.632 | -0.09988 | 0.1005 | -8.021 | 16.632
13.744 | 34.13 | 18.166 | -0.099879 | 0.1 | -7.82 | 18.166
13.744 | 34.33 | 19.881 | -0.099881 | 0.0985 | -7.62 | 19.881
13.744 | 34.527 | 22.222 | -0.099889 | 0.0995 | -7.423 | 22.222
13.744 | 34.726 | 23.878 | -0.09988 | 0.1005 | -7.224 | 23.878
13.744 | 34.927 | 26.545 | -0.099881 | 0.1005 | -7.023 | 26.545
13.744 | 35.128 | 29.478 | -0.099894 | 0.099 | -6.822 | 29.478
13.744 | 35.326 | 32.239 | -0.099889 | 0.1005 | -6.624 | 32.239
13.744 | 35.527 | 35.231 | -0.099877 | 0.1015 | -6.423 | 35.231
13.744 | 35.73 | 39.01 | -0.099875 | 0.1005 | -6.22 | 39.01
13.744 | 35.931 | 42.857 | -0.099883 | 0.0995 | -6.019 | 42.857
13.744 | 36.13 | 46.744 | -0.099873 | 0.0995 | -5.82 | 46.744
13.744 | 36.329 | 50.858 | -0.09988 | 0.1005 | -5.621 | 50.858
13.744 | 36.53 | 54.877 | -0.09988 | 0.1005 | -5.42 | 54.877
13.744 | 36.731 | 59.22 | -0.099883 | 0.098 | -5.219 | 59.22
13.744 | 36.927 | 63.293 | -0.099873 | 0.1 | -5.023 | 63.293
13.744 | 37.127 | 67.615 | -0.099894 | 0.1005 | -4.823 | 67.615
13.744 | 37.328 | 72.308 | -0.099873 | 0.0995 | -4.622 | 72.308
13.744 | 37.527 | 76.375 | -0.099873 | 0.099 | -4.423 | 76.375
13.744 | 37.725 | 80.713 | -0.099892 | 0.1005 | -4.225 | 80.713
13.744 | 37.926 | 84.685 | -0.099885 | 0.101 | -4.024 | 84.685
13.744 | 38.128 | 87.991 | -0.099882 | 0.101 | -3.822 | 87.991
13.744 | 38.33 | 91.992 | -0.099874 | 0.099 | -3.62 | 91.992
13.744 | 38.528 | 94.785 | -0.099875 | 0.1 | -3.422 | 94.785
13.744 | 38.728 | 97.628 | -0.099879 | 0.101 | -3.222 | 97.628
13.744 | 38.93 | 100.292 | -0.099881 | 0.1005 | -3.02 | 100.292
13.744 | 39.131 | 102.266 | -0.099891 | 0.098 | -2.819 | 102.266
13.744 | 39.327 | 104.822 | -0.099884 | 0.0995 | -2.623 | 104.822
13.744 | 39.526 | 106.313 | -0.099876 | 0.1005 | -2.424 | 106.313
13.744 | 39.727 | 107.818 | -0.099886 | 0.1 | -2.223 | 107.818
13.744 | 39.927 | 109.113 | -0.099885 | 0.0985 | -2.023 | 109.113
13.744 | 40.124 | 109.769 | -0.099887 | 0.1005 | -1.826 | 109.769
13.744 | 40.325 | 111.078 | -0.09988 | 0.1015 | -1.625 | 111.078
13.744 | 40.528 | 111.474 | -0.099879 | 0.1015 | -1.422 | 111.474
13.744 | 40.731 | 111.981 | -0.099893 | 0.099 | -1.219 | 111.981
13.744 | 40.929 | 112.436 | -0.099888 | 0.1 | -1.021 | 112.436
13.744 | 41.129 | 113.007 | -0.099874 | 0.1005 | -0.821 | 113.007
13.744 | 41.33 | 113.221 | -0.099878 | 0.1005 | -0.62 | 113.221
13.744 | 41.531 | 113.413 | -0.099895 | 0.099 | -0.419 | 113.413
13.744 | 41.729 | 113.59 | -0.099874 | 0.099 | -0.221 | 113.59
13.744 | 41.927 | 113.723 | -0.099883 | 0.1005 | -0.023 | 113.723
13.744 | 42.128 | 113.846 | -0.099887 | 0.1 | 0.178 | 113.846
13.744 | 42.328 | 113.953 | -0.099874 | 0.099 | 0.378 | 113.953
13.744 | 42.526 | 114.026 | -0.099882 | 0.0995 | 0.576 | 114.026
13.744 | 42.725 | 114.052 | -0.099879 | 0.1015 | 0.775 | 114.052
13.744 | 42.928 | 114.034 | -0.099891 | 0.101 | 0.978 | 114.034
13.744 | 43.13 | 113.931 | -0.099854 | 0.0995 | 1.18 | 113.931
13.744 | 43.329 | 113.737 | -0.099874 | 0.0995 | 1.379 | 113.737
13.744 | 43.528 | 113.433 | -0.099882 | 0.101 | 1.578 | 113.433
13.744 | 43.73 | 112.901 | -0.099874 | 0.1005 | 1.78 | 112.901
13.744 | 43.931 | 112.315 | -0.099876 | 0.0995 | 1.981 | 112.315
13.744 | 44.13 | 110.957 | -0.099869 | 0.099 | 2.18 | 110.957
13.744 | 44.328 | 109.725 | -0.099885 | 0.1005 | 2.378 | 109.725
13.744 | 44.529 | 108.319 | -0.099881 | 0.1 | 2.579 | 108.319
13.744 | 44.729 | 106.768 | -0.099874 | 0.0985 | 2.779 | 106.768
13.744 | 44.926 | 104.958 | -0.099882 | 0.0995 | 2.976 | 104.958
13.744 | 45.125 | 102.332 | -0.099881 | 0.101 | 3.175 | 102.332
13.744 | 45.327 | 99.764 | -0.09987 | 0.1015 | 3.377 | 99.764
13.744 | 45.53 | 96.922 | -0.099873 | 0.0995 | 3.58 | 96.922
13.744 | 45.729 | 93.722 | -0.099877 | 0.0995 | 3.779 | 93.722
13.744 | 45.928 | 89.971 | -0.099878 | 0.101 | 3.978 | 89.971
13.744 | 46.13 | 86.172 | -0.099881 | 0.101 | 4.18 | 86.172
13.744 | 46.332 | 82.288 | -0.099867 | 0.099 | 4.382 | 82.288
13.744 | 46.53 | 77.943 | -0.099887 | 0.099 | 4.58 | 77.943
13.744 | 46.728 | 73.297 | -0.09987 | 0.1005 | 4.778 | 73.297
13.744 | 46.929 | 68.99 | -0.099853 | 0.1005 | 4.979 | 68.99
13.744 | 47.13 | 64.239 | -0.099879 | 0.0985 | 5.18 | 64.239
13.744 | 47.327 | 60.023 | -0.099881 | 0.0995 | 5.377 | 60.023
13.744 | 47.526 | 55.622 | -0.099884 | 0.1005 | 5.576 | 55.622
13.744 | 47.727 | 51.487 | -0.099863 | 0.101 | 5.777 | 51.487
13.744 | 47.929 | 47.17 | -0.099873 | 0.1 | 5.979 | 47.17
13.744 | 48.129 | 43.196 | -0.099851 | 0.0995 | 6.179 | 43.196
13.744 | 48.328 | 39.403 | -0.099873 | 0.101 | 6.378 | 39.403
13.744 | 48.53 | 36.271 | -0.099858 | 0.1005 | 6.58 | 36.271
13.744 | 48.731 | 32.504 | -0.099877 | 0.1 | 6.781 | 32.504
13.744 | 48.931 | 29.853 | -0.099872 | 0.099 | 6.981 | 29.853
13.744 | 49.129 | 27.191 | -0.099877 | 0.1005 | 7.179 | 27.191
13.744 | 49.33 | 24.565 | -0.099872 | 0.1 | 7.38 | 24.565
13.744 | 49.53 | 22.653 | -0.099875 | 0.0995 | 7.58 | 22.653
13.744 | 49.729 | 20.714 | -0.099842 | 0.0985 | 7.779 | 20.714
13.744 | 49.926 | 18.863 | -0.099877 | 0.101 | 7.976 | 18.863
13.744 | 50.128 | 17.161 | -0.09987 | 0.101 | 8.178 | 17.161
13.744 | 50.33 | 15.571 | -0.099872 | 0.099 | 8.38 | 15.571
13.744 | 50.528 | 14.184 | -0.099878 | 0.0995 | 8.578 | 14.184
13.744 | 50.727 | 12.869 | -0.099865 | 0.1015 | 8.777 | 12.869
13.744 | 50.93 | 12.282 | -0.099873 | 0.1005 | 8.98 | 12.282
13.744 | 51.131 | 10.889 | -0.099876 | 0.0995 | 9.181 | 10.889
13.744 | 51.33 | 10.376 | -0.099876 | 0.0995 | 9.38 | 10.376
13.744 | 51.529 | 9.714 | -0.099874 | 0.1005 | 9.579 | 9.714
13.744 | 51.73 | 9.014 | -0.099882 | 0.1005 | 9.78 | 9.014
13.744 | 51.931 | 8.364 | -0.09988 | 0.2485 | 9.981 | 8.364
13.744 | 52.428 | 6.749 | -0.09988 | 0.25 | 10.478 | 6.749
13.744 | 52.928 | 6.02 | -0.099872 | 0.251 | 10.978 | 6.02
13.744 | 53.43 | 5.234 | -0.099882 | 0.2495 | 11.48 | 5.234
13.744 | 53.929 | 4.531 | -0.099887 | 0.251 | 11.979 | 4.531
13.744 | 54.431 | 3.934 | -0.099874 | 0.2485 | 12.481 | 3.934
13.744 | 54.928 | 3.439 | -0.099887 | 0.249 | 12.978 | 3.439
13.744 | 55.426 | 3.046 | -0.099887 | 0.2525 | 13.476 | 3.046
13.744 | 55.931 | 2.721 | -0.09987 | 0.2495 | 13.981 | 2.721
13.744 | 56.43 | 2.432 | -0.099878 | 0.2505 | 14.48 | 2.432
13.744 | 56.931 | 2.195 | -0.099864 | 0.2495 | 14.981 | 2.195
13.744 | 57.43 | 1.993 | -0.099876 | 0.2485 | 15.48 | 1.993
13.744 | 57.927 | 1.816 | -0.099877 | 0.252 | 15.977 | 1.816
13.744 | 58.431 | 1.664 | -0.099874 | 0.25 | 16.481 | 1.664
13.744 | 58.931 | 1.531 | -0.099851 | 0.249 | 16.981 | 1.531
13.744 | 59.429 | 1.418 | -0.09987 | 0.251 | 17.479 | 1.418
13.744 | 59.931 | 1.311 | -0.099877 | 0.248 | 17.981 | 1.311
13.744 | 60.427 | 1.225 | -0.099878 | 0.2515 | 18.477 | 1.225
13.744 | 60.93 | 1.143 | -0.099881 | 0.2505 | 18.98 | 1.143
13.744 | 61.431 | 1.07 | -0.099883 | 0.249 | 19.481 | 1.07
13.744 | 61.929 | 1.001 | -0.099875 | 0.251 | 19.979 | 1.001
13.744 | 62.431 | 0.941 | -0.099877 | 0.2485 | 20.481 | 0.941
13.744 | 62.928 | 0.895 | -0.099867 | 0.25 | 20.978 | 0.895
13.744 | 63.428 | 0.849 | -0.099872 | 0.251 | 21.478 | 0.849
13.744 | 63.93 | 0.808 | -0.099868 | 0.249 | 21.98 | 0.808
13.744 | 64.428 | 0.767 | -0.099869 | 0.251 | 22.478 | 0.767
13.744 | 64.93 | 0.732 | -0.099871 | 0.2495 | 22.98 | 0.732
13.744 | 65.429 | 0.696 | -0.099877 | 0.249 | 23.479 | 0.696
13.744 | 65.927 | 0.664 | -0.09988 | 0.252 | 23.977 | 0.664
13.744 | 66.431 | 0.634 | -0.099867 | 0.249 | 24.481 | 0.634
13.744 | 66.929 | 0.614 | -0.099858 | 0.25 | 24.979 | 0.614

## SW Run -2A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002012.stp | ! data file
2020-08-11 | ! date
13:34 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=-2A
! end File Header |  |  |  |  | I | -1.9995 | A
! begin Step Header |  |  |  |  | BdL | 1349.5008 | g-cm
0 | ! set current |  |  |  | B_center | 231.644 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.8258 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.939 | 0.744 | -0.19993 | 0.2465 | -25.011 | 0.744
13.744 | 17.432 | 0.63 | -0.19995 | 0.249 | -24.518 | 0.63
13.744 | 17.93 | 0.622 | -0.19992 | 0.2505 | -24.02 | 0.622
13.744 | 18.431 | 0.657 | -0.19995 | 0.25 | -23.519 | 0.657
13.744 | 18.931 | 0.704 | -0.19996 | 0.2485 | -23.019 | 0.704
13.744 | 19.428 | 0.758 | -0.19995 | 0.251 | -22.522 | 0.758
13.744 | 19.93 | 0.821 | -0.19995 | 0.2505 | -22.02 | 0.821
13.744 | 20.431 | 0.881 | -0.19996 | 0.249 | -21.519 | 0.881
13.744 | 20.929 | 0.95 | -0.19996 | 0.2515 | -21.021 | 0.95
13.744 | 21.432 | 1.03 | -0.19994 | 0.248 | -20.518 | 1.03
13.744 | 21.928 | 1.112 | -0.19995 | 0.2505 | -20.022 | 1.112
13.744 | 22.429 | 1.208 | -0.19995 | 0.251 | -19.521 | 1.208
13.744 | 22.931 | 1.315 | -0.19995 | 0.2485 | -19.019 | 1.315
13.744 | 23.428 | 1.433 | -0.19995 | 0.251 | -18.522 | 1.433
13.744 | 23.93 | 1.57 | -0.19994 | 0.2495 | -18.02 | 1.57
13.744 | 24.429 | 1.718 | -0.19995 | 0.2485 | -17.521 | 1.718
13.744 | 24.926 | 1.888 | -0.19994 | 0.2525 | -17.024 | 1.888
13.744 | 25.431 | 2.083 | -0.19995 | 0.249 | -16.519 | 2.083
13.744 | 25.929 | 2.304 | -0.19995 | 0.2505 | -16.021 | 2.304
13.744 | 26.43 | 2.553 | -0.19994 | 0.25 | -15.52 | 2.553
13.744 | 26.93 | 2.856 | -0.19994 | 0.248 | -15.02 | 2.856
13.744 | 27.426 | 3.194 | -0.19994 | 0.252 | -14.524 | 3.194
13.744 | 27.93 | 3.582 | -0.19994 | 0.2495 | -14.02 | 3.582
13.744 | 28.429 | 4.077 | -0.19994 | 0.2495 | -13.521 | 4.077
13.744 | 28.928 | 4.622 | -0.19993 | 0.2505 | -13.022 | 4.622
13.744 | 29.429 | 5.331 | -0.19994 | 0.2485 | -12.521 | 5.331
13.744 | 29.926 | 6.157 | -0.19994 | 0.2515 | -12.024 | 6.157
13.744 | 30.429 | 7.663 | -0.19995 | 0.2505 | -11.521 | 7.663
13.744 | 30.93 | 9.123 | -0.19994 | 0.249 | -11.02 | 9.123
13.744 | 31.428 | 10.708 | -0.19994 | 0.2505 | -10.522 | 10.708
13.744 | 31.929 | 13.271 | -0.19995 | 0.0985 | -10.021 | 13.271
13.744 | 32.126 | 14.515 | -0.19995 | 0.1 | -9.824 | 14.515
13.744 | 32.326 | 15.746 | -0.19995 | 0.101 | -9.624 | 15.746
13.744 | 32.528 | 17.058 | -0.19994 | 0.1 | -9.422 | 17.058
13.744 | 32.728 | 18.511 | -0.19995 | 0.0995 | -9.222 | 18.511
13.744 | 32.927 | 20.179 | -0.19993 | 0.1005 | -9.023 | 20.179
13.744 | 33.128 | 21.842 | -0.19995 | 0.1015 | -8.822 | 21.842
13.744 | 33.331 | 24.549 | -0.19995 | 0.1005 | -8.619 | 24.549
13.744 | 33.532 | 26.998 | -0.19995 | 0.0985 | -8.418 | 26.998
13.744 | 33.729 | 29.601 | -0.19994 | 0.0995 | -8.221 | 29.601
13.744 | 33.928 | 32.309 | -0.19995 | 0.101 | -8.022 | 32.309
13.744 | 34.13 | 35.571 | -0.19995 | 0.1 | -7.82 | 35.571
13.744 | 34.33 | 39.621 | -0.19995 | 0.098 | -7.62 | 39.621
13.744 | 34.526 | 43.633 | -0.19995 | 0.1 | -7.424 | 43.633
13.744 | 34.726 | 47.871 | -0.19995 | 0.1005 | -7.224 | 47.871
13.744 | 34.927 | 53.297 | -0.19994 | 0.1005 | -7.023 | 53.297
13.744 | 35.128 | 58.884 | -0.19995 | 0.099 | -6.822 | 58.884
13.744 | 35.326 | 64.911 | -0.19995 | 0.1005 | -6.624 | 64.911
13.744 | 35.527 | 71.675 | -0.19995 | 0.1015 | -6.423 | 71.675
13.744 | 35.73 | 78.837 | -0.19995 | 0.1005 | -6.22 | 78.837
13.744 | 35.931 | 86.047 | -0.19995 | 0.099 | -6.019 | 86.047
13.744 | 36.129 | 94.423 | -0.19995 | 0.1 | -5.821 | 94.423
13.744 | 36.329 | 102.611 | -0.19994 | 0.1005 | -5.621 | 102.611
13.744 | 36.53 | 111.455 | -0.19996 | 0.1 | -5.42 | 111.455
13.744 | 36.73 | 120.332 | -0.19995 | 0.0985 | -5.22 | 120.332
13.744 | 36.927 | 129.29 | -0.19995 | 0.0995 | -5.023 | 129.29
13.744 | 37.126 | 138.542 | -0.19995 | 0.1005 | -4.824 | 138.542
13.744 | 37.327 | 147.873 | -0.19996 | 0.1 | -4.623 | 147.873
13.744 | 37.527 | 156.177 | -0.19995 | 0.0985 | -4.423 | 156.177
13.744 | 37.724 | 164.589 | -0.19995 | 0.101 | -4.226 | 164.589
13.744 | 37.926 | 172.882 | -0.19995 | 0.101 | -4.024 | 172.882
13.744 | 38.128 | 180.51 | -0.19995 | 0.101 | -3.822 | 180.51
13.744 | 38.33 | 187.577 | -0.19996 | 0.099 | -3.62 | 187.577
13.744 | 38.528 | 193.817 | -0.19995 | 0.1 | -3.422 | 193.817
13.744 | 38.728 | 199.851 | -0.19996 | 0.101 | -3.222 | 199.851
13.744 | 38.93 | 205.404 | -0.19994 | 0.1 | -3.02 | 205.404
13.744 | 39.13 | 209.814 | -0.19995 | 0.0985 | -2.82 | 209.814
13.744 | 39.327 | 213.799 | -0.19995 | 0.0995 | -2.623 | 213.799
13.744 | 39.526 | 217.666 | -0.19995 | 0.1005 | -2.424 | 217.666
13.744 | 39.727 | 220.402 | -0.19994 | 0.1 | -2.223 | 220.402
13.744 | 39.927 | 223.042 | -0.19994 | 0.0985 | -2.023 | 223.042
13.744 | 40.124 | 224.841 | -0.19995 | 0.1005 | -1.826 | 224.841
13.744 | 40.325 | 226.597 | -0.19996 | 0.1015 | -1.625 | 226.597
13.744 | 40.528 | 228.156 | -0.19995 | 0.101 | -1.422 | 228.156
13.744 | 40.73 | 229.425 | -0.19995 | 0.099 | -1.22 | 229.425
13.744 | 40.928 | 229.811 | -0.19996 | 0.1 | -1.022 | 229.811
13.744 | 41.128 | 230.333 | -0.19994 | 0.101 | -0.822 | 230.333
13.744 | 41.33 | 230.794 | -0.19995 | 0.1005 | -0.62 | 230.794
13.744 | 41.531 | 231.15 | -0.19993 | 0.099 | -0.419 | 231.15
13.744 | 41.729 | 231.438 | -0.19995 | 0.0985 | -0.221 | 231.438
13.744 | 41.926 | 231.644 | -0.19995 | 0.1005 | -0.024 | 231.644
13.744 | 42.127 | 231.792 | -0.19994 | 0.1 | 0.177 | 231.792
13.744 | 42.327 | 231.879 | -0.19993 | 0.099 | 0.377 | 231.879
13.744 | 42.525 | 231.915 | -0.19995 | 0.1 | 0.575 | 231.915
13.744 | 42.725 | 231.877 | -0.19994 | 0.1015 | 0.775 | 231.877
13.744 | 42.928 | 231.701 | -0.19994 | 0.101 | 0.978 | 231.701
13.744 | 43.13 | 231.424 | -0.19995 | 0.0995 | 1.18 | 231.424
13.744 | 43.329 | 230.968 | -0.19995 | 0.0995 | 1.379 | 230.968
13.744 | 43.528 | 229.693 | -0.19995 | 0.101 | 1.578 | 229.693
13.744 | 43.73 | 228.46 | -0.19995 | 0.1005 | 1.78 | 228.46
13.744 | 43.931 | 227.042 | -0.19995 | 0.0995 | 1.981 | 227.042
13.744 | 44.13 | 225.233 | -0.19995 | 0.099 | 2.18 | 225.233
13.744 | 44.328 | 222.533 | -0.19994 | 0.1005 | 2.378 | 222.533
13.744 | 44.529 | 219.638 | -0.19995 | 0.1 | 2.579 | 219.638
13.744 | 44.729 | 215.716 | -0.19995 | 0.0985 | 2.779 | 215.716
13.744 | 44.926 | 211.957 | -0.19995 | 0.1 | 2.976 | 211.957
13.744 | 45.126 | 206.776 | -0.19993 | 0.101 | 3.176 | 206.776
13.744 | 45.328 | 201.335 | -0.19994 | 0.101 | 3.378 | 201.335
13.744 | 45.53 | 195.34 | -0.19994 | 0.0995 | 3.58 | 195.34
13.744 | 45.729 | 188.918 | -0.19994 | 0.0995 | 3.779 | 188.918
13.744 | 45.928 | 181.526 | -0.19995 | 0.101 | 3.978 | 181.526
13.744 | 46.13 | 173.603 | -0.19995 | 0.101 | 4.18 | 173.603
13.744 | 46.332 | 165.59 | -0.19995 | 0.099 | 4.382 | 165.59
13.744 | 46.53 | 156.854 | -0.19996 | 0.099 | 4.58 | 156.854
13.744 | 46.728 | 148.113 | -0.19995 | 0.1005 | 4.778 | 148.113
13.744 | 46.929 | 138.905 | -0.19995 | 0.1 | 4.979 | 138.905
13.744 | 47.129 | 129.874 | -0.19995 | 0.099 | 5.179 | 129.874
13.744 | 47.327 | 120.746 | -0.19995 | 0.099 | 5.377 | 120.746
13.744 | 47.525 | 111.229 | -0.19995 | 0.101 | 5.575 | 111.229
13.744 | 47.727 | 102.792 | -0.19995 | 0.101 | 5.777 | 102.792
13.744 | 47.929 | 93.985 | -0.19993 | 0.1 | 5.979 | 93.985
13.744 | 48.129 | 85.725 | -0.19994 | 0.0995 | 6.179 | 85.725
13.744 | 48.328 | 78.55 | -0.19995 | 0.101 | 6.378 | 78.55
13.744 | 48.53 | 70.986 | -0.19993 | 0.1005 | 6.58 | 70.986
13.744 | 48.731 | 64.528 | -0.19995 | 0.0995 | 6.781 | 64.528
13.744 | 48.93 | 58.724 | -0.19995 | 0.099 | 6.98 | 58.724
13.744 | 49.128 | 52.912 | -0.19995 | 0.101 | 7.178 | 52.912
13.744 | 49.33 | 47.833 | -0.19994 | 0.1 | 7.38 | 47.833
13.744 | 49.53 | 43.205 | -0.19993 | 0.099 | 7.58 | 43.205
13.744 | 49.728 | 39.243 | -0.19993 | 0.099 | 7.778 | 39.243
13.744 | 49.926 | 36.071 | -0.19992 | 0.101 | 7.976 | 36.071
13.744 | 50.128 | 32.462 | -0.19995 | 0.1005 | 8.178 | 32.462
13.744 | 50.329 | 29.579 | -0.19995 | 0.0995 | 8.379 | 29.579
13.744 | 50.528 | 26.921 | -0.19995 | 0.0995 | 8.578 | 26.921
13.744 | 50.727 | 24.438 | -0.19993 | 0.101 | 8.777 | 24.438
13.744 | 50.929 | 22.655 | -0.19995 | 0.101 | 8.979 | 22.655
13.744 | 51.131 | 20.798 | -0.19995 | 0.0995 | 9.181 | 20.798
13.744 | 51.33 | 18.888 | -0.19994 | 0.099 | 9.38 | 18.888
13.744 | 51.528 | 17.379 | -0.19995 | 0.101 | 9.578 | 17.379
13.744 | 51.73 | 15.806 | -0.19995 | 0.1005 | 9.78 | 15.806
13.744 | 51.931 | 14.379 | -0.19994 | 0.248 | 9.981 | 14.379
13.744 | 52.427 | 11.955 | -0.19995 | 0.2505 | 10.477 | 11.955
13.744 | 52.928 | 10.246 | -0.19995 | 0.251 | 10.978 | 10.246
13.744 | 53.43 | 8.573 | -0.19995 | 0.2495 | 11.48 | 8.573
13.744 | 53.929 | 7.216 | -0.19993 | 0.251 | 11.979 | 7.216
13.744 | 54.431 | 6.552 | -0.19993 | 0.2485 | 12.481 | 6.552
13.744 | 54.928 | 5.787 | -0.19994 | 0.2495 | 12.978 | 5.787
13.744 | 55.427 | 5.075 | -0.19994 | 0.252 | 13.477 | 5.075
13.744 | 55.931 | 4.472 | -0.19995 | 0.249 | 13.981 | 4.472
13.744 | 56.429 | 3.961 | -0.19995 | 0.251 | 14.479 | 3.961
13.744 | 56.931 | 3.526 | -0.19994 | 0.2495 | 14.981 | 3.526
13.744 | 57.43 | 3.167 | -0.19995 | 0.2485 | 15.48 | 3.167
13.744 | 57.927 | 2.862 | -0.19995 | 0.252 | 15.977 | 2.862
13.744 | 58.431 | 2.586 | -0.19995 | 0.2495 | 16.481 | 2.586
13.744 | 58.93 | 2.36 | -0.19994 | 0.2495 | 16.98 | 2.36
13.744 | 59.429 | 2.157 | -0.19994 | 0.251 | 17.479 | 2.157
13.744 | 59.931 | 1.978 | -0.19995 | 0.2485 | 17.981 | 1.978
13.744 | 60.428 | 1.827 | -0.19995 | 0.251 | 18.478 | 1.827
13.744 | 60.93 | 1.687 | -0.19995 | 0.2505 | 18.98 | 1.687
13.744 | 61.431 | 1.573 | -0.19996 | 0.249 | 19.481 | 1.573
13.744 | 61.929 | 1.462 | -0.19995 | 0.251 | 19.979 | 1.462
13.744 | 62.431 | 1.361 | -0.19995 | 0.2485 | 20.481 | 1.361
13.744 | 62.928 | 1.282 | -0.19993 | 0.25 | 20.978 | 1.282
13.744 | 63.428 | 1.205 | -0.19995 | 0.251 | 21.478 | 1.205
13.744 | 63.93 | 1.145 | -0.19994 | 0.249 | 21.98 | 1.145
13.744 | 64.428 | 1.087 | -0.19994 | 0.251 | 22.478 | 1.087
13.744 | 64.93 | 1.019 | -0.19994 | 0.2495 | 22.98 | 1.019
13.744 | 65.429 | 0.964 | -0.19994 | 0.249 | 23.479 | 0.964
13.744 | 65.927 | 0.913 | -0.19993 | 0.252 | 23.977 | 0.913
13.744 | 66.431 | 0.873 | -0.19994 | 0.249 | 24.481 | 0.873
13.744 | 66.929 | 0.831 | -0.19994 | 0.25 | 24.979 | 0.831

## SW Run -2.8A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002013.stp | ! data file
2020-08-11 | ! date
13:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=-2.8A
! end File Header |  |  |  |  | I | -2.7992 | A
! begin Step Header |  |  |  |  | BdL | 1891.1695 | g-cm
0 | ! set current |  |  |  | B_center | 325.902 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.8029 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.939 | 0.991 | -0.27992 | 0.2465 | -25.011 | 0.991
13.744 | 17.432 | 0.827 | -0.27993 | 0.2485 | -24.518 | 0.827
13.744 | 17.929 | 0.815 | -0.27991 | 0.251 | -24.021 | 0.815
13.744 | 18.431 | 0.851 | -0.27992 | 0.25 | -23.519 | 0.851
13.744 | 18.931 | 0.912 | -0.27993 | 0.248 | -23.019 | 0.912
13.744 | 19.427 | 0.978 | -0.27992 | 0.252 | -22.523 | 0.978
13.744 | 19.931 | 1.054 | -0.27992 | 0.2495 | -22.019 | 1.054
13.744 | 20.43 | 1.147 | -0.27992 | 0.249 | -21.52 | 1.147
13.744 | 20.928 | 1.24 | -0.27992 | 0.2515 | -21.022 | 1.24
13.744 | 21.431 | 1.337 | -0.27991 | 0.249 | -20.519 | 1.337
13.744 | 21.929 | 1.436 | -0.27991 | 0.25 | -20.021 | 1.436
13.744 | 22.429 | 1.566 | -0.27991 | 0.251 | -19.521 | 1.566
13.744 | 22.931 | 1.703 | -0.27992 | 0.2485 | -19.019 | 1.703
13.744 | 23.428 | 1.859 | -0.27992 | 0.251 | -18.522 | 1.859
13.744 | 23.93 | 2.041 | -0.27991 | 0.249 | -18.02 | 2.041
13.744 | 24.428 | 2.234 | -0.27991 | 0.249 | -17.522 | 2.234
13.744 | 24.926 | 2.464 | -0.27992 | 0.252 | -17.024 | 2.464
13.744 | 25.43 | 2.716 | -0.27992 | 0.2495 | -16.52 | 2.716
13.744 | 25.929 | 3.004 | -0.27992 | 0.2505 | -16.021 | 3.004
13.744 | 26.43 | 3.345 | -0.27992 | 0.25 | -15.52 | 3.345
13.744 | 26.93 | 3.731 | -0.27992 | 0.248 | -15.02 | 3.731
13.744 | 27.426 | 4.199 | -0.27992 | 0.252 | -14.524 | 4.199
13.744 | 27.93 | 4.736 | -0.27992 | 0.25 | -14.02 | 4.736
13.744 | 28.43 | 5.371 | -0.27992 | 0.2485 | -13.52 | 5.371
13.744 | 28.927 | 6.156 | -0.27992 | 0.251 | -13.023 | 6.156
13.744 | 29.429 | 7.582 | -0.27993 | 0.2485 | -12.521 | 7.582
13.744 | 29.926 | 8.884 | -0.27993 | 0.2515 | -12.024 | 8.884
13.744 | 30.429 | 10.398 | -0.27992 | 0.2505 | -11.521 | 10.398
13.744 | 30.93 | 12.076 | -0.27993 | 0.2485 | -11.02 | 12.076
13.744 | 31.427 | 14.841 | -0.27992 | 0.251 | -10.523 | 14.841
13.744 | 31.929 | 17.829 | -0.27992 | 0.0985 | -10.021 | 17.829
13.744 | 32.126 | 19.46 | -0.27992 | 0.1 | -9.824 | 19.46
13.744 | 32.326 | 21.235 | -0.27992 | 0.101 | -9.624 | 21.235
13.744 | 32.528 | 23.559 | -0.27992 | 0.1 | -9.422 | 23.559
13.744 | 32.728 | 25.309 | -0.27993 | 0.0995 | -9.222 | 25.309
13.744 | 32.927 | 28.021 | -0.27992 | 0.1005 | -9.023 | 28.021
13.744 | 33.128 | 30.571 | -0.27992 | 0.101 | -8.822 | 30.571
13.744 | 33.33 | 33.597 | -0.27993 | 0.1005 | -8.62 | 33.597
13.744 | 33.531 | 36.995 | -0.27993 | 0.099 | -8.419 | 36.995
13.744 | 33.729 | 40.997 | -0.27993 | 0.0995 | -8.221 | 40.997
13.744 | 33.928 | 45.147 | -0.27992 | 0.1005 | -8.022 | 45.147
13.744 | 34.129 | 49.982 | -0.27992 | 0.1005 | -7.821 | 49.982
13.744 | 34.33 | 55.24 | -0.27991 | 0.098 | -7.62 | 55.24
13.744 | 34.526 | 60.931 | -0.27992 | 0.1 | -7.424 | 60.931
13.744 | 34.726 | 67.23 | -0.27992 | 0.1005 | -7.224 | 67.23
13.744 | 34.927 | 74.321 | -0.27992 | 0.1005 | -7.023 | 74.321
13.744 | 35.128 | 82.143 | -0.27992 | 0.099 | -6.822 | 82.143
13.744 | 35.326 | 90.605 | -0.27992 | 0.1005 | -6.624 | 90.605
13.744 | 35.527 | 100.157 | -0.27992 | 0.101 | -6.423 | 100.157
13.744 | 35.729 | 110.314 | -0.27992 | 0.101 | -6.221 | 110.314
13.744 | 35.931 | 121.01 | -0.27993 | 0.099 | -6.019 | 121.01
13.744 | 36.129 | 132.09 | -0.27992 | 0.0995 | -5.821 | 132.09
13.744 | 36.328 | 144.47 | -0.27992 | 0.101 | -5.622 | 144.47
13.744 | 36.53 | 156.825 | -0.27992 | 0.1 | -5.42 | 156.825
13.744 | 36.73 | 169.332 | -0.27992 | 0.0985 | -5.22 | 169.332
13.744 | 36.927 | 182.173 | -0.27991 | 0.0995 | -5.023 | 182.173
13.744 | 37.126 | 194.981 | -0.27992 | 0.1005 | -4.824 | 194.981
13.744 | 37.327 | 207.638 | -0.27992 | 0.0995 | -4.623 | 207.638
13.744 | 37.526 | 220.049 | -0.27992 | 0.099 | -4.424 | 220.049
13.744 | 37.724 | 231.996 | -0.27992 | 0.1005 | -4.226 | 231.996
13.744 | 37.925 | 243.23 | -0.27992 | 0.1015 | -4.025 | 243.23
13.744 | 38.128 | 254.179 | -0.27991 | 0.101 | -3.822 | 254.179
13.744 | 38.33 | 264.226 | -0.27992 | 0.099 | -3.62 | 264.226
13.744 | 38.528 | 273.616 | -0.27992 | 0.1 | -3.422 | 273.616
13.744 | 38.728 | 281.928 | -0.27992 | 0.1005 | -3.222 | 281.928
13.744 | 38.929 | 288.99 | -0.27992 | 0.1005 | -3.021 | 288.99
13.744 | 39.13 | 295.851 | -0.27993 | 0.098 | -2.82 | 295.851
13.744 | 39.326 | 301.705 | -0.27992 | 0.1 | -2.624 | 301.705
13.744 | 39.526 | 306.105 | -0.27991 | 0.1005 | -2.424 | 306.105
13.744 | 39.727 | 310.894 | -0.27992 | 0.1 | -2.223 | 310.894
13.744 | 39.927 | 313.953 | -0.27992 | 0.0985 | -2.023 | 313.953
13.744 | 40.124 | 317.015 | -0.27992 | 0.101 | -1.826 | 317.015
13.744 | 40.326 | 319.513 | -0.27992 | 0.101 | -1.624 | 319.513
13.744 | 40.528 | 321.079 | -0.27992 | 0.101 | -1.422 | 321.079
13.744 | 40.73 | 322.647 | -0.27992 | 0.099 | -1.22 | 322.647
13.744 | 40.928 | 323.908 | -0.27992 | 0.1 | -1.022 | 323.908
13.744 | 41.128 | 324.293 | -0.27992 | 0.101 | -0.822 | 324.293
13.744 | 41.33 | 324.879 | -0.27992 | 0.1005 | -0.62 | 324.879
13.744 | 41.531 | 325.267 | -0.27992 | 0.099 | -0.419 | 325.267
13.744 | 41.729 | 325.788 | -0.27992 | 0.0985 | -0.221 | 325.788
13.744 | 41.926 | 325.902 | -0.27992 | 0.1005 | -0.024 | 325.902
13.744 | 42.127 | 326.011 | -0.27992 | 0.1 | 0.177 | 326.011
13.744 | 42.327 | 326.08 | -0.27992 | 0.099 | 0.377 | 326.08
13.744 | 42.525 | 326.08 | -0.27992 | 0.1 | 0.575 | 326.08
13.744 | 42.725 | 325.976 | -0.27992 | 0.1015 | 0.775 | 325.976
13.744 | 42.928 | 325.715 | -0.27992 | 0.1005 | 0.978 | 325.715
13.744 | 43.129 | 325.262 | -0.27992 | 0.1 | 1.179 | 325.262
13.744 | 43.329 | 323.96 | -0.27993 | 0.0995 | 1.379 | 323.96
13.744 | 43.528 | 322.715 | -0.27992 | 0.101 | 1.578 | 322.715
13.744 | 43.73 | 321.216 | -0.27992 | 0.1005 | 1.78 | 321.216
13.744 | 43.931 | 318.627 | -0.27992 | 0.099 | 1.981 | 318.627
13.744 | 44.129 | 315.853 | -0.27992 | 0.0995 | 2.179 | 315.853
13.744 | 44.328 | 312.707 | -0.27992 | 0.1005 | 2.378 | 312.707
13.744 | 44.529 | 307.874 | -0.27992 | 0.1 | 2.579 | 307.874
13.744 | 44.729 | 302.909 | -0.27992 | 0.0985 | 2.779 | 302.909
13.744 | 44.926 | 297.013 | -0.27992 | 0.0995 | 2.976 | 297.013
13.744 | 45.125 | 290.622 | -0.27992 | 0.101 | 3.175 | 290.622
13.744 | 45.327 | 283.121 | -0.27991 | 0.101 | 3.377 | 283.121
13.744 | 45.529 | 274.455 | -0.27992 | 0.0995 | 3.579 | 274.455
13.744 | 45.728 | 265.219 | -0.27992 | 0.1 | 3.778 | 265.219
13.744 | 45.928 | 254.853 | -0.27991 | 0.101 | 3.978 | 254.853
13.744 | 46.13 | 243.627 | -0.27993 | 0.1005 | 4.18 | 243.627
13.744 | 46.331 | 232.405 | -0.27992 | 0.0995 | 4.381 | 232.405
13.744 | 46.53 | 219.863 | -0.27992 | 0.099 | 4.58 | 219.863
13.744 | 46.728 | 207.7 | -0.27991 | 0.1005 | 4.778 | 207.7
13.744 | 46.929 | 194.478 | -0.27992 | 0.1 | 4.979 | 194.478
13.744 | 47.129 | 181.661 | -0.27992 | 0.099 | 5.179 | 181.661
13.744 | 47.327 | 169.02 | -0.27992 | 0.099 | 5.377 | 169.02
13.744 | 47.525 | 155.943 | -0.27991 | 0.101 | 5.575 | 155.943
13.744 | 47.727 | 143.381 | -0.27991 | 0.101 | 5.777 | 143.381
13.744 | 47.929 | 131.48 | -0.27992 | 0.0995 | 5.979 | 131.48
13.744 | 48.128 | 120.017 | -0.27991 | 0.0995 | 6.178 | 120.017
13.744 | 48.327 | 109.607 | -0.27992 | 0.101 | 6.377 | 109.607
13.744 | 48.529 | 99.445 | -0.27991 | 0.101 | 6.579 | 99.445
13.744 | 48.731 | 90.241 | -0.27992 | 0.0995 | 6.781 | 90.241
13.744 | 48.93 | 81.349 | -0.27992 | 0.099 | 6.98 | 81.349
13.744 | 49.128 | 74.105 | -0.27992 | 0.101 | 7.178 | 74.105
13.744 | 49.33 | 66.811 | -0.27992 | 0.1 | 7.38 | 66.811
13.744 | 49.53 | 60.461 | -0.27992 | 0.099 | 7.58 | 60.461
13.744 | 49.728 | 54.651 | -0.27993 | 0.099 | 7.778 | 54.651
13.744 | 49.926 | 49.375 | -0.27992 | 0.1005 | 7.976 | 49.375
13.744 | 50.127 | 45.029 | -0.27992 | 0.101 | 8.177 | 45.029
13.744 | 50.329 | 40.666 | -0.27992 | 0.0995 | 8.379 | 40.666
13.744 | 50.528 | 37.151 | -0.27992 | 0.0995 | 8.578 | 37.151
13.744 | 50.727 | 33.955 | -0.27993 | 0.101 | 8.777 | 33.955
13.744 | 50.929 | 30.927 | -0.27992 | 0.101 | 8.979 | 30.927
13.744 | 51.131 | 28.238 | -0.27993 | 0.0995 | 9.181 | 28.238
13.744 | 51.33 | 25.495 | -0.27992 | 0.099 | 9.38 | 25.495
13.744 | 51.528 | 23.711 | -0.27992 | 0.101 | 9.578 | 23.711
13.744 | 51.73 | 21.836 | -0.27992 | 0.1005 | 9.78 | 21.836
13.744 | 51.931 | 20.134 | -0.27992 | 0.248 | 9.981 | 20.134
13.744 | 52.427 | 16.337 | -0.27992 | 0.2505 | 10.477 | 16.337
13.744 | 52.928 | 13.584 | -0.27992 | 0.251 | 10.978 | 13.584
13.744 | 53.43 | 11.743 | -0.27991 | 0.249 | 11.48 | 11.743
13.744 | 53.928 | 9.973 | -0.27991 | 0.2515 | 11.978 | 9.973
13.744 | 54.431 | 8.428 | -0.27992 | 0.2485 | 12.481 | 8.428
13.744 | 54.928 | 7.089 | -0.27992 | 0.249 | 12.978 | 7.089
13.744 | 55.426 | 6.548 | -0.27991 | 0.2525 | 13.476 | 6.548
13.744 | 55.931 | 5.835 | -0.27991 | 0.249 | 13.981 | 5.835
13.744 | 56.429 | 5.182 | -0.27991 | 0.251 | 14.479 | 5.182
13.744 | 56.931 | 4.592 | -0.27991 | 0.249 | 14.981 | 4.592
13.744 | 57.429 | 4.121 | -0.27991 | 0.2485 | 15.479 | 4.121
13.744 | 57.926 | 3.704 | -0.27991 | 0.2525 | 15.976 | 3.704
13.744 | 58.431 | 3.333 | -0.27992 | 0.2495 | 16.481 | 3.333
13.744 | 58.93 | 3.032 | -0.27992 | 0.2495 | 16.98 | 3.032
13.744 | 59.429 | 2.759 | -0.27992 | 0.2505 | 17.479 | 2.759
13.744 | 59.93 | 2.521 | -0.27992 | 0.2485 | 17.98 | 2.521
13.744 | 60.427 | 2.329 | -0.27992 | 0.2515 | 18.477 | 2.329
13.744 | 60.93 | 2.142 | -0.27993 | 0.2505 | 18.98 | 2.142
13.744 | 61.431 | 1.987 | -0.27991 | 0.249 | 19.481 | 1.987
13.744 | 61.929 | 1.845 | -0.27992 | 0.251 | 19.979 | 1.845
13.744 | 62.431 | 1.723 | -0.27992 | 0.248 | 20.481 | 1.723
13.744 | 62.927 | 1.608 | -0.27993 | 0.25 | 20.977 | 1.608
13.744 | 63.427 | 1.528 | -0.27991 | 0.2515 | 21.477 | 1.528
13.744 | 63.93 | 1.454 | -0.27991 | 0.249 | 21.98 | 1.454
13.744 | 64.428 | 1.34 | -0.27992 | 0.251 | 22.478 | 1.34
13.744 | 64.93 | 1.263 | -0.27993 | 0.2495 | 22.98 | 1.263
13.744 | 65.429 | 1.194 | -0.27993 | 0.249 | 23.479 | 1.194
13.744 | 65.927 | 1.129 | -0.27992 | 0.2515 | 23.977 | 1.129
13.744 | 66.43 | 1.071 | -0.27992 | 0.2495 | 24.48 | 1.071
13.744 | 66.929 | 1.002 | -0.27992 | 0.25 | 24.979 | 1.002

## SW Run -3A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002014.stp | ! data file
2020-08-11 | ! date
14:10 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=-3A
! end File Header |  |  |  |  | I | -3.0003 | A
! begin Step Header |  |  |  |  | BdL | 2026.7307 | g-cm
0 | ! set current |  |  |  | B_center | 349.534 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.7984 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.94 | 1.016 | -0.30003 | 0.246 | -25.01 | 1.016
13.744 | 17.432 | 0.853 | -0.30003 | 0.248 | -24.518 | 0.853
13.744 | 17.928 | 0.85 | -0.30004 | 0.251 | -24.022 | 0.85
13.744 | 18.43 | 0.899 | -0.30003 | 0.25 | -23.52 | 0.899
13.744 | 18.93 | 0.955 | -0.30004 | 0.248 | -23.02 | 0.955
13.744 | 19.426 | 1.033 | -0.30003 | 0.251 | -22.524 | 1.033
13.744 | 19.928 | 1.117 | -0.30003 | 0.2505 | -22.022 | 1.117
13.744 | 20.429 | 1.208 | -0.30003 | 0.2495 | -21.521 | 1.208
13.744 | 20.928 | 1.31 | -0.30003 | 0.251 | -21.022 | 1.31
13.744 | 21.43 | 1.415 | -0.30004 | 0.2485 | -20.52 | 1.415
13.744 | 21.927 | 1.534 | -0.30003 | 0.2505 | -20.023 | 1.534
13.744 | 22.428 | 1.663 | -0.30003 | 0.2505 | -19.522 | 1.663
13.744 | 22.929 | 1.811 | -0.30003 | 0.2505 | -19.021 | 1.811
13.744 | 23.43 | 1.977 | -0.30003 | 0.25 | -18.52 | 1.977
13.744 | 23.93 | 2.161 | -0.30002 | 0.249 | -18.02 | 2.161
13.744 | 24.428 | 2.384 | -0.30004 | 0.249 | -17.522 | 2.384
13.744 | 24.926 | 2.618 | -0.30004 | 0.252 | -17.024 | 2.618
13.744 | 25.43 | 2.887 | -0.30004 | 0.2495 | -16.52 | 2.887
13.744 | 25.929 | 3.201 | -0.30005 | 0.2505 | -16.021 | 3.201
13.744 | 26.43 | 3.576 | -0.30004 | 0.2495 | -15.52 | 3.576
13.744 | 26.929 | 3.982 | -0.30003 | 0.2485 | -15.021 | 3.982
13.744 | 27.426 | 4.455 | -0.30003 | 0.252 | -14.524 | 4.455
13.744 | 27.93 | 5.045 | -0.30003 | 0.2495 | -14.02 | 5.045
13.744 | 28.429 | 5.711 | -0.30002 | 0.249 | -13.521 | 5.711
13.744 | 28.927 | 7.085 | -0.30004 | 0.251 | -13.023 | 7.085
13.744 | 29.429 | 7.693 | -0.30003 | 0.2485 | -12.521 | 7.693
13.744 | 29.926 | 9.27 | -0.30003 | 0.251 | -12.024 | 9.27
13.744 | 30.428 | 10.916 | -0.30004 | 0.251 | -11.522 | 10.916
13.744 | 30.93 | 12.846 | -0.30002 | 0.2485 | -11.02 | 12.846
13.744 | 31.427 | 15.826 | -0.30003 | 0.251 | -10.523 | 15.826
13.744 | 31.929 | 19.436 | -0.30003 | 0.0985 | -10.021 | 19.436
13.744 | 32.126 | 20.956 | -0.30003 | 0.1 | -9.824 | 20.956
13.744 | 32.326 | 22.733 | -0.30004 | 0.1005 | -9.624 | 22.733
13.744 | 32.527 | 24.759 | -0.30004 | 0.1005 | -9.423 | 24.759
13.744 | 32.728 | 27.349 | -0.30003 | 0.0995 | -9.222 | 27.349
13.744 | 32.927 | 29.819 | -0.30003 | 0.1005 | -9.023 | 29.819
13.744 | 33.128 | 32.844 | -0.30003 | 0.101 | -8.822 | 32.844
13.744 | 33.33 | 35.891 | -0.30003 | 0.1005 | -8.62 | 35.891
13.744 | 33.531 | 39.797 | -0.30002 | 0.099 | -8.419 | 39.797
13.744 | 33.729 | 43.779 | -0.30003 | 0.0995 | -8.221 | 43.779
13.744 | 33.928 | 48.093 | -0.30003 | 0.1005 | -8.022 | 48.093
13.744 | 34.129 | 53.399 | -0.30005 | 0.1 | -7.821 | 53.399
13.744 | 34.329 | 59.153 | -0.30004 | 0.0985 | -7.621 | 59.153
13.744 | 34.526 | 65.206 | -0.30004 | 0.0995 | -7.424 | 65.206
13.744 | 34.725 | 71.898 | -0.30004 | 0.101 | -7.225 | 71.898
13.744 | 34.927 | 79.736 | -0.30004 | 0.1005 | -7.023 | 79.736
13.744 | 35.128 | 87.888 | -0.30003 | 0.099 | -6.822 | 87.888
13.744 | 35.326 | 97.257 | -0.30004 | 0.1005 | -6.624 | 97.257
13.744 | 35.527 | 107.15 | -0.30004 | 0.101 | -6.423 | 107.15
13.744 | 35.729 | 118.384 | -0.30004 | 0.101 | -6.221 | 118.384
13.744 | 35.931 | 130.125 | -0.30004 | 0.099 | -6.019 | 130.125
13.744 | 36.129 | 142.042 | -0.30003 | 0.0995 | -5.821 | 142.042
13.744 | 36.328 | 154.868 | -0.30004 | 0.101 | -5.622 | 154.868
13.744 | 36.53 | 168.385 | -0.30004 | 0.1 | -5.42 | 168.385
13.744 | 36.73 | 181.884 | -0.30004 | 0.098 | -5.22 | 181.884
13.744 | 36.926 | 195.337 | -0.30003 | 0.1 | -5.024 | 195.337
13.744 | 37.126 | 209.291 | -0.30004 | 0.1005 | -4.824 | 209.291
13.744 | 37.327 | 222.507 | -0.30003 | 0.0995 | -4.623 | 222.507
13.744 | 37.526 | 236.321 | -0.30004 | 0.099 | -4.424 | 236.321
13.744 | 37.724 | 249.016 | -0.30004 | 0.1005 | -4.226 | 249.016
13.744 | 37.925 | 261.013 | -0.30003 | 0.1015 | -4.025 | 261.013
13.744 | 38.128 | 272.982 | -0.30003 | 0.101 | -3.822 | 272.982
13.744 | 38.33 | 283.39 | -0.30003 | 0.099 | -3.62 | 283.39
13.744 | 38.528 | 293.089 | -0.30003 | 0.1 | -3.422 | 293.089
13.744 | 38.728 | 302.163 | -0.30004 | 0.1005 | -3.222 | 302.163
13.744 | 38.929 | 310.495 | -0.30004 | 0.1005 | -3.021 | 310.495
13.744 | 39.13 | 317.562 | -0.30004 | 0.098 | -2.82 | 317.562
13.744 | 39.326 | 323.457 | -0.30004 | 0.0995 | -2.624 | 323.457
13.744 | 39.525 | 328.912 | -0.30003 | 0.101 | -2.425 | 328.912
13.744 | 39.727 | 333.244 | -0.30004 | 0.1 | -2.223 | 333.244
13.744 | 39.927 | 337.245 | -0.30004 | 0.098 | -2.023 | 337.245
13.744 | 40.123 | 340.133 | -0.30003 | 0.101 | -1.827 | 340.133
13.744 | 40.325 | 342.679 | -0.30003 | 0.1015 | -1.625 | 342.679
13.744 | 40.528 | 344.343 | -0.30004 | 0.101 | -1.422 | 344.343
13.744 | 40.73 | 346.004 | -0.30003 | 0.099 | -1.22 | 346.004
13.744 | 40.928 | 347.379 | -0.30003 | 0.1 | -1.022 | 347.379
13.744 | 41.128 | 347.852 | -0.30004 | 0.101 | -0.822 | 347.852
13.744 | 41.33 | 348.399 | -0.30004 | 0.1 | -0.62 | 348.399
13.744 | 41.53 | 348.951 | -0.30003 | 0.099 | -0.42 | 348.951
13.744 | 41.728 | 349.26 | -0.30003 | 0.099 | -0.222 | 349.26
13.744 | 41.926 | 349.534 | -0.30005 | 0.1005 | -0.024 | 349.534
13.744 | 42.127 | 349.67 | -0.30003 | 0.1 | 0.177 | 349.67
13.744 | 42.327 | 349.741 | -0.30004 | 0.099 | 0.377 | 349.741
13.744 | 42.525 | 349.73 | -0.30004 | 0.1 | 0.575 | 349.73
13.744 | 42.725 | 349.609 | -0.30002 | 0.101 | 0.775 | 349.609
13.744 | 42.927 | 349.331 | -0.30003 | 0.101 | 0.977 | 349.331
13.744 | 43.129 | 348.75 | -0.30003 | 0.0995 | 1.179 | 348.75
13.744 | 43.328 | 347.503 | -0.30004 | 0.1 | 1.378 | 347.503
13.744 | 43.528 | 346.062 | -0.30003 | 0.101 | 1.578 | 346.062
13.744 | 43.73 | 344.333 | -0.30002 | 0.1005 | 1.78 | 344.333
13.744 | 43.931 | 341.836 | -0.30003 | 0.099 | 1.981 | 341.836
13.744 | 44.129 | 338.81 | -0.30002 | 0.099 | 2.179 | 338.81
13.744 | 44.327 | 334.768 | -0.30004 | 0.1005 | 2.377 | 334.768
13.744 | 44.528 | 330.622 | -0.30003 | 0.1005 | 2.578 | 330.622
13.744 | 44.729 | 324.821 | -0.30004 | 0.098 | 2.779 | 324.821
13.744 | 44.925 | 319.049 | -0.30004 | 0.1 | 2.975 | 319.049
13.744 | 45.125 | 311.419 | -0.30003 | 0.101 | 3.175 | 311.419
13.744 | 45.327 | 303.269 | -0.30004 | 0.101 | 3.377 | 303.269
13.744 | 45.529 | 294.165 | -0.30004 | 0.0995 | 3.579 | 294.165
13.744 | 45.728 | 284.082 | -0.30004 | 0.1 | 3.778 | 284.082
13.744 | 45.928 | 273.311 | -0.30004 | 0.101 | 3.978 | 273.311
13.744 | 46.13 | 261.315 | -0.30004 | 0.1005 | 4.18 | 261.315
13.744 | 46.331 | 248.93 | -0.30004 | 0.099 | 4.381 | 248.93
13.744 | 46.529 | 235.813 | -0.30003 | 0.0995 | 4.579 | 235.813
13.744 | 46.728 | 222.752 | -0.30004 | 0.1005 | 4.778 | 222.752
13.744 | 46.929 | 208.697 | -0.30004 | 0.1 | 4.979 | 208.697
13.744 | 47.129 | 194.433 | -0.30003 | 0.0985 | 5.179 | 194.433
13.744 | 47.326 | 180.833 | -0.30004 | 0.0995 | 5.376 | 180.833
13.744 | 47.525 | 167.093 | -0.30004 | 0.101 | 5.575 | 167.093
13.744 | 47.727 | 153.586 | -0.30003 | 0.101 | 5.777 | 153.586
13.744 | 47.929 | 141.087 | -0.30004 | 0.0995 | 5.979 | 141.087
13.744 | 48.128 | 128.544 | -0.30003 | 0.0995 | 6.178 | 128.544
13.744 | 48.327 | 117.414 | -0.30004 | 0.101 | 6.377 | 117.414
13.744 | 48.529 | 106.612 | -0.30004 | 0.101 | 6.579 | 106.612
13.744 | 48.731 | 96.508 | -0.30004 | 0.0995 | 6.781 | 96.508
13.744 | 48.93 | 87.482 | -0.30004 | 0.099 | 6.98 | 87.482
13.744 | 49.128 | 78.944 | -0.30003 | 0.1005 | 7.178 | 78.944
13.744 | 49.329 | 71.518 | -0.30003 | 0.1005 | 7.379 | 71.518
13.744 | 49.53 | 64.411 | -0.30003 | 0.0985 | 7.58 | 64.411
13.744 | 49.727 | 58.504 | -0.30004 | 0.0995 | 7.777 | 58.504
13.744 | 49.926 | 52.858 | -0.30003 | 0.1005 | 7.976 | 52.858
13.744 | 50.127 | 47.968 | -0.30003 | 0.101 | 8.177 | 47.968
13.744 | 50.329 | 43.477 | -0.30002 | 0.0995 | 8.379 | 43.477
13.744 | 50.528 | 39.347 | -0.30001 | 0.0995 | 8.578 | 39.347
13.744 | 50.727 | 36.148 | -0.30003 | 0.101 | 8.777 | 36.148
13.744 | 50.929 | 32.501 | -0.30004 | 0.101 | 8.979 | 32.501
13.744 | 51.131 | 29.806 | -0.30004 | 0.0995 | 9.181 | 29.806
13.744 | 51.33 | 27.231 | -0.30004 | 0.099 | 9.38 | 27.231
13.744 | 51.528 | 25.187 | -0.30004 | 0.1005 | 9.578 | 25.187
13.744 | 51.729 | 23.347 | -0.30005 | 0.1005 | 9.779 | 23.347
13.744 | 51.93 | 20.934 | -0.30004 | 0.2485 | 9.98 | 20.934
13.744 | 52.427 | 17.317 | -0.30004 | 0.25 | 10.477 | 17.317
13.744 | 52.927 | 14.617 | -0.30004 | 0.251 | 10.977 | 14.617
13.744 | 53.429 | 12.077 | -0.30004 | 0.2495 | 11.479 | 12.077
13.744 | 53.928 | 10.414 | -0.30004 | 0.2515 | 11.978 | 10.414
13.744 | 54.431 | 8.822 | -0.30005 | 0.2485 | 12.481 | 8.822
13.744 | 54.928 | 8.027 | -0.30004 | 0.249 | 12.978 | 8.027
13.744 | 55.426 | 6.631 | -0.30003 | 0.2525 | 13.476 | 6.631
13.744 | 55.931 | 6.12 | -0.30003 | 0.249 | 13.981 | 6.12
13.744 | 56.429 | 5.471 | -0.30003 | 0.2505 | 14.479 | 5.471
13.744 | 56.93 | 4.886 | -0.30004 | 0.2495 | 14.98 | 4.886
13.744 | 57.429 | 4.364 | -0.30004 | 0.2485 | 15.479 | 4.364
13.744 | 57.926 | 3.92 | -0.30004 | 0.2525 | 15.976 | 3.92
13.744 | 58.431 | 3.544 | -0.30002 | 0.2495 | 16.481 | 3.544
13.744 | 58.93 | 3.206 | -0.30003 | 0.249 | 16.98 | 3.206
13.744 | 59.428 | 2.922 | -0.30004 | 0.251 | 17.478 | 2.922
13.744 | 59.93 | 2.67 | -0.30003 | 0.2485 | 17.98 | 2.67
13.744 | 60.427 | 2.454 | -0.30004 | 0.2515 | 18.477 | 2.454
13.744 | 60.93 | 2.266 | -0.30004 | 0.25 | 18.98 | 2.266
13.744 | 61.43 | 2.092 | -0.30003 | 0.249 | 19.48 | 2.092
13.744 | 61.928 | 1.945 | -0.30004 | 0.251 | 19.978 | 1.945
13.744 | 62.43 | 1.806 | -0.30003 | 0.2485 | 20.48 | 1.806
13.744 | 62.927 | 1.687 | -0.30002 | 0.25 | 20.977 | 1.687
13.744 | 63.427 | 1.584 | -0.30004 | 0.2515 | 21.477 | 1.584
13.744 | 63.93 | 1.486 | -0.30002 | 0.249 | 21.98 | 1.486
13.744 | 64.428 | 1.398 | -0.30003 | 0.2505 | 22.478 | 1.398
13.744 | 64.929 | 1.316 | -0.30004 | 0.2495 | 22.979 | 1.316
13.744 | 65.428 | 1.242 | -0.30004 | 0.249 | 23.478 | 1.242
13.744 | 65.926 | 1.179 | -0.30003 | 0.252 | 23.976 | 1.179
13.744 | 66.43 | 1.118 | -0.30005 | 0.2495 | 24.48 | 1.118
13.744 | 66.929 | 1.066 | -0.30004 | 0.25 | 24.979 | 1.066

## SW Run -4A
! begin File Header |  |  |  |  | ZcL | 41.95
MFX38mm002015.stp | ! data file
2020-08-11 | ! date
14:31 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
161 | ! Z number of points
2 | ! number of devices
0 | ! hall probe 1 offset
Note: Integral, SW, Hyst x 1 (+/-4A), I=-4A
! end File Header |  |  |  |  | I | -3.9987 | A
! begin Step Header |  |  |  |  | BdL | 2702.3001 | g-cm
0 | ! set current |  |  |  | B_center | 466.778 | g
0 | ! actual current at start of step |  |  |  | Eff_L | 5.7893 | cm
0 | ! actual current at end of step
! end Step Header
Position (X, Z) |  | Hall Probe #1 | Voltmeter
13.744 | 16.939 | 1.256 | -0.39987 | 0.2465 | -25.011 | 1.256
13.744 | 17.432 | 1.104 | -0.39987 | 0.249 | -24.518 | 1.104
13.744 | 17.93 | 1.106 | -0.39987 | 0.2495 | -24.02 | 1.106
13.744 | 18.429 | 1.163 | -0.39987 | 0.25 | -23.521 | 1.163
13.744 | 18.929 | 1.251 | -0.39987 | 0.248 | -23.021 | 1.251
13.744 | 19.425 | 1.339 | -0.39987 | 0.2515 | -22.525 | 1.339
13.744 | 19.928 | 1.442 | -0.39987 | 0.25 | -22.022 | 1.442
13.744 | 20.428 | 1.558 | -0.39988 | 0.2495 | -21.522 | 1.558
13.744 | 20.927 | 1.681 | -0.39987 | 0.251 | -21.023 | 1.681
13.744 | 21.429 | 1.817 | -0.39987 | 0.2485 | -20.521 | 1.817
13.744 | 21.926 | 1.968 | -0.39986 | 0.2505 | -20.024 | 1.968
13.744 | 22.427 | 2.14 | -0.39987 | 0.251 | -19.523 | 2.14
13.744 | 22.929 | 2.321 | -0.39987 | 0.249 | -19.021 | 2.321
13.744 | 23.427 | 2.534 | -0.39986 | 0.251 | -18.523 | 2.534
13.744 | 23.929 | 2.778 | -0.39987 | 0.2495 | -18.021 | 2.778
13.744 | 24.428 | 3.05 | -0.39987 | 0.2485 | -17.522 | 3.05
13.744 | 24.925 | 3.368 | -0.39988 | 0.2525 | -17.025 | 3.368
13.744 | 25.43 | 3.714 | -0.39987 | 0.249 | -16.52 | 3.714
13.744 | 25.928 | 4.117 | -0.39987 | 0.2505 | -16.022 | 4.117
13.744 | 26.429 | 4.586 | -0.39986 | 0.25 | -15.521 | 4.586
13.744 | 26.929 | 5.133 | -0.39987 | 0.248 | -15.021 | 5.133
13.744 | 27.425 | 5.77 | -0.39986 | 0.252 | -14.525 | 5.77
13.744 | 27.929 | 6.537 | -0.39987 | 0.25 | -14.021 | 6.537
13.744 | 28.429 | 7.948 | -0.39986 | 0.249 | -13.521 | 7.948
13.744 | 28.927 | 8.688 | -0.39988 | 0.2505 | -13.023 | 8.688
13.744 | 29.428 | 10.327 | -0.39986 | 0.2485 | -12.522 | 10.327
13.744 | 29.925 | 12.087 | -0.39987 | 0.2515 | -12.025 | 12.087
13.744 | 30.428 | 14.671 | -0.39987 | 0.251 | -11.522 | 14.671
13.744 | 30.93 | 17.336 | -0.39988 | 0.2485 | -11.02 | 17.336
13.744 | 31.427 | 21.015 | -0.39987 | 0.251 | -10.523 | 21.015
13.744 | 31.929 | 25.527 | -0.39987 | 0.098 | -10.021 | 25.527
13.744 | 32.125 | 27.626 | -0.39987 | 0.1 | -9.825 | 27.626
13.744 | 32.325 | 30.165 | -0.39987 | 0.101 | -9.625 | 30.165
13.744 | 32.527 | 33.134 | -0.39986 | 0.1 | -9.423 | 33.134
13.744 | 32.727 | 36.124 | -0.39986 | 0.0995 | -9.223 | 36.124
13.744 | 32.926 | 39.382 | -0.39987 | 0.1005 | -9.024 | 39.382
13.744 | 33.127 | 43.604 | -0.39987 | 0.1015 | -8.823 | 43.604
13.744 | 33.33 | 47.786 | -0.39987 | 0.1005 | -8.62 | 47.786
13.744 | 33.531 | 52.808 | -0.39987 | 0.0985 | -8.419 | 52.808
13.744 | 33.728 | 58.272 | -0.39987 | 0.1 | -8.222 | 58.272
13.744 | 33.928 | 64.086 | -0.39987 | 0.1005 | -8.022 | 64.086
13.744 | 34.129 | 70.791 | -0.39987 | 0.1 | -7.821 | 70.791
13.744 | 34.329 | 78.343 | -0.39987 | 0.098 | -7.621 | 78.343
13.744 | 34.525 | 86.534 | -0.39986 | 0.1 | -7.425 | 86.534
13.744 | 34.725 | 95.681 | -0.39987 | 0.1005 | -7.225 | 95.681
13.744 | 34.926 | 106.377 | -0.39986 | 0.1005 | -7.024 | 106.377
13.744 | 35.127 | 117.299 | -0.39987 | 0.099 | -6.823 | 117.299
13.744 | 35.325 | 129.48 | -0.39987 | 0.1005 | -6.625 | 129.48
13.744 | 35.526 | 143.477 | -0.39987 | 0.1015 | -6.424 | 143.477
13.744 | 35.729 | 158.066 | -0.39988 | 0.1005 | -6.221 | 158.066
13.744 | 35.93 | 173.452 | -0.39987 | 0.099 | -6.02 | 173.452
13.744 | 36.128 | 189.238 | -0.39987 | 0.1 | -5.822 | 189.238
13.744 | 36.328 | 206.681 | -0.39988 | 0.1005 | -5.622 | 206.681
13.744 | 36.529 | 224.667 | -0.39987 | 0.1 | -5.421 | 224.667
13.744 | 36.729 | 242.446 | -0.39987 | 0.0985 | -5.221 | 242.446
13.744 | 36.926 | 260.992 | -0.39986 | 0.0995 | -5.024 | 260.992
13.744 | 37.125 | 279.436 | -0.39986 | 0.101 | -4.825 | 279.436
13.744 | 37.327 | 297.41 | -0.39985 | 0.0995 | -4.623 | 297.41
13.744 | 37.526 | 315.228 | -0.39987 | 0.099 | -4.424 | 315.228
13.744 | 37.724 | 332.601 | -0.39987 | 0.1005 | -4.226 | 332.601
13.744 | 37.925 | 349.241 | -0.39986 | 0.101 | -4.025 | 349.241
13.744 | 38.127 | 364.705 | -0.39986 | 0.101 | -3.823 | 364.705
13.744 | 38.329 | 379.078 | -0.39986 | 0.099 | -3.621 | 379.078
13.744 | 38.527 | 392.114 | -0.39988 | 0.1 | -3.423 | 392.114
13.744 | 38.727 | 404.146 | -0.39986 | 0.101 | -3.223 | 404.146
13.744 | 38.929 | 414.608 | -0.39986 | 0.1 | -3.021 | 414.608
13.744 | 39.129 | 423.988 | -0.39985 | 0.0985 | -2.821 | 423.988
13.744 | 39.326 | 432.102 | -0.39987 | 0.0995 | -2.624 | 432.102
13.744 | 39.525 | 439.644 | -0.39986 | 0.1005 | -2.425 | 439.644
13.744 | 39.726 | 445.554 | -0.39987 | 0.1 | -2.224 | 445.554
13.744 | 39.926 | 450.66 | -0.39987 | 0.0985 | -2.024 | 450.66
13.744 | 40.123 | 454.701 | -0.39987 | 0.1005 | -1.827 | 454.701
13.744 | 40.324 | 457.85 | -0.39987 | 0.1015 | -1.626 | 457.85
13.744 | 40.527 | 460.668 | -0.39987 | 0.101 | -1.423 | 460.668
13.744 | 40.729 | 462.283 | -0.39987 | 0.099 | -1.221 | 462.283
13.744 | 40.927 | 463.918 | -0.39987 | 0.1005 | -1.023 | 463.918
13.744 | 41.128 | 465.218 | -0.39986 | 0.1005 | -0.822 | 465.218
13.744 | 41.329 | 465.62 | -0.39987 | 0.1005 | -0.621 | 465.62
13.744 | 41.53 | 466.078 | -0.39987 | 0.099 | -0.42 | 466.078
13.744 | 41.728 | 466.531 | -0.39987 | 0.099 | -0.222 | 466.531
13.744 | 41.926 | 466.778 | -0.39987 | 0.1005 | -0.024 | 466.778
13.744 | 42.127 | 466.96 | -0.39987 | 0.1 | 0.177 | 466.96
13.744 | 42.327 | 467.019 | -0.39986 | 0.0985 | 0.377 | 467.019
13.744 | 42.524 | 466.966 | -0.39987 | 0.1005 | 0.574 | 466.966
13.744 | 42.725 | 466.729 | -0.39985 | 0.101 | 0.775 | 466.729
13.744 | 42.927 | 466.347 | -0.39987 | 0.101 | 0.977 | 466.347
13.744 | 43.129 | 465.676 | -0.39986 | 0.0995 | 1.179 | 465.676
13.744 | 43.328 | 463.962 | -0.39987 | 0.0995 | 1.378 | 463.962
13.744 | 43.527 | 462.253 | -0.39987 | 0.101 | 1.577 | 462.253
13.744 | 43.729 | 459.468 | -0.39987 | 0.1005 | 1.779 | 459.468
13.744 | 43.93 | 456.229 | -0.39988 | 0.0995 | 1.98 | 456.229
13.744 | 44.129 | 452.261 | -0.39987 | 0.099 | 2.179 | 452.261
13.744 | 44.327 | 446.961 | -0.39986 | 0.1005 | 2.377 | 446.961
13.744 | 44.528 | 440.657 | -0.39987 | 0.1 | 2.578 | 440.657
13.744 | 44.728 | 433.548 | -0.39987 | 0.0985 | 2.778 | 433.548
13.744 | 44.925 | 424.984 | -0.39987 | 0.0995 | 2.975 | 424.984
13.744 | 45.124 | 415.773 | -0.39986 | 0.1015 | 3.174 | 415.773
13.744 | 45.327 | 404.42 | -0.39986 | 0.101 | 3.377 | 404.42
13.744 | 45.529 | 392.394 | -0.39987 | 0.0995 | 3.579 | 392.394
13.744 | 45.728 | 379.005 | -0.39987 | 0.0995 | 3.778 | 379.005
13.744 | 45.927 | 364.232 | -0.39987 | 0.1015 | 3.977 | 364.232
13.744 | 46.13 | 348.387 | -0.39987 | 0.1005 | 4.18 | 348.387
13.744 | 46.331 | 331.811 | -0.39987 | 0.099 | 4.381 | 331.811
13.744 | 46.529 | 314.669 | -0.39987 | 0.099 | 4.579 | 314.669
13.744 | 46.727 | 296.961 | -0.39987 | 0.1005 | 4.777 | 296.961
13.744 | 46.928 | 278.487 | -0.39986 | 0.1005 | 4.978 | 278.487
13.744 | 47.129 | 259.254 | -0.39988 | 0.0985 | 5.179 | 259.254
13.744 | 47.326 | 240.947 | -0.39986 | 0.099 | 5.376 | 240.947
13.744 | 47.524 | 222.763 | -0.39988 | 0.1015 | 5.574 | 222.763
13.744 | 47.727 | 204.803 | -0.39987 | 0.1005 | 5.777 | 204.803
13.744 | 47.928 | 187.49 | -0.39987 | 0.1 | 5.978 | 187.49
13.744 | 48.128 | 171.341 | -0.39985 | 0.0995 | 6.178 | 171.341
13.744 | 48.327 | 156.457 | -0.39986 | 0.101 | 6.377 | 156.457
13.744 | 48.529 | 141.514 | -0.39986 | 0.101 | 6.579 | 141.514
13.744 | 48.731 | 128.393 | -0.39987 | 0.0995 | 6.781 | 128.393
13.744 | 48.93 | 115.898 | -0.39987 | 0.099 | 6.98 | 115.898
13.744 | 49.128 | 105.122 | -0.39986 | 0.1005 | 7.178 | 105.122
13.744 | 49.329 | 94.907 | -0.39987 | 0.1005 | 7.379 | 94.907
13.744 | 49.53 | 85.761 | -0.39987 | 0.0985 | 7.58 | 85.761
13.744 | 49.727 | 77.916 | -0.39987 | 0.099 | 7.777 | 77.916
13.744 | 49.925 | 70.407 | -0.39987 | 0.101 | 7.975 | 70.407
13.744 | 50.127 | 63.726 | -0.39987 | 0.101 | 8.177 | 63.726
13.744 | 50.329 | 57.857 | -0.39987 | 0.0995 | 8.379 | 57.857
13.744 | 50.528 | 52.389 | -0.39987 | 0.0995 | 8.578 | 52.389
13.744 | 50.727 | 47.322 | -0.39987 | 0.101 | 8.777 | 47.322
13.744 | 50.929 | 43.159 | -0.39986 | 0.1005 | 8.979 | 43.159
13.744 | 51.13 | 39.308 | -0.39987 | 0.0995 | 9.18 | 39.308
13.744 | 51.329 | 36.167 | -0.39986 | 0.099 | 9.379 | 36.167
13.744 | 51.527 | 33.033 | -0.39987 | 0.101 | 9.577 | 33.033
13.744 | 51.729 | 30.285 | -0.39986 | 0.1005 | 9.779 | 30.285
13.744 | 51.93 | 27.825 | -0.39986 | 0.248 | 9.98 | 27.825
13.744 | 52.426 | 22.71 | -0.39987 | 0.2505 | 10.476 | 22.71
13.744 | 52.927 | 18.964 | -0.39987 | 0.251 | 10.977 | 18.964
13.744 | 53.429 | 15.913 | -0.39987 | 0.249 | 11.479 | 15.913
13.744 | 53.927 | 13.274 | -0.39987 | 0.2515 | 11.977 | 13.274
13.744 | 54.43 | 11.593 | -0.39987 | 0.2485 | 12.48 | 11.593
13.744 | 54.927 | 10.017 | -0.39986 | 0.2495 | 12.977 | 10.017
13.744 | 55.426 | 8.617 | -0.39987 | 0.252 | 13.476 | 8.617
13.744 | 55.93 | 7.919 | -0.39988 | 0.2495 | 13.98 | 7.919
13.744 | 56.429 | 7.053 | -0.39988 | 0.2505 | 14.479 | 7.053
13.744 | 56.93 | 6.293 | -0.39987 | 0.2495 | 14.98 | 6.293
13.744 | 57.429 | 5.607 | -0.39986 | 0.2485 | 15.479 | 5.607
13.744 | 57.926 | 5.026 | -0.39988 | 0.252 | 15.976 | 5.026
13.744 | 58.43 | 4.518 | -0.39987 | 0.25 | 16.48 | 4.518
13.744 | 58.93 | 4.09 | -0.39988 | 0.249 | 16.98 | 4.09
13.744 | 59.428 | 3.714 | -0.39987 | 0.251 | 17.478 | 3.714
13.744 | 59.93 | 3.391 | -0.39988 | 0.248 | 17.98 | 3.391
13.744 | 60.426 | 3.111 | -0.39987 | 0.2515 | 18.476 | 3.111
13.744 | 60.929 | 2.862 | -0.39988 | 0.2505 | 18.979 | 2.862
13.744 | 61.43 | 2.632 | -0.39987 | 0.249 | 19.48 | 2.632
13.744 | 61.928 | 2.441 | -0.39987 | 0.251 | 19.978 | 2.441
13.744 | 62.43 | 2.266 | -0.39987 | 0.2485 | 20.48 | 2.266
13.744 | 62.927 | 2.108 | -0.39987 | 0.25 | 20.977 | 2.108
13.744 | 63.427 | 1.974 | -0.39987 | 0.2515 | 21.477 | 1.974
13.744 | 63.93 | 1.846 | -0.39987 | 0.2485 | 21.98 | 1.846
13.744 | 64.427 | 1.733 | -0.39987 | 0.251 | 22.477 | 1.733
13.744 | 64.929 | 1.628 | -0.39988 | 0.2495 | 22.979 | 1.628
13.744 | 65.428 | 1.534 | -0.39988 | 0.249 | 23.478 | 1.534
13.744 | 65.926 | 1.454 | -0.39987 | 0.252 | 23.976 | 1.454
13.744 | 66.43 | 1.379 | -0.39987 | 0.249 | 24.48 | 1.379
13.744 | 66.928 | 1.305 | -0.39988 | 0.25 | 24.978 | 1.305

## B v I
! begin File Header |  |  |  | Index | Current (A) | Voltage (V) | Field (G)
MFX38mm002017.stp | ! data file |  |  | 0 | -3.9986 | -15.576 | 466.84
2020-08-12 | ! date |  |  | 16 | -3.0002 | -11.655 | 353.18
8:26 | ! start time |  |  | 32 | -2.7991 | -10.871 | 329.99
5.1 | ! stepper code version |  |  | 48 | -1.9995 | -7.756 | 237.27
1 | ! number Z of currents |  |  | 64 | -0.99876 | -3.869 | 120.31
1 | ! X number of points |  |  | 80 | -0.00106 | 0 | 3.14
1 | ! Z number of points |  |  | 96 | 0.99896 | 3.91 | -114.61
2 | ! number of devices |  |  | 112 | 2.0005 | 7.83 | -232.62
41.95 | ! position [z] |  |  | 128 | 2.7995 | 10.958 | -326.73
13.75 | ! position [x] |  |  | 144 | 3.0007 | 11.746 | -350.39
0 | ! hall probe 1 offset |  |  | 160 | 3.9995 | 15.61 | -467.79
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets |  |  |  | 176 | 3.0008 | 11.746 | -354.05
! end File Header |  |  |  | 192 | 2.7995 | 10.958 | -330.82
! set curr | actual curr | Hall Probe #1 | Voltmeter | 208 | 2.0006 | 7.83 | -238.17
0 | 0 | 466.84 | -0.39986 | 224 | 0.99885 | 3.91 | -121.12
! begin File Header |  |  |  | 240 | -0.00121 | 0 | -3.68
MFX38mm002017.stp | ! data file |  |  | 256 | -0.99882 | -3.869 | 113.82
2020-08-12 | ! date |  |  | 272 | -1.9994 | -7.756 | 231.75
8:26 | ! start time |  |  | 288 | -2.7992 | -10.871 | 325.9
5.1 | ! stepper code version |  |  | 304 | -3.0002 | -11.655 | 349.54
1 | ! number Z of currents |  |  | 320 | -3.9985 | -15.576 | 466.81
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 353.18 | -0.30002
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:27 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 329.99 | -0.27991
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:27 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 237.27 | -0.19995
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:27 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 120.31 | -0.099876
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:28 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 3.14 | -0.000106
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:28 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -114.61 | 0.099896
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:28 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -232.62 | 0.20005
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:29 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -326.73 | 0.27995
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:29 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -350.39 | 0.30007
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:29 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -467.79 | 0.39995
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:30 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -354.05 | 0.30008
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:30 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -330.82 | 0.27995
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:30 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -238.17 | 0.20006
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:31 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -121.12 | 0.099885
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:31 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -3.68 | -0.000121
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:31 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 113.82 | -0.099882
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:32 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 231.75 | -0.19994
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:32 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 325.9 | -0.27992
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:32 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 349.54 | -0.30002
! begin File Header
MFX38mm002017.stp | ! data file
2020-08-12 | ! date
8:33 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), 1A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 466.81 | -0.39985

## B v I (Null .2A inc)
! begin File Header |  |  |  | Index | Current (A) | Field (G) | Voltage (V)
MFX38mm002020.stp | ! data file |  |  | 0 | 3.9995 | -299.11 | 16.38
2020-08-12 | ! date |  |  | 16 | 3.7977 | -285.26 | 15.548
11:06 | ! start time |  |  | 32 | 3.5991 | -271.27 | 14.731
5.1 | ! stepper code version |  |  | 48 | 3.3998 | -257.1 | 13.914
1 | ! number Z of currents |  |  | 64 | 3.2005 | -242.77 | 13.096
1 | ! X number of points |  |  | 80 | 3.0007 | -228.29 | 12.28
1 | ! Z number of points |  |  | 96 | 2.7996 | -213.61 | 11.455
2 | ! number of devices |  |  | 112 | 2.5979 | -198.82 | 10.632
38.95 | ! position [z] |  |  | 128 | 2.3988 | -184.15 | 9.818
13.75 | ! position [x] |  |  | 144 | 2.1996 | -169.35 | 9.005
0 | ! hall probe 1 offset |  |  | 160 | 2.0003 | -154.55 | 8.19
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments |  |  |  | 176 | 1.8008 | -139.68 | 7.374
! end File Header |  |  |  | 192 | 1.5989 | -124.59 | 6.549
! set curr | actual curr | Hall Probe #1 | Voltmeter | 208 | 1.3971 | -109.49 | 5.724
0 | 0 | -299.11 | 0.39995 | 224 | 1.197 | -94.44 | 4.908
! begin File Header |  |  |  | 240 | 0.99874 | -79.55 | 4.097
MFX38mm002020.stp | ! data file |  |  | 256 | 0.79993 | -64.57 | 3.283
2020-08-12 | ! date |  |  | 272 | 0.60184 | -49.64 | 2.474
11:06 | ! start time |  |  | 288 | 0.40074 | -34.46 | 1.65
5.1 | ! stepper code version |  |  | 304 | 0.19941 | -19.26 | 0.825
1 | ! number Z of currents |  |  | 320 | -0.00128 | -4.09 | 0
1 | ! X number of points |  |  | 336 | -0.20031 | 11 | -0.812
1 | ! Z number of points |  |  | 352 | -0.39932 | 26.08 | -1.628
2 | ! number of devices |  |  | 368 | -0.59883 | 41.21 | -2.448
38.95 | ! position [z] |  |  | 384 | -0.79893 | 56.37 | -3.269
13.75 | ! position [x] |  |  | 400 | -0.9989 | 71.53 | -4.092
0 | ! hall probe 1 offset |  |  | 416 | -1.1989 | 86.69 | -4.915
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments |  |  |  | 432 | -1.3993 | 101.91 | -5.74
! end File Header |  |  |  | 448 | -1.5992 | 117.08 | -6.562
! set curr | actual curr | Hall Probe #1 | Voltmeter | 464 | -1.7994 | 132.28 | -7.388
0 | 0 | -285.26 | 0.37977 | 480 | -1.9996 | 147.45 | -8.213
! begin File Header |  |  |  | 496 | -2.1989 | 162.55 | -9.035
MFX38mm002020.stp | ! data file |  |  | 512 | -2.3988 | 177.68 | -9.86
2020-08-12 | ! date |  |  | 528 | -2.5987 | 192.84 | -10.686
11:07 | ! start time |  |  | 544 | -2.7992 | 208.02 | -11.512
5.1 | ! stepper code version |  |  | 560 | -3.0004 | 223.23 | -12.341
1 | ! number Z of currents |  |  | 576 | -3.201 | 238.38 | -13.167
1 | ! X number of points |  |  | 592 | -3.4001 | 253.43 | -13.987
1 | ! Z number of points |  |  | 608 | -3.5998 | 268.48 | -14.808
2 | ! number of devices |  |  | 624 | -3.7995 | 283.56 | -15.628
38.95 | ! position [z] |  |  | 640 | -3.9988 | 298.58 | -16.445
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -271.27 | 0.35991
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:07 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -257.1 | 0.33998
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:07 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -242.77 | 0.32005
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:08 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -228.29 | 0.30007
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:08 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -213.61 | 0.27996
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:08 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -198.82 | 0.25979
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:09 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -184.15 | 0.23988
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:09 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -169.35 | 0.21996
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:09 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -154.55 | 0.20003
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:10 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -139.68 | 0.18008
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:10 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -124.59 | 0.15989
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:10 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -109.49 | 0.13971
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:11 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -94.44 | 0.1197
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:11 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -79.55 | 0.099874
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:11 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -64.57 | 0.079993
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:12 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -49.64 | 0.060184
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:12 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -34.46 | 0.040074
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:12 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -19.26 | 0.019941
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:12 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -4.09 | -0.000128
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:13 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 11 | -0.020031
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:13 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 26.08 | -0.039932
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:13 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 41.21 | -0.059883
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:14 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 56.37 | -0.079893
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:14 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 71.53 | -0.09989
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:14 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 86.69 | -0.11989
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:15 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 101.91 | -0.13993
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:15 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 117.08 | -0.15992
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:15 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 132.28 | -0.17994
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:16 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 147.45 | -0.19996
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:16 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 162.55 | -0.21989
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:16 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 177.68 | -0.23988
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:17 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 192.84 | -0.25987
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:17 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 208.02 | -0.27992
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:18 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 223.23 | -0.30004
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:18 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 238.38 | -0.3201
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:18 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 253.43 | -0.34001
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:19 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 268.48 | -0.35998
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:19 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 283.56 | -0.37995
! begin File Header
MFX38mm002020.stp | ! data file
2020-08-12 | ! date
11:19 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
38.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, CW/Bucked, Hyst x 4 (+/-4A) .2A increments
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 298.58 | -0.39988

## B v I (SW .2A inc)
! begin File Header |  |  |  | Index | Current (A) | Field (G) | Voltage (V)
MFX38mm002016.stp | ! data file |  |  | 0 | 3.9998 | -467.63 | 15.61
2020-08-12 | ! date |  |  | 16 | 3.7979 | -445.18 | 14.85
7:33 | ! start time |  |  | 32 | 3.5992 | -422.63 | 14.077
5.1 | ! stepper code version |  |  | 48 | 3.4001 | -399.89 | 13.302
1 | ! number Z of currents |  |  | 64 | 3.2006 | -377.02 | 12.525
1 | ! X number of points |  |  | 80 | 3.0008 | -354.08 | 11.746
1 | ! Z number of points |  |  | 96 | 2.7995 | -330.85 | 10.958
2 | ! number of devices |  |  | 112 | 2.5983 | -307.57 | 10.17
41.95 | ! position [z] |  |  | 128 | 2.399 | -284.5 | 9.39
13.75 | ! position [x] |  |  | 144 | 2.2001 | -261.42 | 8.611
0 | ! hall probe 1 offset |  |  | 160 | 2.0006 | -238.19 | 7.83
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets |  |  |  | 176 | 1.801 | -214.94 | 7.048
! end File Header |  |  |  | 192 | 1.599 | -191.38 | 6.258
! set curr | actual curr | Hall Probe #1 | Voltmeter | 208 | 1.3972 | -167.82 | 5.468
0 | 0 | -467.63 | 0.39998 | 224 | 1.1972 | -144.38 | 4.684
! begin File Header |  |  |  | 240 | 0.99905 | -121.19 | 3.91
MFX38mm002016.stp | ! data file |  |  | 256 | 0.80018 | -97.85 | 3.133
2020-08-12 | ! date |  |  | 272 | 0.60214 | -74.62 | 2.36
7:34 | ! start time |  |  | 288 | 0.40111 | -50.99 | 1.575
5.1 | ! stepper code version |  |  | 304 | 0.19967 | -27.34 | 0.79
1 | ! number Z of currents |  |  | 320 | -0.00101 | -3.74 | 0
1 | ! X number of points |  |  | 336 | -0.20016 | 19.69 | -0.766
1 | ! Z number of points |  |  | 352 | -0.39908 | 43.11 | -1.54
2 | ! number of devices |  |  | 368 | -0.59874 | 66.64 | -2.315
41.95 | ! position [z] |  |  | 384 | -0.79875 | 90.22 | -3.093
13.75 | ! position [x] |  |  | 400 | -0.99876 | 113.78 | -3.869
0 | ! hall probe 1 offset |  |  | 416 | -1.1988 | 137.35 | -4.645
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets |  |  |  | 432 | -1.3991 | 160.96 | -5.423
! end File Header |  |  |  | 448 | -1.5989 | 184.51 | -6.199
! set curr | actual curr | Hall Probe #1 | Voltmeter | 464 | -1.7993 | 208.15 | -6.979
0 | 0 | -445.18 | 0.37979 | 480 | -1.9995 | 231.74 | -7.756
! begin File Header |  |  |  | 496 | -2.1988 | 255.2 | -8.53
MFX38mm002016.stp | ! data file |  |  | 512 | -2.3987 | 278.74 | -9.308
2020-08-12 | ! date |  |  | 528 | -2.5985 | 302.27 | -10.086
7:34 | ! start time |  |  | 544 | -2.7991 | 325.87 | -10.871
5.1 | ! stepper code version |  |  | 560 | -3.0003 | 349.56 | -11.655
1 | ! number Z of currents |  |  | 576 | -3.2008 | 373.14 | -12.441
1 | ! X number of points |  |  | 592 | -3.3998 | 396.53 | -13.22
1 | ! Z number of points |  |  | 608 | -3.5996 | 419.99 | -14.004
2 | ! number of devices |  |  | 624 | -3.7993 | 443.43 | -14.792
41.95 | ! position [z] |  |  | 640 | -3.9986 | 466.84 | -15.576
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -422.63 | 0.35992
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:35 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -399.89 | 0.34001
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:35 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -377.02 | 0.32006
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:36 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -354.08 | 0.30008
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:36 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -330.85 | 0.27995
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:36 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -307.57 | 0.25983
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:37 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -284.5 | 0.2399
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:37 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -261.42 | 0.22001
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:38 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -238.19 | 0.20006
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:38 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -214.94 | 0.1801
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:39 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -191.38 | 0.1599
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:39 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -167.82 | 0.13972
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:39 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -144.38 | 0.11972
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:40 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -121.19 | 0.099905
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:40 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -97.85 | 0.080018
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:41 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -74.62 | 0.060214
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:41 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -50.99 | 0.040111
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:41 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -27.34 | 0.019967
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:42 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | -3.74 | -0.000101
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:43 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 19.69 | -0.020016
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:43 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 43.11 | -0.039908
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:44 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 66.64 | -0.059874
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:45 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 90.22 | -0.079875
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:45 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 113.78 | -0.099876
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:46 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 137.35 | -0.11988
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:46 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 160.96 | -0.13991
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:46 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 184.51 | -0.15989
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:47 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 208.15 | -0.17993
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:47 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 231.74 | -0.19995
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:48 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 255.2 | -0.21988
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:48 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 278.74 | -0.23987
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:48 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 302.27 | -0.25985
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:49 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 325.87 | -0.27991
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:49 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 349.56 | -0.30003
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:49 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 373.14 | -0.32008
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 396.53 | -0.33998
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 419.99 | -0.35996
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:50 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 443.43 | -0.37993
! begin File Header
MFX38mm002016.stp | ! data file
2020-08-12 | ! date
7:51 | ! start time
5.1 | ! stepper code version
1 | ! number Z of currents
1 | ! X number of points
1 | ! Z number of points
2 | ! number of devices
41.95 | ! position [z]
13.75 | ! position [x]
0 | ! hall probe 1 offset
Note: B v I, SW, Hyst x 4 (+/-4A), .2A incremets
! end File Header
! set curr | actual curr | Hall Probe #1 | Voltmeter
0 | 0 | 466.84 | -0.39986