## Sheet1
Apperture A1
Coordiate System: | The coordinate system is based on:
Z axis is defined as the line thorugh thecenter of the US and DS Flanges
Roll is controlled by the top of the cube
Origin was established using the intersection between the bisected plane of US and DS Flange and the z axis.
| Name | X[mm] | Y[mm] | Z[mm]
| US_FLANGE | 0 | 0 | -74.76
| DS_FLANGE | 0 | 0 | 74.76
At atmosphere: | Name | X[mm] | Y[mm]
| 2mm | 0.21 | 0.26
| 3mm | 0.47 | 0.17
| 4mm | 0.65 | 0.23
Measured values under vacuum : | Name | X[mm] | Y[mm]
| 2mm | -0.29 | 0.08
| 3mm | -0.03 | -0.01
| 4mm | 0.15 | 0.05
The dimensions shown above are based on the test performed under vacuum, showing the translation of 0.500mm to the right( in x) and 0.180mm down (in y).
Fiducial Locations | Name | X[mm] | Y[mm] | Z[mm]
| FIDA | 50.84 | 65 | -51.11
| FIDB | 50.97 | 65.19 | 50.42
| FIDC | -50.67 | 65.25 | 50.51
| FIDD | -50.76 | 65.04 | -51.02
| FIDE | 50.83 | 50.47 | -65.26
| FIDF | -50.76 | 50.55 | -65.19
| FIDG | -50.78 | -51.06 | -65.05
| FIDH | 50.87 | -51.12 | -65.1
| FIDJ | -50.63 | 50.94 | 64.99
| FIDK | 50.97 | 50.95 | 64.92
| FIDL | 50.96 | -50.63 | 65.11
| FIDM | -50.69 | -50.63 | 65.14
Spacer Dimensions (under vacuum) | Name | Dim. [mm] | Joe diff from Home | Scott reading scale | Scott Encoder (steps)
| Home | 100.49 | 0
| 2mm | 27.29 | 73.2 | "74mm" | -1466800
| 3mm | 37.16 | 63.33 | "64mm" | -1268800
| 4mm | 47.62 | 52.87 | "54mm" | -1065000
Apperture A1 vertical postion is obtained using an electric motor. Encoder values were obtained by Scott Windham. The values shown above are as found values. If it becomes necessary to adjust vertical position by hand, 0.180 mm should be added to these values.