## Slide 1: Helicity Decoder Board
Thanks to:
Chris Cuevas, Ed Jastrzembski, Bryan Moffit, and Paul King
Riad Suleiman

## Slide 2: Helicity Generator Board
Now:
  Free Clock Mode at 30 Hz Default, 500 µs T_Settle, 33330 µs T_Stable
  Quartet pattern, and 8-window delay.
New Helicity settings for MOLLER:
    Free Clock Mode at 1920 Hz, 10 µs T_Settle, 510.85 µs T_Stable
    64-window patterns: Thue-Morse-64, 16-Quad, 32-Pair
Board is supported by Jefferson Lab Fast Electronics Group

## Slide 3: Helicity Generator Settings
1
30
29
28
27
…
7
…
1
Switch
Open: Pseudo-random Patterns
Closed: Toggle Pattern
30-Bit Shift Register
Pseudo-random Bit out
Pseudo-random Bit
(0 or 1)
Shift Left
OR
XOR
XOR
XOR

## Slide 4: Helicity Decoder Board
MOLLER fast helicity reversal with delay reporting could cause mis-identification of real helicity of some of events in counting mode (these are random events), diluting measured asymmetry
Board developed by: MOLLER Collaboration (Ohio University), Jefferson Lab Fast Electronics Group and Accelerator Division
Will be used by all data acquisition systems using helicity signals (parity and non-parity)
Board will not reconstruct real helicity or real 30-Bit Shift Register
Board has four input signals: T_Settle, Delayed Helicity, Helicity Pattern, and Pair Sync
For each counting event:
  Now: Delayed Helicity and Pattern are recorded in Rings of helicity-gated SIS3801 Scaler. Offline analysis then attempts to construct real helicity, a complicated process due to delay and randomness of events.
  New Board: board will keep records of helicities spanning last 30 patterns – to be able to construct 30-bit shift register seed of pseudo-random helicity generator – and position within pattern. And for each event, these records are read along with event. Offline analysis (anytime>n-patterns) will calculate real helicity using helicity predictor. Will be able to correctly identify helicity of every event.

## Slide 5: Electron Beam Travel Time
With 10 µs T_Settle time, travel time from photocathode (Pockels Cell) to Hall target becomes relevant. Also, time it takes for helicity board signals to propagate to Laser Hut and to Halls.
It takes electron beam about 4.3 µs per pass to reach Hall
Total distance at 5 passes is about 6.5 km
Fiber length from Injector Service Building (ISB) IN01B05 to Counting House was measured to be 1705 ft. With n=1.45, it takes 2.5 µs for helicity signals to travel in fibers to Hall DAQ
  All these travel times will be accounted for in this new Helicity Decoder Board such that recorded events have correct helicity at Physics Interaction Time
Without Helicity Decoder: Experiment deadtime is (4.3 per pass + 10.0)/510.85 instead of 10.0/510.85
Built 20 Production Boards, tested and ready for distribution:
  Five in use: 2 MOLLER, 1 Hall B, 1 Hall D, 1 Injector Mott Polarimeter
  Contact Chris Cuevas to get your board

## Slide 6: Decoder Board I/O and Features
Input:
  T_Settle
  Delayed Helicity
  Helicity Pattern
  Pair Sync
  External 250 MHz FADC Clock (ECL) – Board has 250 MHz internal clock
  Trigger and Sync Signals
Output:
  Four NIM helicity outputs: T_Settle, Delayed, Pattern and Pair Sync
  Data words for each trigger event
Add programable common delay (0-32 µs) for input helicity signals:
  To fix time delays due to beam travel time and helicity signal distribution to Hall
  Will be measured for each board in a specific DAQ
  Will sync helicity of beam at Physics Interaction Time in Hall
NIM or Fiber, programmable

## Slide 7: Board Programmable Common Delay
How to set input common time delay:
  Helicity Board: 60 Hz Beam Sync (354 µs)
  Electron Beam: VL, 60 Hz Beam Sync, 50 ns pulses
  First, check relative timing of electron VL Generator and Helicity signal on laser table
  Adjust board common delay to match relative timing at laser table
VL
Generator
RTP
Cell
Delayed
Helicity
Fiber
HV Supply
CEBAF
Hall
Target
DAQ
Helicity Generator
ISB
60 Hz Beam Sync
Real
Helicity
Fiber
Check relative timing
Helicity Decoder Board
Check relative timing
Trigger
Board Common Input Delay (0-32 µs)