## Slide 1: Viewer resistance measurements
Marcy Stutzman, Oct 18 2022

## Slide 2: Viewer induced beam motion
Harp and viewer both installed at IHA0I02 location
  Harp typically does not share the cross with a viewer
Harp does not completely retract from cross
  Crashed top hat of viewer into harp
YAG coated with conductive ITO coating, 90 nm thick, to bleed charge
Removed top hat at ITV0I02

## Slide 3: Evidence for charge/discharge beam motion

## Slide 4: 4 point resistance measurements
Contact resistance dominates for conductive films
Use 4 point measurement technique

## Slide 5: Results
Broken sample – old viewer from GTS, no coating
Resistance matches that of open probe (∞)
Chart: IV curves for surface resistivity
| 1.3224924961538461E-2 | -2.8157791714285723E-3 | -5.6214898461538453E-3
open | 0.01 | 0.001 | 5e-06
broken 1 | 0.01 | 0.001 | 5e-06
broken 2 | 0.01 | 0.001 | 5e-06
broken 3 | 0.01 | 0.001 | 5e-06

## Slide 6: Results 2
New viewer with fiducials, ITO coating
Sheet resistance
Rs = π/ln(2) *V/I (ohms/square)
528.5 ohms/square
Viewer should discharge any beam incident on it
Chart: IV curves for surface resistivity
| 1.3224924961538461E-2 | -2.8157791714285723E-3 | -5.6214898461538453E-3
open | 0.01 | 0.001 | 5e-06
broken 1 | 0.01 | 0.001 | 5e-06
broken 2 | 0.01 | 0.001 | 5e-06
broken 3 | 0.01 | 0.001 | 5e-06
new screen | 0.01 | 0.001 | 5e-06

## Slide 7: Solution
Adjust harp retracted position
  White: schematic of current
  Red: new fully retracted position
Install viewer with top hat again
Puzzle over why conductive viewer is charging