## Slide 1: Electron Beam Setup for Bubble Chamber
Hofler, Accelerator-Physics Experiment Liaison (APEL)
J. Benesch, B. Cade, A. Freyberger, J. Grames, R. Kazimi, M. McCaughn, D. Moser, M. Poelker, & R. Suleiman

## Slide 2: Improvements since 2015 Engineering Run
Dipole
  More accurate dipole field map [1,2]
  Dipole current readback resolution
    Readback with Digital Multimeter (Keithley 2000)
    Higher resolution (1 mA) than trim card readback (5 mA)
  Hall probe in dipole
    Measure field in dipole
    Digital Teslameter (Group3 Technology DTM-151)
    MPT-231 Hall probe (Group3 Technology MPT-231)
  Stray magnetic field shielding in 5D line [3]
Beam diagnostics
  Wire scanner beam profile monitor (harp) in 5D line
    Measure beam size upstream of radiator
[1] R. Suleiman, BTeam presentation, 2016
[2] J. Benesch, JLAB-TN-017, 2015
[3] R. Suleiman, e-log 3392568, 2016
[4] Bubble Chamber wiki
[4] (http://wiki.jlab.org/ciswiki/index.php/Bubble_Chamber)
2D line
Straight ahead line
Dipole
Harp location
5D line
for Bubble Chamber Experiment (3.63 m from dipole to center of chamber)

## Slide 3: Improvements since 2015 Engineering Run
Dipole
  More accurate dipole field map [1,2]
  Dipole current readback resolution
    Readback with Digital Multimeter (Keithley 2000)
    Higher resolution (1 mA) than trim card readback (5 mA)
  Hall probe in dipole
    Measure field in dipole
    Digital Teslameter (Group3 Technology DTM-151)
    MPT-231 Hall probe (Group3 Technology MPT-231)
  Stray magnetic field shielding in 5D line [3]
Beam diagnostics
  Wire scanner beam profile monitor (harp) in 5D line
    Measure beam size upstream of radiator
Procedures for Operations [4]
  Set energy
    Use dipole Hall probe readback to set momentum delta and reproduce setting
  Measure energy
  Measure energy spread with harp in 2D line
  Set optics to produce requested spot size on radiator
    Quadrupole scan technique
    Wire scanner beam profile monitors (harps) in straight ahead and 2D lines
[1] R. Suleiman, BTeam presentation, 2016
[2] J. Benesch, JLAB-TN-017, 2015
[3] R. Suleiman, e-log 3392568, 2016
[4] Bubble Chamber wiki
[4] (http://wiki.jlab.org/ciswiki/index.php/Bubble_Chamber)

## Slide 4: Energy Measurement
Procedure:
A priori, center beam in quadrupole closest to BPM and note positions (except 2D—no quads).
Turn OFF and cycle all unnecessary magnets around hysteresis loop.
Degauss dipole.
Establish straight ahead orbit with restricted corrector magnet set.
Adjust dipole setting (and cycle) for zero orbit in each dump BPM.
BPM=Beam Position Monitor
Dipole
2D
-30o
5D
25o
BPM
BPM
BPM
BPM
For fixed incoming momentum, = ,
Fits of dipole settings give p and errors in angle and field.
Note dipole setpoint
  < 0 for beam in 2D line
  > 0 for beam in 5D line
Accelerator coordinate system is LH
Beam
Direction

## Slide 5: Measured Energies
Design p (MeV/c) | Design K.E. (MeV) | 2D line dipole setting (G cm) | 5D (Bubble) line dipole setting (G cm) | Measured p (MeV/c) | Measured K.E. (MeV)
5.84 | 5.35 | -9937.637 | 8168.800 | 5.885 | 5.395
5.74 | 5.25 | -9865.500 | 8099.000 | 5.840 | 5.355
5.64 | 5.15 | -9632.300 | 7909.200 | 5.705 | 5.215
5.54 | 5.05 | -9468.500 | 7771.400 | 5.605 | 5.115
5.44 | 4.95 | -9320.700 | 7646.800 | 5.515 | 5.030
5.34 | 4.85 | -9135.993 | 7490.000 | 5.405 | 4.920
5.24 | 4.75 | -8957.675 | 7338.900 | 5.300 | 4.815

## Slide 6: Beam Energy Errors
Dipole current uncertainty is 1 mA
  Average dipole current for beam in 5D line 3.177 A
  Average dipole current for beam in 2D line -3.817 A
  0.03% uncertainty
Field map uncertainty is 7 G cm [1,3]
Magnet Model (to find momentum from field map) uncertainty is ~0.1% [2,3]
Errors in bend angle and stray fields are still under study
Error | Uncertainty
Dipole Current | 0.03%
Field Map [1,3] | 0.1%
Magnet Model [2,3] | 0.1%
Total | 0.14
Errors Summary (so far):
[1] R. Suleiman, BTeam presentation, 2016
[2] J. Benesch, JLAB-TN-017, 2015
[3] J. Grames, JLAB-TN-17-001, 2017

## Slide 7: Summary and Future Plans
To reduce variability in setting and measuring the energy
  Put controller for Hall probe in a shielded box and move to floor (eliminate controller resets)
  Reduce number of magnets used in straight ahead orbit (reduce ambient fields)
  Improve ability to measure and reproduce beam orbit angles through dipole
    fiducialize 2D and 5D harps AND/OR
    add quads to 2D line to establish orbit reference
To set and verify beam spot position on the radiator
  Install retractable X-ray screen in front of the radiator
To reliably set the beam size at the radiator
  Improve optics model to more accurately reflect beam dynamics affecting beam size
Efforts since 2015 Engineering Run to improve dipole
  field characterization
  current and field monitoring
    ⇒ Successful setting and measurement of energies during the 2018 Engineering Run
Minor Suggested Changes for Future Engineering Runs