## Slide 1: Foil thickness measurements more uncertainties captured
29July2015

## Slide 2: Old slide: First pass at data error bars
Chart: Lebow thickness (nm)
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 750 | 750 | 625 | 625 | 625 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 355 | 355 | 355 | 355 | 355 | 225 | 225 | 225 | 225 | 50 | 50 | 50
all decent data | 971.556 | 928.094 | 916.089 | 968.0135 | 941.007 | 966.716 | 749.135 | 752.624 | 808.417 | 809.928 | 807.487 | 861.042 | 876.638 | 844.427 | 824.635 | 836.0366 | 782 | 772 | 539 | 576 | 552 | 508 | 430.7273 | 465 | 463 | 534 | 494 | 477 | 477 | 489 | 361 | 388 | 394 | 391 | 400.61 | 211.068 | 215.514 | 217.646 | 223.289 | 51.254 | 53.553 | 51.824
 | 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200
Foil | Lebow (nm)
5385B | 1000
3057C | 870
5134B | 750
7028B | 625
5275C | 500
5613A | 350
7029B | 225
6809B | 50

## Slide 3: Sources of uncertainty in measurement
Inherent limitations of the sample preparation and measurement in the FESEM
  FESEM resolution: 1.2 nm machine specifications
  Image tilt
  Uncertainty in magnification/working distance
Limitations on analyzing the image
  Pixel resolution: ± 4 pixels estimated, varies from ± 2.5 – 40 nm depending on scale
  Reproducibility between different analyses of same image
  Differences between images that should be the same
  Differences between images that should be similar (tilt, translation)
Errors or problems introduced in the preparation of the samples
  Sample mounting flatness, uniformity
  Any deformation during cleaving? Did it snap nicely or pull?
Uncertainties that have to be bounded by Lebow specifications
  Uniformity across entire sample: quote <= 2%
  Uniformity between samples within a batch: quote <= 5%
Boldface items included previously

## Slide 4: FESEM measurement uncertainties
1b:
Tilt (pitch) of ±5° introduced – none of the images appear to have pitch greater than ±5°
For a tilt of ±5°, we’d expect to measure 99.6% of the actual thickness of the foil.
0.4% uncertainty added to all for possible misalignments in tilt
1c:
Tilting can change the working distance of the edge of the foil.
Calculated for sample stage 25 mm, hinged at back edge. Change in working distance no more than 0.1 mm for tilts of ±5°, or ~1% of ~10 mm working distance
Tilt induced working distance uncertainty of 1% added
1d:
Working distance knob specifies working distance to 0.1 mm. This adjusts focus
Out of focus images can be about 1 “click” of the focus knob off.
This accounted for by 0.1 mm/working distance for each image ~1%

## Slide 5: Image analysis uncertainties
2a:
Line length assumed to be no better than ±2 pixels at each end,
Pixel – nm conversion for each image used
**Note – this has been assumed constant across images with different magnifications and still needs to be addressed**
2b:
The same image was analyzed 2-3 times typically.
Each measurement has an uncertainty
Percent uncertainty for each measurement were used, average percent uncertainty of the weighted mean of the measurements is quoted here to get one number for each foil
2c:
Some of the foils have multiple images of the same area at nominally the same tilt and location: perhaps focus or magnification was changed, but these should yield the same value since it is the same foil.
The standard deviation of data sets in nominally identically conditions used here (smaller set of data than in 2d).
2d:
Some of the foils have images made of different spots: different samples (center, edge) or translations along one sample.
The standard deviation of all the data sets at different locations is used to get this uncertainty

## Slide 6: Sample prep and Lebow consistency
3a: Sample mounting flatness and uniformity need to be considered and assigned an uncertainty. Not sure how to quantify this
3b: if we are systematically stretching or pulling the foil when we try to cleave the gold along with the silicon substrate, we should assign a value to this: it will only make them thinner, and we see a trend toward all the thicker foils measuring thinner than the Lebow quote. This systematic could affect the fit of the line.
4a: Lebow quotes the non-uniformity across each foil at less than 2%. If we are getting non-uniformities in 2d greater than 2%, it is likely our technique rather than the foils and we need to reconcile these numbers.
4b: Lebow only guarantees that sibling foils are identical to less than 5%. We could have cases where this could be our leading cause, and if we find elastic rates that deviate significantly from the FESEM measurements, this could be a factor.

## Slide 7: Error contributions
foil | Lebow (nm) | Meas (nm) | 1a: FE-SEM res. | 1b: tilt 0.4% meathick | 1c: working distance due to tilt | 1d: working distance or focus – 0.1 mm/ WD | 2a: pixel res. | 2b: re-analy. same image | 2c: diff btwn similar images | 2d: varia. btwn diff configs | 3: ??? stretch, foil mnt issues | 4a: lebow foil uniform 2% | 4b: Lebow batch uniform. 5%
5385 | 1000 | 943.7 | 1.2 | 3.7 | 9.2 | 8.5 | 40.0 | 31.0 | 29.0 | 6.5 |  | 20.0 | 50.0
3057 | 870 | 836.8 | 1.2 | 3.3 | 8.3 | 8.3 | 8.0 | 13.8 | 7.1 | 5.6 |  | 17.4 | 43.5
5134 | 750 | 774.6 | 1.2 | 3.1 | 7.7 | 6.9 | 10.0 | 13.2 | 9.1 | 9.1 |  | 15.0 | 37.5
7028 | 625 | 561.2 | 1.2 | 2.2 | 5.6 | 4.6 | 8.0 | 12.6 | 8.0 | 8.0 |  | 12.5 | 31.3
5275 | 500 | 487.6 | 1.2 | 2.0 | 4.9 | 4.3 | 8.0 | 8.4 | 3.4 | 2.9 |  | 10.0 | 25.0
5613 | 355 | 389.4 | 1.2 | 1.6 | 3.9 | 3.4 | 8.0 | 6.6 | 4.5 | 4.5 |  | 7.1 | 17.8
7029 | 225 | 215.2 | 1.2 | 0.9 | 2.2 | 1.9 | 2.6 | 2.7 | 1.9 | 1.9 |  | 4.5 | 11.3
6809 | 50 | 52.0 | 1.2 | 0.2 | 0.5 | 0.4 | 2.6 | 3.2 | 2.3 | 2.3 |  | 1.0 | 2.5
All measurements in nm
FESEM
uncertainties
Image analysis
uncertainties
Quoted batch
and
sibling uniformity

## Slide 8: Combining uncertainties
Target ladder position | foil | Meas (nm) | FESEM (all from section 1) | image anayl (all from section 2) | Lebow (from section 4: larger only) | Total uncertainty (nm)
15 | 5385 | 943.7 | 13.1 | 58.7 | 50.0 | 78.2
3 | 3057 | 836.8 | 12.3 | 18.3 | 43.5 | 48.8
4 | 5134 | 774.6 | 10.9 | 21.0 | 37.5 | 44.3
2 | 7028 | 561.2 | 7.7 | 18.7 | 31.3 | 37.2
5 | 5275 | 487.6 | 6.9 | 12.5 | 25.0 | 28.8
8, 14 | 5613 | 389.4 | 5.5 | 12.1 | 17.8 | 22.2
1 | 7029 | 215.2 | 3.2 | 4.6 | 11.3 | 12.6
13 | 6809 | 52.0 | 1.4 | 5.3 | 2.5 | 6.0

## Slide 9: Incorporating more sources of uncertainty
Chart: Nominal thickness (nm)
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 870 | 750 | 750 | 750 | 750 | 625 | 625 | 625 | 625 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 355 | 355 | 355 | 355 | 355 | 355 | 225 | 225 | 225 | 225 | 50 | 50 | 50 | 50 | 50
each image measured and averaged | 931.9272 | 915.0999 | 941.5649 | 969.0364 | 750.2828 | 808.4341 | 873.5087 | 839.798 | 833.1549 | 779.0541 | 771.048 | 567.3849 | 553.6801 | 458.4521 | 467.0947 | 517.7847 | 478.3559 | 492.0324 | 390.4303 | 391.1418 | 212.8328 | 217.3946 | 50.4647 | 52.9721
all data included | 943.7092 | 836.758 | 774.5652 | 561.1785 | 487.5807 | 389.4437 | 215.1707 | 52.0269
 | 1000 | 1000 | 1000 | 1000 | 870 | 870 | 870 | 870 | 750 | 750 | 625 | 625 | 500 | 500 | 500 | 500 | 355 | 355 | 225 | 225 | 50 | 50

## Slide 10: Incorporating more sources of uncertainty
Chart: Measured foil thickness (nm)
| 943.70920684302757 | 836.75801889050945 | 774.56515052895634 | 561.17847713696767 | 487.58073632237307 | 389.44371198870277 | 215.1706754356743 | 52.026928944137495 | 389.44371198870277
 | 33.7739 | 34.6218 | 35.6177 | 37.2456 | 38.6084 | 39.1848 | 40.9586 | 43.5056 | 39.182 | 43.2215
fit to guide | 44.0308 | 43.3453 | 42.6808 | 42.0363 | 41.411 | 40.8041 | 40.2147 | 39.642 | 39.0855 | 38.5443 | 38.0179 | 37.5058 | 37.0072 | 36.5217 | 36.0488 | 35.588 | 35.1388 | 34.7008 | 34.2736 | 33.8568 | 33.45 | 33.0528 | 32.665

## Slide 11: Foil “870 nm” looked at again
Chart: Thickness (nm)
| 1 | 2 | july | 1 | july | 1 | july | 1 | july | 0 deg | 0 deg | 0 deg | 2.5 deg | 2.5 | 0 deg | 0deg | 0 deg | 0deg | spot 1 | spot 1 | spot1 | spot 2 | spot2 | spot 2 | spot 2 | spot 4 | spot4 | center | center | center | center | center | edge | edge | edge | edge
 | 749.135 | 752.624 | 743.961 | 876.638 | 868.33 | 844.427 | 834.98 | 836.0366 | 828.008
mean | 836
Mean: 836
Multiple analyses of the same image are grouped