## Slide 1: UCSB Photocathode tests in MicroMott
Marcy Stutzman
14 December 2023

## Slide 2: MicroMott functional again
Root cause:
Metallization on CEM changed
Existing design puts HV at or nearly touching support
Mechanical support
Biased 0-300V
Collector: 3kV
Back: 3kV
Front: Ground

## Slide 3: MicroMott functional again
Root cause:
Metallization on CEM changed
Existing design puts HV at or nearly touching support
Mechanical support
Biased 0-300V
Collector: 3kV
Back: 3kV
Front: Ground
Problem area

## Slide 4: UCSB cathode results
Sample 144
  InAlGaAs/AlGaAs SSL
  Higher strain
Sample 199: in progress
  InAlGaAs/AlGaAs SSL
  Standard strain, 505C growth temp
Chart: wavelength (nm)
| 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730 | 740 | 750 | 760 | 770 | 780 | 790 | 800
QE 199 heat 1 610 lp | 7.2046 | 6.7556 | 6.2933 | 5.9109 | 5.4862 | 5.0668 | 4.4671 | 3.8085 | 2.1441 | 1.1996 | 0.9862 | 0.88418 | 0.80622 | 0.72084 | 0.60996 | 0.44441 | 0.22478 | 0.11053 | 0.1078
QE 199 heat 1 715 nm longpass | 1.4827 | 1.3454 | 1.2524 | 1.1769 | 1.0985 | 1.0548 | 0.98585 | 0.92706 | 0.85017 | 0.77035 | 0.68754 | 0.57491 | 0.43882 | 0.29963 | 0.20739 | 0.14664 | 0.13368 | 0.14181 | 0
polarization | 78.96 | 81.66 | 84.93 | 74.58 | 85.48 | 90.17
Chart: wavelength (nm)
| 710 | 715 | 720 | 725 | 730 | 735 | 740 | 745 | 750 | 755 | 760 | 765 | 770 | 775 | 780 | 785 | 790 | 795 | 800
QE 3rd heat 715 longpass | 1.2772 | 1.2035 | 1.1156 | 1.0276 | 0.95272 | 0.89064 | 0.85213 | 0.74945 | 0.54954 | 0.32539 | 0.16064 | 0.072842 | 0.031896 | 0.015595 | 0.008547 | 0.0056438 | 0.0041193 | 0.0035632 | 0.0032566
QE 2nd heat | 3.7007 | 3.5186 | 3.3808 | 3.2541 | 3.1288 | 2.9835 | 2.8386 | 2.6825 | 2.5576 | 2.424 | 2.2918 | 2.1624 | 2.0213 | 1.8949 | 1.7675 | 1.653 | 1.5391 | 1.4421 | 1.3553 | 1.2841 | 1.2227 | 1.1594 | 1.0816 | 1.004 | 0.93009 | 0.87155 | 0.82506 | 0.75532 | 0.61726 | 0.42449 | 0.22932 | 0.11005 | 0.04977 | 0.025661 | 0.011477 | 0.0098738 | 0.0067617 | 0.0084749 | 0.0084724
1st heat QE | 0.027
2nd heat | 80.4 | 86.14 | 87.2 | 89.6 | 86.26 | 82.8 | 83.8013 | 78.7 | 81.69 | 84 | 81.39 | 73.42 | 75.3 | 75.5 | 85.35 | 92.8 | 84.14
pol 3rd heat | 75.2 | 75.77 | 76.09 | 75.303 | 74 | 76 | 70.48 | 69.39 | 53.4
avg 2nd heat | 80.4 | 86.14 | 87.2 | 89.6 | 84.4671 | 80.6301 | 81.8561 | 80.7207 | 82.8688
first run | 79.78

## Slide 5: UCSB sample 1 (520C): 2022
HH1
LH1
HH2
 | Polarization | QE@max
Sample 1: 520C | 80% | 0.15%
Sample 2: 505C | 90% | 0.3%
Sample 3: high strain | 89.5% | 0.42%
Polarimeter back in operation
UCSB samples show promise
They can add DBR structures
If they get a good formula for InAlGaAs, easy to get commercial growers -- no Phosphorus residue