## Slide 1: RTP Driver Prototype Crosstalk tests
Jim Kortze 20 October 2024

## Slide 2: Summary:
Tests done to identify HV channel crosstalk
Tests conducted using Rigol MSO5204 oscilloscope and 1000:1 HV divider box.
For each test, the ground lead of the divider box was connected in parallel to the common lead of the RTP box and crystal being tested, and the lead under test connected to the HV divider signal input. All leads not under test were connected to their respective crystal connections
For each lead, four measurements were made: All supplies on, DUT supplies on with all others off, Pos on/ Neg off, Pos off/ Neg on

## Slide 3: RTP1 POS LEAD ALL ON
518.06V/-929.22V

## Slide 4: RTP1 POS LEAD OTHERS OFF
518.06V/-929.22V

## Slide 5: RTP1 POS LEAD POS ON/ NEG OFF
518.06V/-8.223

## Slide 6: RTP1 POS LEAD POS OFF/ NEG ON
16.446V/-929.22V

## Slide 7: RTP1 NEG LEAD ALL ON
929.22V/ -518.06V

## Slide 8: RTP1 NEG LEAD OTHERS OFF
929.22V/ -518.06V

## Slide 9: RTP1 NEG LEAD POS ON/ NEG OFF
929.22V/ -15.446V

## Slide 10: RTP1 NEG LEAD POS OFF/ NEG ON
16.446V/ -518.06V

## Slide 11: RTP2 POS LEAD ALL ON
509.84V/ -921.00V

## Slide 12: RTP2 POS LEAD OTHERS OFF
509.84V/ -929.22V

## Slide 13: RTP2 POS LEAD POS ON/ NEG OFF
518.06V/ -16.446V

## Slide 14: RTP2 POS LEAD POS OFF/ NEG ON
16.446V/ -929.22V

## Slide 15: RTP2 NEG LEAD ALL ON
921.00V/ -518.06V

## Slide 16: RTP2 NEG LEAD OTHERS OFF
921.00V/ -518.06V

## Slide 17: RTP2 NEG LEAD POS ON/ NEG OFF
929.22V/ -16.446V

## Slide 18: RTP2 NEG LEAD POS OFF/ NEG ON
16.446V/ -526.28V

## Slide 19: Conclusions
No notable signs of crosstalk between measurements
The largest variance between supply measurements was less than 2%, which is attributed to measurement noise