## Slide 1: DSG NPS Contributions Status
Aaron Brown
Detector Support Group
07/09/2020

## Slide 2: Contents
DSG Contributors
CAEN High Voltage Testing
High Voltage Divider Cables
EPICS Controls & Monitoring Screens
Upcoming Work
Conclusion

## Slide 3: DSG Contributors
CAEN high voltage testing
  Hardware
    Mindy Leffel, Marc McMullen
  Testing and analysis
    Aaron Brown, Pablo Campero, George Jacobs, Marc McMullen
High voltage divider cables fabrication
  Mindy Leffel
Environmental monitoring
  Peter Bonneau, Tyler Lemon
EPICS CSS-BOY screens development and testing
  Mary Ann Antonioli, Aaron Brown, Brian Eng, Tyler Lemon

## Slide 4: CAEN High Voltage Testing
In process of testing 34 CAEN high voltage modules and two SY4527 crates
Tests include:
  Stability using GECO 2020 (in progress – 32 of 34 completed)
  Stability using built-in EPICS server (planned)
  Current Trip (planned)
    Verify channels trip
    If Imeas. > Iset
52-pin Radiall Connector Adapter
CAEN High Voltage Modules
DSG SHV Load Chassis

## Slide 5: CAEN High Voltage Testing
Stability test details
  1500 V
  2 MΩ resistor load
  24 hours
  Failures found
Two firmware updates
  Module firmware update fixed channel latency issues
  Crate firmware update fixed EPICS communication issues

## Slide 6: CAEN High Voltage Module Testing
Average readback voltage of each channel of module #262
  All channels within manufacturer’s specifications
  Voltage spread (V/1500 V) ~0.02%

## Slide 7: CAEN High Voltage Module Testing
Average readback voltage of each channel of module #262

## Slide 8: CAEN High Voltage Module Testing
Average readback current of each channel of module #262
  All channels above expected current value
  Offset perhaps due to lower resistor value, 1.98 M instead of 2 M
  Current spread (I/757 A ) ~ 1%
Issue is being investigated

## Slide 9: High Voltage Divider Cables
Fabricating 1100 high voltage divider cables
390 cables fabricated (35% completed)
All components on hand
Cable fabrication sequence: left to right
Finished cable

## Slide 10: EPICS Controls & Monitoring Screens
Development of EPICS CSS-BOY screens
Screens will include
  Overview Screen
  Environmental Readback Screens
  Chiller Temperature and Flow Screens
1080 PMT HV indicators
In a 3x3 matrix

## Slide 11: EPICS Controls & Monitoring Screens
High Voltage Readback Screen
30 x 36 grid of voltage and current readback of each PMT
PMT located in
Column -4, Row -2 from the origin

## Slide 12: Upcoming Work
Continue cable fabrication
Continue CAEN high voltage system testing
  Crates for EPICS interface
  Verify module specifications
Investigation of sensors and instrumentation for environmental monitoring
Continue development of EPICS CSS-BOY controls & monitoring screens
Development of WEDM screens

## Slide 13: DSG Website
NPS section of DSG Hall C Technical Documentation
Page will be updated as progress is made

## Slide 14: Conclusion
Contributions to NPS by DSG
  Testing CAEN crates and modules
  Fabricating and testing high voltage divider cables
  Developing environmental monitoring
  Developing EPICS controls & monitoring screens
Significant progress is being made

## Slide 15
Thank You!