## Slide 1
NPS Wrapping procedure
2021/02/02

## Slide 2
Reflective sheet 70 µ
No overlap (gap 0,5mm to be adapted for all crystals sizes
No reflective sheet on both extremities (will be glued on the plastic block part for front side and no reflective sheet side PMt)
Reflective sheet 200 mm length

## Slide 3
Wrapping tape Crystal side PMt : tape 60 µ
Mylar aluminised tape 60µ (17 mm width)
No overlap for tape
Side PMt
17 mm
Cutting tape 90°/ cutting reflective sheet

## Slide 4
Mylar aluminised tape 60µ (17 mm width)
(no overlap)
Wrapping tape crystal side reflective sheet :tape 60 µ
Cutting tape 90°/ cutting reflective sheet
17 mm
17 mm

## Slide 5
Wrapping Longitudinal tape 60µ
Longitudinal Tape mylar aluminised 60µ 12 mm width
No overlap for tape gap 1 mm reasonable
12 mm

## Slide 6
Total wrapping thickness
Total vertical wrapping thickness = 260 µ (2x70 + 2x60)
Total horizontal wrapping thickness = 260µ (2x70 + 2x60)

## Slide 7
Tedlar Wrapping
Mylar aluminised tape 12 mm width:
@ 90 ° / wrapping tape
TEDLAR wrapping
Front stacking area 17 mm
Rear stacking area 17 mm
Mu metal area 40 mm
Wrapping tape
1 mm gap between stacking area and Tedlar (no overlap for good stacking)
1 mm gap between Mu metal area and Tedlar (no overlap for good stacking)

## Slide 8
Mylar 60µ 17mm width for adjustment
Mylar 60µ 17mm width for adjustment
Number of layers variable ; no overlap on both Top and bottom sides
Adjustment tape : mylar 60µ 17mm width

## Slide 9
Adjustment Tape : mylar 60µ 17mm width
For a alveoli 21mm inner dimension (nominal dimension) : the Gap is (between alveoli and crystal wrapped)
Gap min (21- 20,86=0,14mm)
Gap Max (21- 20,69=0,31mm)
Mean Gap Siccas ( 21-20,80 = 0,20mm)
Mean Gap Crytur ( 21-20,72 = 0,28mm)
Mean gap :
Siccas 0,20 ≈ 3 x 0,06 = 18µ 1 turn +1/4 turn of tape 60 microns
Crytur 0,28 ≈ 4 x 0,06 = 24µ 1 turn + ½ turn of tape 60 microns
60µ
These gaps are based from Jlab measurement values
See next slide for calculations

## Slide 10
The measurement for CRYTUR gives :
mean value : 20,46mm (20,72 with wrapping)
Max value : (20,46 + 3σ) = 20,46+ 3*0,01 =20,49mm (20,75 with wrapping)
min Value : (20,46-3 σ) = 20,46-3*0,01 = 20,43 mm (20,69 with wrapping)
The measurement for SICCAS gives :
mean value : 20,54mm (20,80 with wrapping)
Max value : (20,54 + 3σ) = 20,54+ 3*0,02 =20,60mm (20,86 with wrapping)
min Value : (20,54 -3 σ) = 20,54 - 3*0,02 = 20,48 mm (20,74 with wrapping)
For a alveoli 21mm inner dimension (nominal dimension) : pitch 21,5 mm the Gap is
Gap min (21- 20,86=0,14mm)
Gap Max (21- 20,69=0,31mm)
Mean Gap Siccas ( 21-20,80 = 0,20mm)
Mean Gap Crytur ( 21-20,72 = 0,28mm)
The extrema gives with wrapping :
Max value : Siccas 20,86mm
min Value : Crytur 20,69mm
Gap between alveoli and crystal (theorical value)
Measurement values from Jlab