## Slide 1
NPS mechanical design
06/20/2019 Jlab meeting

## Slide 2
Design presentation
Wrapping crystals
Pitch alveolis
Mechanical tests and simulations for carbon alveolis
Fiber tests and mounting
Cables and PCB adjustement (Moke up )
Cooling

## Slide 3
Design presentation

## Slide 4
General front Axio view
Calorimeter

## Slide 5
General back axo view
30 Interface PCB and cables

## Slide 6
Crystal +PMt+cabling channel
Reference plate
Carbon plates honeycomb structure
PMt support plate
independant PMt assembly
Interface Pcb (HV,Anode and LED)
Fiber Cut view

## Slide 7
Independent PMt assembly
Unscrewing of 1 alone screw
Then pull The Block
Cut view dismounting Screw
Long Captive screw

## Slide 8
Calorimeter
36 raws of 30 crystals =1080 crystals 20,5 x 20,5 x 200 mm
Cooling copper plates around the calorimeter + insulation with Foam in order to maintain the T° @ 18°=/- 0,1 °(Study in progress)
Pitch between 2 crystals = 21,5 mm in each direction (gap between 2 crystals without Wrapping = 1mm)

## Slide 9
Calorimeter support
Carbon alveolis (20 mm width)
Aluminum frames (20 mm width)
Spacers and threaded rods between the 2 frames
Back side alveolis
Front side alveolis
Holes for plastic stop parts

## Slide 10
Wrapping crystals

## Slide 11
Total wrapping thickness
Total vertical wrapping thickness = 260 µ (2x70 + 2x60)
Total horizontal wrapping thickness = 260µ (2x70 + 2x60)
Hypothesis : no air between crystal and reflective sheet

## Slide 12
Front reflective sheet
Front reflective sheet glued on the plastic block part

## Slide 13
Pitch alveoli

## Slide 14
Crytur dimensions

## Slide 15
Siccas dimensions

## Slide 16
We choose an alveoli 21mm inner dimensions: so the pitch is 21,5 mm (gap between 2 crystals : 1mm)
Gap min (21- 20,91 =0,09mm)
Gap Max (21- 20,74=0,26mm)
Mean Gap Siccas ( 21-20,85 = 0,15mm)
Mean Gap Crytur ( 21-20,77 = 0,23mm)
Gap between alveoli and crystal (practical value)
Max value : Siccas = 20,91 mm
mean Value : Siccas = 20,85 mm
Min Value : Siccas = 20,79 mm
max Value : Crytur= 20,80 mm
mean Value :Crytur= 20,77 mm
min Value : Crytur= 20,74 mm
The final value of the crystals with 260µ for wrapping +50µ (du to air and corners of wrapping) is:

## Slide 17
Mechanical tests and simulations on carbon honeycomb structure

## Slide 18
Shearing simulation on 1 pillar with no gap under the pillar
NO GAP
Shearing Max = 1,4Mpa
X 36
P/2 = 7 N
P/2
P/2
1,4 Mpa @ each stage
For each horizontal plate (even the bottom one) the stress will be the result of 1 crystal weight only = 1,4 Mpa
No gap
Max stress = 1,4 Mpa
Shear resistance carbon plate = 90Mpa (Good Fellow value) security coeff =64
Very secure

## Slide 19
Shearing simulation on 1 pillar
36P/2
contact
pillars
Shearing
Shearing
125 N
Max stress = 48Mpa
Shear resistance carbon plate = 90Mpa (Good Fellow value) security coeff =2
Safe but not secure enouth
In reality its necessary to have no gap under the Vertical plates in order to reduce the stress
In this simulation we want to see what would happen if we have gap under the vertical carbon plates :All the load will be supported by the bottom horizontal plate.

## Slide 20
Test in workshop: schearing in case of gap under the vertical carbon plates
additionnal weight to obtain 76 Kg (3 columns weight)
Gap
Conclusion : No damage; it confirm the very good Shearing resistance of these carbon plates

## Slide 21
Test in workshop: deformation alveoli with 1 crystal
Bending : 4 measurement points : A=0,05mm ; B0,02mm ; C=0,05mm ; D=0,04mm
A
B
C
D
Aluminum part
additionnal weight to obtain 0,7 Kg (PBWo4)
Conclusion: The deformation is very small

## Slide 22
Buckling
Hypothesis : first raw without crystal (no real scenario)
We consider that all the stages above the first raw are compact
The vertical plates are blocked strongly at the bottom
F=125x30 =3750N
Buckling factor = 1,36 (> 1 ) Ok
With the moke up under load we tried to remove the 3 bottom crystals:
1st conclusion : it was possible to remove them
2nd conclusion: after a small time we have seen a movement of the vertical plates : buckling. Probably the load was not exactly centered so
It is evident not to remove one bottom or intermediate raw when crystals above !

## Slide 23
Fiber tests and mounting

## Slide 24
Fiber side crystal
Silicon Cable gland in order to block the fiber
Metal ferrule screwded in aluminum support

## Slide 25
Fiber side PCB
inserted and no glued in the plastic Ferrule
Pinched between 2 aluminum plates
Protected with plastic tube
Advantage:
1) The fibers can be dismounted if necessary
2) The PCB can be dismounted separately if necessary
Length fiber = 51cm (has to be adjusted with the PCB prototype)

## Slide 26
Cables and PCB adjustement (Moke up )

## Slide 27
Interface PCB

## Slide 28
Interface PCB
PCB in front of its PMt column (Top view)
PCB shifted 15 cm of its PMt column (Top view)
15 cm
Access for PMt assembly exchange

## Slide 29
108O Lemo connectors for signal
HV PCB
LED distribution
Electronique Input/Output

## Slide 30
Cooling

## Slide 31
Thermical aspect
Heating Zone : 500mW/channel = 540 W
Crystals Zone : @ 18°+/-0,1°
External Zone : T° = 22°c
Intermediate Zone : Thermical Resistance
Intermediate Zone : Made of Glass , Vacuum (PMt)
Depron (Thermal insulator)
+ cold Copper plate
inserted and no glued in the plastic Ferrule

## Slide 32
Cooling
Electronic cooling ; Device delivered
Test in progress with the moke-up
Crystals cooling ; Study in progress
Test will be done with the moke-up