## Slide 1: CPS Prototype Testing
100 to 300 Watts
June 29, 2023

## Slide 2: Brazing - End View Initial
Cusil 72%Ag 28%Cu Braze
Melting point 780C
K= 371W/m-K Modeled in Ideas thermal solution
Back End
Front End

## Slide 3: Setup

## Slide 4: Test Setup
Cartridge Heater Resistance 28 and 28.3 ohms

## Slide 5
A thermal analysis was performed using both Ideas and Fluent software. Both agreed well with each other on temperature profiles
Materials: C1020 OFHC Copper
Absorber
Base
inlet
outlet
water tube
Top section
Heating elements
 | Surface area of a heating tube |  | m^2 | 7.200E-03
 | Diameter of Heater |  | m | 1.905E-02
 | Diameter of H20 tube |  | m | 3.200E-03
 | Area of Tube |  | m^2 | 8.040E-06
 | Mesh element Size |  | m | 6.350E-03
 | Density of Water |  | kg/m^3 | 9.982E+02
 | Cp of H20 |  | J/kg-K | 4.182E+03
 | Conductivity H20 |  | W/m.K | 6.000E-01
 | Density of Cu |  | kg/m^3 | 8.978E+03
 | Cp of Cu |  | J/kg-K | 3.810E+02
 | Conductivity Cu |  | W/m.K | 2.876E+02
 | Fluent model -Tube conditions were prefect - no roughness factor |  |  | 

## Slide 6: Thermal Camera Captures

## Slide 7: Measured vs Model
Chart: Top Temp
| 100 | 105.7 | 252 | 265
Measured | 57.7 | 65.3 | 117.4 | 126.4
Model | 58.9 | 130.6 | 231.1
Top
End
Side

## Slide 8: Measured vs Model
Chart: Side Temp
| 252 | 265
Measured | 41.1 | 41.7
Model | 26.1 | 34.8 | 46.9
Chart: End Temp
| 100 | 105.7 | 252 | 265
Measured | 28.9 | 31.2
Model | 30.4 | 46.7 | 70.6

## Slide 9: Steady State reached ~ 30 minutes
Chart: Temp Rise for 100 & 105.7W
| 1 | 2 | 3 | 4 | 5 | 7 | 8 | 10
T1 (end) | 22.5 | 23.2222 | 23.9444 | 25 | 25.7222 | 27.5 | 27.8333 | 28.9444
T2 (top) | 27.2222 | 35.5 | 40.3889 | 44.6667 | 48.1111 | 53.2222 | 54.8889 | 57.7222
T2 (top) | 38.4 | 53.6 | 60.6 | 63.2 | 64.6 | 64.8 | 65.3
T1 (end) | 23.8 | 27.3 | 29.6 | 30.3 | 30.9 | 30.8 | 31.2
Model Top | 68.7 | 68.7 | 68.7 | 68.7 | 68.7 | 68.7 | 68.7 | 68.7
Chart: Temp Rise for 252 & 265W
| 1 | 5 | 10 | 15 | 20 | 25 | 30
T1 (side) | 23.8333 | 31.6111 | 37.6667 | 39.7778 | 41 | 41.1667 | 41.1111
T2 (top) | 50.8889 | 87.8333 | 105.8333 | 114.1667 | 117.3333 | 118.2222 | 117.4444
T1 (top) 265W | 126.4444
Model Top | 130 | 130 | 130 | 130 | 130 | 130 | 130

## Slide 10: Water Flow
Pressure measured values:
Flow 0.625GPM
Static Pressure ~50psig
Flowing inlet pressure 42psig.
Flowing outlet pressure 0 psig
Fluent Model:
Pressure drop 22psig for all 0.625GPM cases (No surface roughness used in model.)
Temperatures:
Inlet 22.2C
Outlet 22.2C (did not increase for the low power test - up to 300W)
Fluent Model:
Inlet 23.3C
Outlet Increased by 2.1C for 300W and 4.1C for 600W

## Slide 11
Chart: Fluent LCW Temp Rise
| 300 | 600 | 1200
Fluent | 2.1 | 4.1 | 8.3

## Slide 12: Thermal Expansion
B
A
Chart: Thermal Deflection Point A
| 100 | 105.7 | 252 | 265
Measured | 0 | 0.5 | 0.5
Model A | 0.38 | 1.1 | 1.95
Setup of dial indicators needs improvement
A mil is 0.001”

## Slide 13
Resulting temperature profile of the fluid.
Power 300W
The maximum temperature of the fluid was 297.5K (24.4oC).
The ΔT between the inlet and outlet was 297.5K – 295.4K = 2.1K

## Slide 14
Resulting profile of the water pressure
Power = 590W The pressure of the fluid at the outlet was 7587Pa (1.1psi).
The pressure of the fluid at the inlet was 159543Pa (23.1psi).
ΔP = 23.1 – 1.1 = 22.0 psi.

## Slide 15: Thermal Contact Paste
Thermal Paste was used around Heater elements to provide better contact surface.
At higher heat loads approaching 600W, The paste stared to flow and give off an odor.
This concerned people nearby and further testing at higher heat loads was halted.

## Slide 16: Notes
Testing is still in the infancy stage.
Test setup and measurement devices could be improved.
Higher heat loads should be tested for thermal deflections, solder integrity, water cooling effects and model proofing.
Fluent model needs to be fine tuned for flow parameters to match test case.

## Slide 17: Extra Slides

## Slide 18: Heaters
Chart: Current / Heater element Power is total
| 10 | 20 | 30 | 47 | 65 | 70 | 80 | 90 | 100
Current | 0.41 | 0.86 | 1.36 | 2.17 | 3 | 3.031 | 3.464 | 3.897 | 4.35
Power | 10.184 | 42.6169 | 105.6648 | 265.064 | 509.1053 | 571.1474 | 736.8061 | 923.5723 | 1128.4688