## Slide 1: DSG NPS Update
Aaron Brown
Detector Support Group
8/26/2020

## Slide 2: Contents
Fabrication, Procurement, and Testing
HV Rating Concerns and Research
  HV Schematic
  HV Connection Parts List
Conclusion

## Slide 3: DSG NPS Work Update
HV divider cables fabrication status
  700 of1100 fabricated
Procurement status for HV multi-conductor cable fabrication
  From CAEN: 40 Radiall 52-pin connectors, insertion/extraction tool, and crimper tool
    PR submitted and signed, lead time: ~12 weeks
    Assigned buyer: Albert DeChristopher
  From General Wire Products: Forty-two 42-conductor cables
    Cut to 142’
    PR submitted, lead time: 4-5 weeks
    Assigned buyer: Jami Anthony
CAEN voltage stability testing
  All 34 modules tested with GECO 2020
    All channels of 12 modules passed testing
    Four modules have major problems
    Eighteen modules have at least one problem channel
  22 of 34 modules tested with EPICS
    Analyzing data

## Slide 4: DSG NPS Work Update (cont’d)
EPICS Controls & Monitoring screens development
  Environment monitoring screens
  Voltage and current readback screen
Researching sensors and related instrumentation for environment monitoring

## Slide 5
HV Connection Schematic
HV board
PMT
IPD1-05-S-K
Fischer S 102 A 018
SMA
PMSD-15-206-K-15.40-D-LUS
140’ GPW 42 cond. 030-23904
Radiall A996 52 pin
IPBD-115-H1-T-D-K
IPBD-108-H1-T-D-K
CAEN 7030TN
RG-174
RG-174
Interface card
Fischer D 102 A 018
SMA
PMT Signal
IPBT-108-H1-T-D-K
IPBT-115-H1-T-D-K
HV supply cable
PMSD-08-206-K-15.40-D-LUS
(12.5 kV)
(424 V)
(424 V)
(5 kV)
PMT HV/Signal Y cable
(565 V)
(565 V)
(600 V)
(600 V)
(8 kV)
(8 kV)
(30 V)
(387 V)
IPBT-108-H1-T-D-K
(424 V)
HV rating concerns
Researching alternates

## Slide 6: NPS Cable Components
NPS Assembly/Board | Component | Part Number | Voltage Rating (V)
HV supply cable | CAEN HV connector | Radiall A996 52 pin | 12.5 kV
 | HV multiconductor - 42 cond. | GWP 030-23904 | 5 kV
 | Samtec 8 pin (male cable mount) | IPBD-08-D-K | 424 V
 | Samtec 15 pin (male cable mount) | IPBD-15-D-K (x2) | 424 V
HV board | Samtec 8 pin (female pcb mount) | IPBT-108-H1-T-D-K (1 per side of patch panel) | 565 V
 | Samtec 15 pin (female pcb mount) | IPBT-115-H1-T-D-K (x2 per side of patch panel) | 565 V
Samtec ribbon cables | Samtec 8 conductor ribbon cable | PMSD-08-206-K-15.40-D-LUS | 600 V
 | Samtec 15 conductor ribbon cable | PMSD-15-206-K-15.40-D-LUS (x2) | 600 V
Interface Card | Samtec 8 pin (female pcb mount) | IPBT-108-H1-T-D-K | 565 V
 | Samtec 15 pin (female pcb mount) | IPBT-115-H1-T-D-K (x2) | 565 V
 | Fischer jack connector | D 102 A018 | 8 kV
 | SMA right angle jack | 142-0701-501 (1 per side of patch panel) | 335 V
PMT HV/Signal Y cable | Fischer plug connector | S 102 A018 | 8 kV
 | SMA female cable mount | R19-003-04-032611011 (1 per side of patch panel) | 335 V
 | Coax cable | RG-174u | 30 V
 | Samtec 5 pin male connector | IPD1-05-S-K | 387 V

## Slide 7: Conclusion
HV divider cable fabrication ongoing
Procurement of HV cable components underway
Controls & Monitoring screens development progressing

## Slide 8
Thank You