## Slide 1: Hall A Tritium Target Part 3: Response to 2010 Committee
Dave Meekins
September 15, 2015

## Slide 2: Summary of Action Items
From the report of the committee 44 action items were derived
  (7) Administrative
  (20) Target cell
  (9) Scattering chamber ventilation and beamline
  (8) ESH tasks
We address these by group
  Green: Fully complete no issues
  Yellow: Complete but issues or N/A
  Incomplete

## Slide 3: Administrative
Appoint a lead person at JLab, a design authority, who will emphasize the engineering aspects of the target system and who will be responsible for the design, fabrication, procurement, installation and operation of the system
Establish an engineering team
Develop clear responsibilities for INL, JLab, and collaboration; and determine who will sign off on safety checkout plans
Administrative requirement that beam blow-up optics are used in experiment
Administrative limit on beam trip rate when cell is in the beam
Make a provision that if the H or He target fails, a failure mode determination is made before the experiment continues with tritium gas
Extensive review of final cell design and test results should be performed by JLab, Argonne and outside experts

## Slide 4: Target Cell-1
Develop a modular target design
Use a 1000 Ci source
Determine maximum target window thickness based on physics requirements
Target cell should comply with ASME Boiler and Pressure Vessel Code, sec VIII Div 2 2007
Cell should be filled offsite and be designed to survive transport.
Cell should survive cyclic loading (beam trips)
W collimator should be better
Valves should have all metal wetted parts
Cell must sustain a full vacuum load
Target system should be designed to be cooled with 90K N2

## Slide 5: Target Cell - 2
Verify that Al 2219 is a suitable material for tritium gas
Determine strength of welded Al 2219
Weld coupons should be tensile and bend tested
Al –ss transition piece should be purchased and elbows should be used
An elastic plastic model of the cell (ASME D2 5.2.4) should be used
Proof tests on more than 3 endcaps and at least one complete cell should be performed
Heat cycling tests with a tritium loaded cell should be performed for a period of 6 months
If insufficient information exists on beam shock wave tests, perform such tests at JLab
Cell must survive off-normal beam conditions for at least 3 times the amount of time that it takes for an interlock to turn the beam off
Determine DOT and DOE regulations for shipping filled target cells to JLab

## Slide 6: Scattering Chamber, Ventilation, Beamline
Secondary containment should be physically isolated from beamline
The scattering chamber pumps shall be vented through tritium stack
The scattering chamber should be monitored for high and low levels of tritium
A U getter bed should be attached to the scattering chamber
An additional long collimator should be placed upstream
Dedicated vent pumps/fans and lines should be installed over the scattering chamber and used for installation and removal procedures

## Slide 7: Scattering Chamber, Ventilation, Beamline -2
Airborne radioactivity detectors interlocked with the vent/fan stack system should be used
Manual “alarm” buttons in the hall and counting house for the ventilation Additional beam raster detector and interlock system

## Slide 8: ESH Tasks
Establish baseline for detectable tritium at the JLab site
Develop algorithm for safety involving amount of tritium, beam current, beam time
Worst-case scenarios for worker exposure and all dose calculations should be analyzed or calculated by qualified personnel
A more detailed assessment of impact of target loss on Hall A should be performed by qualified personnel
Use a 15 m elevation and a vertical tritium stack as necessary for limiting site boundary doses
Use the ICRP-68 dose coefficient of 1.8E-11 Sv/Bq reference for exposure evaluations
Use 10 mrem as the maximum allowed site boundary dose
The risk analysis should follow tables 4.2-4.5 of JLab’s FSAD, rev. 6 and use realistic target failure probabilities

## Slide 9: Summary
Develop clear responsibilities for INL, JLab, and collaboration; and determine who will sign off on safety checkout plans
  Being addressed now
Extensive review of final cell design and test results should be performed by JLab, Argonne and outside experts
  Being addressed now (shall also undergo full DA review as required by JLAB PS policies)
Heat cycling tests with a tritium loaded cell should be performed for a period of 6 months
  Exposure studies underway at SRS/SRNL